## ProtocolsList
Protocol name | Range name | Series description | Ref. kV | kV | Quality ref. mAs | (Eff.) mAs | (Eff.) mAs(Tube A) | (Eff.) mAs(Tube B) | CARE kV | Tissue of Interest (Slider position 1 - 12) | Dose modulation | CARE Dose type | CTDIvol (mGy) | FAST Adjust: Upper limit scan time | FAST Adjust: Lower limit max. mAs | Dose Notification value CTDIvol (mGy) | Dose Notification value DLP (mGy*cm) | Rotation time (s) | Scan time (s) | Delay (s) | Pitch | Feed per scan (mm) | Slice (mm) | Acq. | Effective Slice (mm) | Position increment (mm) | No. of images | Kernel/Algorithm | Window | API | Comment 1 | Comment 2 | Auto transfer 1 | Auto transfer 2 | Auto transfer 3 | syngo.via: Workflow | syngo.via: Data role | Scan start | Scan end | Pulsing | Pulsing start | Pulsing end | Best phase | Phase start | Multi phase
Customized scan protocols
Head
11010_ang_Hea_PED (Child)
11010_ang_Hea_PED (Child) | Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.23 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS
| PreMonitoring |  | 120 | 120 | 390 | 20 |  |  | Off | 3 | Off | CARE Dose4D | 2.78 |  | 25% |  |  | 0.5 |  | 2 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | PreMonitoring 10.0 H30s | [22 empty] | 10 |  | 1 | H30s medium smooth | Head Soft | [13 empty] | manual
| Monitoring |  | 120 | 120 | 390 | 20 |  |  | Off | 3 | Off | CARE Dose4D | 83.43 |  | 25% |  |  | 0.5 |  | 10 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | Monitoring | [22 empty] | 10 |  | 30 | H30s medium smooth | Head Soft | [13 empty] | manual
| art Hea ph010 |  | 120 | 120 | 390 | 175 |  |  | Off | 3 | Off | CARE Dose4D | 26.77 | 14s | 25% | 60 |  | 0.5 |  | 2 | 0.6 |  | 2 | 128x0.6mm | [15 empty] | Off
|  | art Hea ph010 | [22 empty] | 2 | 2 | 61 | J40f medium | Head Soft |  |  | X-CARE | GEPACS | [9 empty] | manual
|  | art Hea cor | [22 empty] | 2 | 2 | 126 | J40f medium | Head |  |  | X-CARE | GEPACS | [9 empty] | manual
|  | art Hea sag | [22 empty] | 2 | 2 | 126 | J40f medium | Head |  |  | X-CARE | GEPACS | [9 empty] | manual
|  | art Hea mip | [22 empty] | 7 | 2 | 126 | J40f medium | Head |  |  | X-CARE | GEPACS | [9 empty] | manual
|  | art Hea cor mip | [22 empty] | 7 | 2 | 126 | J40f medium | Head |  |  | X-CARE | GEPACS | [9 empty] | manual
|  | art Hea sag mip | [22 empty] | 7 | 2 | 126 | J40f medium | Head |  |  | X-CARE | GEPACS | [9 empty] | manual
|  | art Hea ph010 :-D | [22 empty] | 0.6 | 0.6 | 201 | J40f medium | Head Soft |  |  | X-CARE | ROUTER | GEPACS | [8 empty] | manual
11100_Hea_PED (Child)
11100_Hea_PED (Child) | Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.23 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS
| 3m Hea ph110 |  | 120 | 120 | 390 | 145 |  |  | Off | 3 | Off | CARE Dose4D | 22.18 | 14s | 25% | 60 |  | 0.5 |  | 2 | 0.6 |  | 3 | 128x0.6mm | [15 empty] | Off
|  | 3m Hea ph100 | [22 empty] | 3 | 1.5 | 81 | J30f medium smooth | Head Soft |  |  | X-CARE | GEPACS | [9 empty] | manual
|  | 3m Hea bon | [22 empty] | 3 | 1.5 | 81 | H60f sharp FR | Bone |  |  | X-CARE | GEPACS | MCHNeuroBodyAW | [8 empty] | manual
|  | 3m Hea cor | [22 empty] | 3 | 1.5 | 134 | J30f medium smooth | Head Soft |  |  | X-CARE | GEPACS | [9 empty] | manual
|  | 3m Hea sag | [22 empty] | 3 | 1.5 | 134 | J30f medium smooth | Head Soft |  |  | X-CARE | GEPACS | [9 empty] | manual
|  | 3m Hea ph100 :-D | [22 empty] | 1 | 1 | 201 | J30f medium smooth | Head Soft |  |  | X-CARE | ROUTER | GEPACS | [8 empty] | manual
11150_Hea_wo_cst_PED_N09_02P (Child)
11150_Hea_wo_cst_PED_N09_02P (Child) | Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.23 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS
| Hea ph160 |  | 120 | 120 | 390 | 145 |  |  | Off | 3 | Off | CARE Dose4D | 22.18 | 14s | 25% | 60 |  | 0.5 |  | 2 | 0.6 |  | 3 | 128x0.6mm | [15 empty] | Off
|  | Hea ph150 | [22 empty] | 3 | 1.5 | 81 | J30f medium smooth | Head Soft |  |  | X-CARE | GEPACS | [9 empty] | manual
|  | Hea bon | [22 empty] | 3 | 1.5 | 81 | H60f sharp FR | Bone |  |  | X-CARE | GEPACS | MCHNeuroBodyAW | [8 empty] | manual
|  | Hea cor | [22 empty] | 3 | 1.5 | 134 | J30f medium smooth | Head Soft |  |  | X-CARE | GEPACS | [9 empty] | manual
|  | Hea sag | [22 empty] | 3 | 1.5 | 134 | J30f medium smooth | Head Soft |  |  | X-CARE | GEPACS | [9 empty] | manual
|  | Hea ph150 :-D | [22 empty] | 1 | 1 | 201 | J30f medium smooth | Head Soft |  |  | X-CARE | ROUTER | GEPACS | [8 empty] | manual
11360_Sin_fus_nav_wo_cst_PED (Child)
11360_Sin_fus_nav_wo_cst_PED (Child) | Topogram | Topogram |  | 100 |  | 109 |  |  |  | 3 | Off | CARE Dose | 0.18 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS
| Sin bon |  | 120 | 120 | 140 | 140 |  |  | Off | 3 | On | CARE Dose4D | 21.42 | 12s | 25% | 60 |  | 1 |  | 2 | 0.6 |  | 0.6 | 128x0.6mm | [15 empty] | Off
|  | Sin FUSION | [22 empty] | 0.6 | 0.6 | 167 | J30s medium smooth | Soft Tissue |  |  |  | OR_MIDIA_POOL | ROUTER | GEPACS | [7 empty] | manual
|  | Sin bon ph360 :-D | [22 empty] | 0.6 | 0.6 | 167 | J70h very sharp | Bone |  |  |  | ROUTER | GEPACS | [8 empty] | manual
|  | Sin cor bon | [22 empty] | 1 | 1 | 151 | J70h very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Sin sag bon | [22 empty] | 1 | 1 | 151 | J70h very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
11390_Tem_bon_sim_wo_cst_PED (Child)
11390_Tem_bon_sim_wo_cst_PED (Child) | Topogram | Topogram |  | 100 |  | 109 |  |  |  | 3 | Off | CARE Dose | 0.18 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS
| Tem bon |  | 120 | 120 | 180 | 250 |  |  | Off | 5 | Off | CARE Dose4D | 54.8 | 24s | 25% | 60 |  | 1 |  | 2 | 0.8 |  | 2 | 16x0.3mm | [15 empty] | Off
|  | Tem ph390 | [22 empty] | 2 | 2 | 21 | U40u medium | Inner Ear |  |  |  | GEPACS | [9 empty] | manual
|  | R Tem bon | [22 empty] | 0.5 | 0.3 | 134 | U70u sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | L Tem bon | [22 empty] | 0.5 | 0.3 | 134 | U70u sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Tem bon ph390 | [22 empty] | 1 | 1 | 41 | U70u sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | R Tem cor bon | [22 empty] | 0.5 | 0.3 | 267 | U70u sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | L Tem cor bon | [22 empty] | 0.5 | 0.3 | 267 | U70u sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Tem bon ph390 :-D | [22 empty] | 0.5 | 0.3 | 134 | U70u sharp | Bone |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
HeadRoutine (Child)
HeadRoutine (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.16 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head
| Head |  | 100 | 100 | 582 | 291 |  |  | On | 3 | On | CARE Dose4D | 56.22 | 11s | 25% |  |  | 1 |  | 2 | 0.8 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | Head 5.0 H31s | [22 empty] | 5 | 5 | 25 | H31s medium smooth + | Baby Brain | [13 empty] | manual
HeadRoutineSeq (Child)
HeadRoutineSeq (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.16 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head
| HeadSeq |  | 100 | 100 | 582 | 291 |  |  | On | 3 | On | CARE Dose4D | 62.48 |  | 25% |  |  | 1 |  | 2 | 0 | 34 | 5 | 128x0.6mm | [15 empty] | Off
|  | HeadSeq 5.0 H31s | [22 empty] | 5 |  | 27 | H31s medium smooth + | Baby Brain | [13 empty] | manual
Head_IR (Child)
Head_IR (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.16 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head
| Head |  | 100 | 100 | 407 | 204 |  |  | On | 3 | On | CARE Dose4D | 39.33 | 11s | 25% |  |  | 1 |  | 2 | 0.8 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | Head 5.0 J30s 3 | [22 empty] | 5 | 5 | 25 | J30s medium smooth | Baby Brain | [13 empty] | manual
Head_XCARE (Child)
Head_XCARE (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.16 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head
| Head |  | 100 | 100 | 582 | 291 |  |  | On | 3 | On | CARE Dose4D | 56.36 | 14s | 25% |  |  | 1 |  | 2 | 0.6 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | Head 5.0 H31s | [22 empty] | 5 | 5 | 25 | H31s medium smooth + | Baby Brain |  |  | X-CARE | [10 empty] | manual
Head_XCARE_IR (Child)
Head_XCARE_IR (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.16 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head
| Head |  | 100 | 100 | 407 | 204 |  |  | On | 3 | On | CARE Dose4D | 39.45 | 14s | 25% |  |  | 1 |  | 2 | 0.6 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | Head 5.0 J30s 3 | [22 empty] | 5 | 5 | 25 | J30s medium smooth | Baby Brain |  |  | X-CARE | [10 empty] | manual
InnerEar (Child)
InnerEar (Child) | Topogram | Topogram 0.6 T80f |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.16 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Head
| InnerEar |  | 120 | 120 | 180 | 90 |  |  | On | 5 | On | CARE Dose4D | 22.46 | 6s | 25% |  |  | 1 |  | 2 | 0.8 |  | 0.6 | 128x0.6mm | [15 empty] | Off
|  | InnerEar 0.6 H60s | [22 empty] | 0.6 | 0.6 | 67 | H60s sharp FR | Inner Ear | [13 empty] | manual
InnerEarUHR (Child)
InnerEarUHR (Child) | Topogram | Topogram 0.6 T80f |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.16 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Head
| InnerEarUHR |  | 120 | 120 | 180 | 90 |  |  | On | 5 | On | CARE Dose4D | 39.82 | 24s | 25% |  |  | 1 |  | 2 | 0.85 |  | 0.6 | 16x0.6mm | [15 empty] | Off
|  | InnerEarUHR 0.6 U75u | [22 empty] | 0.6 | 0.6 | 67 | U75u very sharp ASA | Inner Ear | [13 empty] | manual
InnerEarUHRSeq (Child)
InnerEarUHRSeq (Child) | Topogram | Topogram 0.6 T80f |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.16 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Head
| InnEarUHRSeq |  | 120 | 120 | 180 | 90 |  |  | On | 5 | On | CARE Dose4D | 41.65 |  | 25% |  |  | 1 |  | 2 | 0 | 4 | 0.6 | 16x0.6mm | [15 empty] | Off
|  | InnEarUHRSeq 0.6 U75u | [22 empty] | 0.6 |  | 68 | U75u very sharp ASA | Inner Ear | [13 empty] | manual
InnerEarZUHR (Child)
InnerEarZUHR (Child) | Topogram | Topogram 0.6 T80f |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.16 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Head
| InnerEarUHR |  | 120 | 120 | 230 | 115 |  |  | On | 5 | On | CARE Dose4D | 51.02 | 24s | 25% |  |  | 1 |  | 2 | 0.85 |  | 0.4 | 16x0.3mm | [15 empty] | Off
|  | InnerEarUHR 0.4 U75u | [22 empty] | 0.4 | 0.4 | 101 | U75u very sharp ASA | Inner Ear | [13 empty] | manual
InnerEarZUHR_IR (Child)
InnerEarZUHR_IR (Child) | Topogram | Topogram 0.6 T80f |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.16 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Head
| InnerEarUHR |  | 120 | 120 | 161 | 81 |  |  | On | 5 | On | CARE Dose4D | 35.66 | 24s | 25% |  |  | 1 |  | 2 | 0.85 |  | 0.4 | 16x0.3mm | [15 empty] | Off
|  | InnerEarUHR 0.4 V80u 3 | [22 empty] | 0.4 | 0.4 | 101 | V80u very sharp | Inner Ear | [13 empty] | manual
NeonateHead (Child)
NeonateHead (Child) | Topogram | Topogram 0.6 T20f |  | 80 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.08 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head
| NeonateHead |  | 100 | 100 | 582 | 150 |  |  | On | 3 | On | CARE Dose4D | 56.22 | 11s | 25% |  |  | 1 |  | 2 | 0.8 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | NeonateHead 5.0 C30s | [22 empty] | 5 | 5 | 25 | C30s med. smooth FR | Baby Brain | [13 empty] | manual
Orbit (Child)
Orbit (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.16 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head
| Orbit |  | 120 | 120 | 125 | 63 |  |  | On | 5 | On | CARE Dose4D | 15.6 | 10s | 25% |  |  | 1 |  | 2 | 0.8 |  | 3 | 128x0.6mm | [15 empty] | Off
|  | Orbit 3.0 H60s | [22 empty] | 3 | 3 | 14 | H60s sharp FR | Base Orbita | [13 empty] | manual
|  | Orbit 3.0 H30s | [22 empty] | 3 | 3 | 14 | H30s medium smooth | Base Orbita | [13 empty] | manual
|  | Orbit 0.6 H60s | [22 empty] | 0.6 | 0.4 | 101 | H60s sharp FR | Base Orbita | [13 empty] | manual
Sinus (Child)
Sinus (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.16 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head
| Sinus |  | 120 | 120 | 50 | 25 |  |  | On | 5 | On | CARE Dose4D | 6.24 | 10s | 25% |  |  | 1 |  | 2 | 0.8 |  | 3 | 128x0.6mm | [15 empty] | Off
|  | Sinus 3.0 H60s | [22 empty] | 3 | 3 | 21 | H60s sharp FR | Sinuses | [13 empty] | manual
|  | Sinus 3.0 H30s | [22 empty] | 3 | 3 | 21 | H30s medium smooth | Sinuses | [13 empty] | manual
|  | Sinus 1.0 H60s | [22 empty] | 1 | 0.7 | 86 | H60s sharp FR | Sinuses | [13 empty] | manual
|  | Sinus 1.0 H30s | [22 empty] | 1 | 0.7 | 86 | H30s medium smooth | Sinuses | [13 empty] | manual
Neck
NeckRoutine (Child)
NeckRoutine (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.07 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Neck |  | 100 | 100 | 233 | 117 |  |  | On | 7 | On | CARE Dose4D | 7.95 | 9s | 25% |  |  | 1 |  | 2 | 0.8 |  | 5 | 32x1.2mm | [15 empty] | Off
|  | Neck 5.0 B30s | [22 empty] | 5 | 5 | 35 | B30s medium smooth | Baby Neck | [13 empty] | manual
|  | Neck 1.5 B60s | [22 empty] | 1.5 | 1 | 171 | B60s sharp | Baby Neck | [13 empty] | manual
Neck_IR (Child)
Neck_IR (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.07 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Neck |  | 100 | 100 | 163 | 82 |  |  | On | 7 | On | CARE Dose4D | 5.57 | 9s | 25% |  |  | 1 |  | 2 | 0.8 |  | 5 | 32x1.2mm | [15 empty] | Off
|  | Neck 5.0 I30s 3 | [22 empty] | 5 | 5 | 35 | I30s medium smooth | Baby Neck | [13 empty] | manual
|  | Neck 1.5 I50s 3 | [22 empty] | 1.5 | 1 | 171 | I50s medium sharp ASA | Baby Neck | [13 empty] | manual
Shoulder
ShoulderRoutine (Child)
ShoulderRoutine (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.07 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Shoulder |  | 100 | 100 | 254 | 127 |  |  | On | 5 | On | CARE Dose4D | 8.44 | 7s | 25% |  |  | 1 |  | 2 | 0.8 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | Shoulder 5.0 B60s | [22 empty] | 5 | 5 | 27 | B60s sharp | Bone | [13 empty] | manual
|  | Shoulder 0.75 B31s | [22 empty] | 0.75 | 0.5 | 261 | B31s medium smooth + | Shoulder | [13 empty] | manual
|  | Shoulder 0.75 B60s | [22 empty] | 0.75 | 0.5 | 261 | B60s sharp | Bone | [13 empty] | manual
Thorax
15130_Che_1_wo_cst_PED_T06_00P (Child)
15130_Che_1_wo_cst_PED_T06_00P (Child) | Topogram | Topogram |  | 100 |  | 109 |  |  |  | 3 | Off | CARE Dose | 0.08 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS
| Che |  | 120 | 100 | 150 | 90 |  |  | Off | 7 | On | CARE Dose4D | 4.97 | 20s | 25% | 30 |  | 0.33 |  | 5 | 1.3 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | Che pc130 | [22 empty] | 3 | 2 | 101 | I41f medium | Chest |  |  |  | GEPACS | [9 empty] | manual
|  | Che sag | [22 empty] | 2 | 2 | 126 | I41f medium | Chest |  |  |  | GEPACS | [9 empty] | manual
|  | Che cor mip | [22 empty] | 10 | 10 | 26 | I41f medium | Chest MIP |  |  |  | GEPACS | [9 empty] | manual
|  | Che pc130 :-D | [22 empty] | 0.6 | 0.6 | 334 | I41f medium | Chest |  |  |  | ROUTER | GEPACS | [8 empty] | manual
15137_Che_1_AbdPel_wo_cst_PED (Child)
15137_Che_1_AbdPel_wo_cst_PED (Child) | Topogram | Topogram |  | 100 |  | 212 |  |  |  | 3 | Off | CARE Dose | 0.08 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS
| Che |  | 120 | 100 | 150 | 90 |  |  | Off | 7 | On | CARE Dose4D | 4.97 | 20s | 25% | 30 |  | 0.33 |  | 5 | 1.3 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | Che pc130 | [22 empty] | 3 | 2 | 101 | I41f medium | Chest |  |  |  | GEPACS | [9 empty] | manual
|  | Che sag | [22 empty] | 2 | 2 | 126 | I41f medium | Chest |  |  |  | GEPACS | [9 empty] | manual
|  | Che cor mip | [22 empty] | 10 | 10 | 26 | I41f medium | Chest MIP |  |  |  | GEPACS | [9 empty] | manual
|  | Che pc130 :-D | [22 empty] | 0.6 | 0.6 | 334 | I41f medium | Chest |  |  |  | ROUTER | GEPACS | [8 empty] | manual
| AbdPel |  | 120 | 100 | 180 | 90 |  |  | Off | 7 | On | CARE Dose4D | 5.96 | 20s | 25% | 30 |  | 0.33 |  | 2 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | AbdPel pa200 | [22 empty] | 3 | 3 | 67 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | AbdPel cor | [22 empty] | 3 | 3 | 61 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | AbdPel sag | [22 empty] | 3 | 3 | 84 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | AbdPel pa200 :-D | [22 empty] | 0.6 | 0.6 | 334 | I30f medium smooth | Abdomen |  |  |  | ROUTER | GEPACS | [8 empty] | manual
15140_Che_2_PED (Child)
15140_Che_2_PED (Child) | Topogram | Topogram |  | 100 |  | 109 |  |  |  | 3 | Off | CARE Dose | 0.08 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS
| art Che |  | 120 | 100 | 150 | 90 |  |  | Off | 7 | On | CARE Dose4D | 4.97 | 20s | 25% | 30 |  | 0.33 |  | 25 | 1.3 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | art Che pc140 | [22 empty] | 3 | 2 | 101 | I41f medium | Chest |  |  |  | GEPACS | [9 empty] | manual
|  | art Che sag | [22 empty] | 2 | 2 | 126 | I41f medium | Chest |  |  |  | GEPACS | [9 empty] | manual
|  | art Che cor mip | [22 empty] | 10 | 10 | 26 | I41f medium | Chest MIP |  |  |  | GEPACS | [9 empty] | manual
|  | art Che pc140 :-D | [22 empty] | 0.6 | 0.6 | 334 | I41f medium | Chest |  |  |  | ROUTER | GEPACS | [8 empty] | manual
15170_Che_2_AbdPel_PED (Child)
15170_Che_2_AbdPel_PED (Child) | Topogram | Topogram |  | 100 |  | 212 |  |  |  | 3 | Off | CARE Dose | 0.08 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS
| art Che |  | 120 | 100 | 150 | 90 |  |  | Off | 7 | On | CARE Dose4D | 4.97 | 20s | 25% | 30 |  | 0.33 |  | 25 | 1.3 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | art Che pc140 | [22 empty] | 3 | 2 | 101 | I41f medium | Chest |  |  |  | GEPACS | [9 empty] | manual
|  | art Che sag | [22 empty] | 2 | 2 | 126 | I41f medium | Chest |  |  |  | GEPACS | [9 empty] | manual
|  | art Che cor mip | [22 empty] | 10 | 10 | 26 | I41f medium | Chest MIP |  |  |  | GEPACS | [9 empty] | manual
|  | art Che pc140 :-D | [22 empty] | 0.6 | 0.6 | 334 | I41f medium | Chest |  |  |  | ROUTER | GEPACS | [8 empty] | manual
| ven AbdPel |  | 120 | 100 | 180 | 90 |  |  | Off | 7 | On | CARE Dose4D | 5.96 | 20s | 25% | 30 |  | 0.33 |  | 60 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven AbdPel pa030 | [22 empty] | 3 | 3 | 67 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel cor | [22 empty] | 3 | 3 | 61 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel sag | [22 empty] | 3 | 3 | 84 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel pa030 :-D | [22 empty] | 0.6 | 0.6 | 334 | I30f medium smooth | Abdomen |  |  |  | ROUTER | GEPACS | [8 empty] | manual
15211_Che_3_PED (Child)
15211_Che_3_PED (Child) | Topogram | Topogram |  | 100 |  | 109 |  |  |  | 3 | Off | CARE Dose | 0.08 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS
| ven Che |  | 120 | 100 | 150 | 90 |  |  | Off | 7 | On | CARE Dose4D | 4.97 | 20s | 25% | 30 |  | 0.33 |  | 75 | 1.3 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven Che pc211 | [22 empty] | 3 | 2 | 101 | I41f medium | Chest |  |  |  | GEPACS | [9 empty] | manual
|  | ven Che sag | [22 empty] | 2 | 2 | 126 | I41f medium | Chest |  |  |  | GEPACS | [9 empty] | manual
|  | ven Che cor mip | [22 empty] | 10 | 10 | 26 | I41f medium | Chest MIP |  |  |  | GEPACS | [9 empty] | manual
|  | ven Che pc211 :-D | [22 empty] | 0.6 | 0.6 | 334 | I41f medium | Chest |  |  |  | ROUTER | GEPACS | [8 empty] | manual
15215_Che_3_AbdPel_PED (Child)
15215_Che_3_AbdPel_PED (Child) | Topogram | Topogram |  | 100 |  | 212 |  |  |  | 3 | Off | CARE Dose | 0.08 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS
| ven Che |  | 120 | 100 | 150 | 90 |  |  | Off | 7 | On | CARE Dose4D | 4.97 | 20s | 25% | 30 |  | 0.33 |  | 75 | 1.3 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven Che pc211 | [22 empty] | 3 | 2 | 101 | I41f medium | Chest |  |  |  | GEPACS | [9 empty] | manual
|  | ven Che sag | [22 empty] | 2 | 2 | 126 | I41f medium | Chest |  |  |  | GEPACS | [9 empty] | manual
|  | ven Che cor mip | [22 empty] | 10 | 10 | 26 | I41f medium | Chest MIP |  |  |  | GEPACS | [9 empty] | manual
|  | ven Che pc211 :-D | [22 empty] | 0.6 | 0.6 | 334 | I41f medium | Chest |  |  |  | ROUTER | GEPACS | [8 empty] | manual
| ven AbdPel |  | 120 | 100 | 180 | 90 |  |  | Off | 7 | On | CARE Dose4D | 5.96 | 20s | 25% | 30 |  | 0.33 |  | 10 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven AbdPel pa030 | [22 empty] | 3 | 3 | 67 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel cor | [22 empty] | 3 | 3 | 61 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel sag | [22 empty] | 3 | 3 | 84 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel pa030 :-D | [22 empty] | 0.6 | 0.6 | 334 | I30f medium smooth | Abdomen |  |  |  | ROUTER | GEPACS | [8 empty] | manual
15335_Che_pec_wo_cst_PED (Child)
15335_Che_pec_wo_cst_PED (Child) | Topogram | Topogram |  | 100 |  | 109 |  |  |  | 3 | Off | CARE Dose | 0.08 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS
| Che pec |  | 120 | 100 | 180 | 90 |  |  | Off | 7 | On | CARE Dose4D | 5.96 | 20s | 25% | 30 |  | 0.33 |  | 5 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | Che pec pc335 | [22 empty] | 3 | 2 | 101 | I41f medium | Chest |  |  |  | GEPACS | [9 empty] | manual
|  | Che sag | [22 empty] | 2 | 2 | 126 | I41f medium | Chest |  |  |  | GEPACS | [9 empty] | manual
|  | Che cor mip | [22 empty] | 10 | 10 | 26 | I41f medium | Chest MIP |  |  |  | GEPACS | [9 empty] | manual
|  | Che pec pc335 :-D | [22 empty] | 0.6 | 0.6 | 334 | I41f medium | Chest |  |  |  | ROUTER | GEPACS | [8 empty] | manual
Flash_Thorax (Child)
Flash_Thorax (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.07 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Fl_Thorax |  | 120 | 120 | 112 |  | 56 |  | On | 7 | On | CARE Dose4D | 6.2 | 5s | 25% |  |  | 0.285 |  | 2 | 3 |  | 3 | 128x0.6mm
|  | Fl_Thorax 3.0 B30f | [22 empty] | 3 | 3 | 51 | B30f medium smooth | Mediastinum
|  | Fl_Thorax 3.0 B60f | [22 empty] | 3 | 3 | 51 | B60f sharp | Baby Lung
|  | Fl_Thorax 1.5 B30f | [22 empty] | 1.5 | 1 | 151 | B30f medium smooth | Mediastinum
|  | Fl_Thorax 1.5 B60f | [22 empty] | 1.5 | 1 | 151 | B60f sharp | Baby Lung
Flash_Thorax_IR (Child)
Flash_Thorax_IR (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.07 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Fl_Thorax |  | 120 | 120 | 68 |  | 34 |  | On | 7 | On | CARE Dose4D | 3.76 | 5s | 25% |  |  | 0.285 |  | 2 | 3 |  | 3 | 128x0.6mm
|  | Fl_Thorax 3.0 I30f 3 | [22 empty] | 3 | 3 | 51 | I30f medium smooth | Mediastinum
|  | Fl_Thorax 3.0 I50f 3 | [22 empty] | 3 | 3 | 51 | I50f medium sharp ASA | Baby Lung
|  | Fl_Thorax 1.5 I30f 3 | [22 empty] | 1.5 | 1 | 151 | I30f medium smooth | Mediastinum
|  | Fl_Thorax 1.5 I50f 3 | [22 empty] | 1.5 | 1 | 151 | I50f medium sharp ASA | Baby Lung
ThoraxRoutine (Child)
ThoraxRoutine (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.07 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Thorax |  | 100 | 100 | 155 | 78 |  |  | On | 7 | On | CARE Dose4D | 5.28 | 5s | 25% |  |  | 0.5 |  | 2 | 1.4 |  | 3 | 32x1.2mm | [15 empty] | Off
|  | Thorax 3.0 B30f | [22 empty] | 3 | 3 | 51 | B30f medium smooth | Mediastinum | [13 empty] | manual
|  | Thorax 3.0 B60f | [22 empty] | 3 | 3 | 51 | B60f sharp | Baby Lung | [13 empty] | manual
|  | Thorax 1.5 B30f | [22 empty] | 1.5 | 1 | 151 | B30f medium smooth | Mediastinum | [13 empty] | manual
|  | Thorax 1.5 B60f | [22 empty] | 1.5 | 1 | 151 | B60f sharp | Baby Lung | [13 empty] | manual
ThoraxSeqHR (Child)
ThoraxSeqHR (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.07 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| ThorHRSeq |  | 120 | 120 | 110 | 55 |  |  | On | 3 | On | CARE Dose4D | 1.24 |  | 25% |  |  | 0.5 |  | 2 | 0 | 10 | 1 | 2x1.0mm | [15 empty] | Off
|  | ThorHRSeq 1.0 B70s | [22 empty] | 1 |  | 32 | B70s very sharp | Baby Lung | [13 empty] | manual
Thorax_IR (Child)
Thorax_IR (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.07 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Thorax |  | 100 | 100 | 93 | 47 |  |  | On | 7 | On | CARE Dose4D | 3.17 | 5s | 25% |  |  | 0.5 |  | 2 | 1.4 |  | 3 | 32x1.2mm | [15 empty] | Off
|  | Thorax 3.0 I30f 3 | [22 empty] | 3 | 3 | 51 | I30f medium smooth | Mediastinum | [13 empty] | manual
|  | Thorax 3.0 I50f 3 | [22 empty] | 3 | 3 | 51 | I50f medium sharp ASA | Baby Lung | [13 empty] | manual
|  | Thorax 1.5 I30f 3 | [22 empty] | 1.5 | 1 | 151 | I30f medium smooth | Mediastinum | [13 empty] | manual
|  | Thorax 1.5 I50f 3 | [22 empty] | 1.5 | 1 | 151 | I50f medium sharp ASA | Baby Lung | [13 empty] | manual
Thorax_XCARE (Child)
Thorax_XCARE (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.07 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Thorax |  | 100 | 100 | 155 | 78 |  |  | On | 7 | On | CARE Dose4D | 5.29 | 5s | 25% |  |  | 0.285 |  | 2 | 0.6 |  | 3 | 32x1.2mm | [15 empty] | Off
|  | Thorax 3.0 B30f | [22 empty] | 3 | 3 | 51 | B30f medium smooth | Mediastinum |  |  | X-CARE | [10 empty] | manual
|  | Thorax 3.0 B60f | [22 empty] | 3 | 3 | 51 | B60f sharp | Baby Lung |  |  | X-CARE | [10 empty] | manual
|  | Thorax 1.5 B30f | [22 empty] | 1.5 | 1 | 151 | B30f medium smooth | Mediastinum |  |  | X-CARE | [10 empty] | manual
|  | Thorax 1.5 B60f | [22 empty] | 1.5 | 1 | 151 | B60f sharp | Baby Lung |  |  | X-CARE | [10 empty] | manual
Thorax_XCARE_IR (Child)
Thorax_XCARE_IR (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.07 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Thorax |  | 100 | 100 | 93 | 47 |  |  | On | 7 | On | CARE Dose4D | 3.17 | 5s | 25% |  |  | 0.285 |  | 2 | 0.6 |  | 3 | 32x1.2mm | [15 empty] | Off
|  | Thorax 3.0 I30f 3 | [22 empty] | 3 | 3 | 51 | I30f medium smooth | Mediastinum |  |  | X-CARE | [10 empty] | manual
|  | Thorax 3.0 I50f 3 | [22 empty] | 3 | 3 | 51 | I50f medium sharp ASA | Baby Lung |  |  | X-CARE | [10 empty] | manual
|  | Thorax 1.5 I30f 3 | [22 empty] | 1.5 | 1 | 151 | I30f medium smooth | Mediastinum |  |  | X-CARE | [10 empty] | manual
|  | Thorax 1.5 I50f 3 | [22 empty] | 1.5 | 1 | 151 | I50f medium sharp ASA | Baby Lung |  |  | X-CARE | [10 empty] | manual
Abdomen
16030_AbdPel_PED_A02_04P (Child)
16030_AbdPel_PED_A02_04P (Child) | Topogram | Topogram |  | 100 |  | 212 |  |  |  | 3 | Off | CARE Dose | 0.08 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS
| ven AbdPel |  | 120 | 100 | 180 | 90 |  |  | Off | 7 | On | CARE Dose4D | 5.96 | 20s | 25% | 30 |  | 0.33 |  | 70 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven AbdPel pa030 | [22 empty] | 3 | 3 | 67 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel cor | [22 empty] | 3 | 3 | 61 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel sag | [22 empty] | 3 | 3 | 84 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel pa030 :-D | [22 empty] | 0.6 | 0.6 | 334 | I30f medium smooth | Abdomen |  |  |  | ROUTER | GEPACS | [8 empty] | manual
16080_AbdPel_liv_bi_PED (Child)
16080_AbdPel_liv_bi_PED (Child) | Topogram | Topogram |  | 100 |  | 212 |  |  |  | 3 | Off | CARE Dose | 0.08 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS
| PreMonitoring |  | 120 | 120 | 240 | 20 |  |  | Off | 8 | Off | CARE Dose4D | 1.23 |  | 25% |  |  | 0.5 |  | 2 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | PreMonitoring 10.0 B30s | [22 empty] | 10 |  | 1 | B30s medium smooth | 400/40 Abd | [13 empty] | manual
| Monitoring |  | 120 | 120 | 240 | 20 |  |  | Off | 8 | Off | CARE Dose4D | 36.98 |  | 25% |  |  | 0.5 |  | 10 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | Monitoring 10.0 B30s | [22 empty] | 10 |  | 30 | B30s medium smooth | 400/40 Abd | [13 empty] | manual
| art Abd |  | 100 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 8.37 | 30s | 25% | 30 |  | 0.5 |  | 16 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | art Abd pa080 | [22 empty] | 3 | 3 | 61 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd pa080 :-D | [22 empty] | 0.6 | 0.6 | 301 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
| ven AbdPel |  | 100 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 8.37 | 30s | 25% | 30 |  | 0.5 |  | 60 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven AbdPel pa080 | [22 empty] | 3 | 3 | 45 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel cor | [22 empty] | 3 | 3 | 84 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel sag | [22 empty] | 3 | 3 | 121 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel pa080 :-D | [22 empty] | 0.6 | 0.6 | 223 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
16110_Abd_liv_tri_Pel_PED (Child)
16110_Abd_liv_tri_Pel_PED (Child) | Topogram | Topogram |  | 100 |  | 212 |  |  |  | 3 | Off | CARE Dose | 0.08 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS
| PreMonitoring |  | 120 | 120 | 240 | 20 |  |  | Off | 8 | Off | CARE Dose4D | 1.23 |  | 25% |  |  | 0.5 |  | 2 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | PreMonitoring 10.0 B30s | [22 empty] | 10 |  | 1 | B30s medium smooth | 400/40 Abd | [13 empty] | manual
| Monitoring |  | 120 | 120 | 240 | 20 |  |  | Off | 8 | Off | CARE Dose4D | 36.98 |  | 25% |  |  | 0.5 |  | 10 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | Monitoring 10.0 B30s | [22 empty] | 10 |  | 30 | B30s medium smooth | 400/40 Abd | [13 empty] | manual
| art Abd |  | 100 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 8.37 | 30s | 25% | 30 |  | 0.5 |  | 16 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | art Abd pa110 | [22 empty] | 3 | 3 | 117 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd cor | [22 empty] | 3 | 3 | 128 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd sag | [22 empty] | 3 | 3 | 121 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd pa110 :-D | [22 empty] | 0.6 | 0.6 | 584 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
| ven AbdPel |  | 100 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 8.37 | 30s | 25% | 30 |  | 0.5 |  | 60 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven AbdPel pa110 | [22 empty] | 3 | 3 | 117 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel cor | [22 empty] | 3 | 3 | 135 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel sag | [22 empty] | 3 | 3 | 121 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel pa110 :-D | [22 empty] | 0.6 | 0.6 | 584 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
| 5m Abd |  | 100 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 8.37 | 30s | 25% | 30 |  | 0.5 |  | 4 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | 5m Abd pa110 | [22 empty] | 3 | 3 | 117 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | 5m Abd pa110 :-D | [22 empty] | 0.6 | 0.6 | 584 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
16200_AbdPel_wo_cst_PED (Child)
16200_AbdPel_wo_cst_PED (Child) | Topogram | Topogram |  | 100 |  | 212 |  |  |  | 3 | Off | CARE Dose | 0.08 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS
| AbdPel |  | 120 | 100 | 180 | 90 |  |  | Off | 7 | On | CARE Dose4D | 5.96 | 20s | 25% | 30 |  | 0.33 |  | 5 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | AbdPel pa200 | [22 empty] | 3 | 3 | 67 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | AbdPel cor | [22 empty] | 3 | 3 | 61 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | AbdPel sag | [22 empty] | 3 | 3 | 84 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | AbdPel pa200 :-D | [22 empty] | 0.6 | 0.6 | 334 | I30f medium smooth | Abdomen |  |  |  | ROUTER | GEPACS | [8 empty] | manual
16407_ang_AbdPel_mes_art_PED (Child)
16407_ang_AbdPel_mes_art_PED (Child) | Topogram | Topogram |  | 100 |  | 212 |  |  |  | 3 | Off | CARE Dose | 0.08 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS
| PreMonitoring |  | 120 | 120 | 240 | 50 |  |  | Off | 8 | Off | CARE Dose4D | 3.08 |  | 25% |  |  | 0.5 |  | 2 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | PreMonitoring | [22 empty] | 10 |  | 1 | B40s medium | 400/40 Abd | [13 empty] | manual
| Monitoring |  | 120 | 120 | 240 | 50 |  |  | Off | 8 | Off | CARE Dose4D | 123.28 |  | 25% |  |  | 0.5 |  | 10 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | Monitoring | [22 empty] | 10 |  | 40 | B40s medium | 400/40 Abd | [13 empty] | manual
| art AbdPel |  | 100 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 8.37 | 30s | 25% | 30 |  | 0.5 |  | 5 | 1 |  | 2 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | art AbdPel pa407 | [22 empty] | 2 | 2 | 235 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art AbdPel cor | [22 empty] | 2 | 2 | 104 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art AbdPel sag | [22 empty] | 2 | 2 | 191 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art AbdPel cor mip | [22 empty] | 24 | 5 | 77 | I30f medium smooth | Abdomen MIP |  |  |  | GEPACS | [9 empty] | manual
|  | art AbdPel pa407 :-D | [22 empty] | 0.6 | 0.6 | 781 | I26f medium smooth ASA | Abdomen MIP |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
| ven AbdPel |  | 100 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 8.37 | 30s | 25% | 30 |  | 0.5 |  | 65 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven AbdPel pa407 | [22 empty] | 3 | 3 | 157 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel cor | [22 empty] | 3 | 3 | 70 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel pa407 :-D | [22 empty] | 0.6 | 0.6 | 781 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
AbdomenRoutine (Child)
AbdomenRoutine (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.07 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Abdomen |  | 100 | 100 | 296 | 148 |  |  | On | 7 | On | CARE Dose4D | 7.9 | 30s | 25% |  |  | 0.5 |  | 2 | 1.4 |  | 5 | 32x1.2mm | [15 empty] | Off
|  | Abdomen 5.0 B30f | [22 empty] | 5 | 5 | 41 | B30f medium smooth | Baby Abdomen | [13 empty] | manual
Abdomen_IR (Child)
Abdomen_IR (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.07 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Abdomen |  | 100 | 100 | 207 | 104 |  |  | On | 7 | On | CARE Dose4D | 7.05 | 30s | 25% |  |  | 0.5 |  | 2 | 1.4 |  | 5 | 32x1.2mm | [15 empty] | Off
|  | Abdomen 5.0 I30f 3 | [22 empty] | 5 | 5 | 41 | I30f medium smooth | Baby Abdomen | [13 empty] | manual
Flash_Abdomen (Child)
Flash_Abdomen (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.07 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Fl_Abdomen |  | 120 | 120 | 212 |  | 106 |  | On | 7 | On | CARE Dose4D | 8.4 | 30s | 25% |  |  | 0.285 |  | 2 | 3 |  | 5 | 128x0.6mm
|  | Fl_Abdomen 5.0 B30f | [22 empty] | 5 | 5 | 41 | B30f medium smooth | Baby Abdomen
Flash_Abdomen_IR (Child)
Flash_Abdomen_IR (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.07 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Fl_Abdomen |  | 120 | 120 | 148 |  | 74 |  | On | 7 | On | CARE Dose4D | 8.18 | 30s | 25% |  |  | 0.285 |  | 2 | 3 |  | 5 | 128x0.6mm
|  | Fl_Abdomen 5.0 I30f 3 | [22 empty] | 5 | 5 | 41 | I30f medium smooth | Baby Abdomen
Spine
13040_Csp_wo_cst_PED (Child)
13040_Csp_wo_cst_PED (Child) | Topogram | Topogram 0.6 T80f |  | 100 |  | 109 |  |  |  | 3 | Off | CARE Dose | 0.08 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Body |  |  |  | GEPACS
| Csp |  | 120 | 120 | 180 | 85 |  |  | Off | 3 | On | CARE Dose4D | 9.97 | 30s | 25% | 30 |  | 0.5 |  | 2 | 0.9 |  | 2 | 128x0.6mm | [15 empty] | Off
|  | Csp bon ps040 | [22 empty] | 2 | 2 | 81 | B60f sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Csp | [22 empty] | 2 | 2 | 81 | B31f medium smooth + | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | Csp cor bon | [22 empty] | 2 | 1.8 | 84 | B60f sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Csp sag bon | [22 empty] | 2 | 1.8 | 84 | B60f sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Csp sag | [22 empty] | 2 | 2 | 76 | B31f medium smooth + | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | Csp bon ps040 :-D | [22 empty] | 0.75 | 0.7 | 229 | B60f sharp | Bone |  |  |  | ROUTER | GEPACS | [8 empty] | manual
13070_Lsp_wo_cst_PED (Child)
13070_Lsp_wo_cst_PED (Child) | Topogram | Topogram |  | 100 |  | 109 |  |  |  | 3 | Off | CARE Dose | 0.08 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Body |  |  |  | GEPACS
| Lsp |  | 120 | 120 | 180 | 85 |  |  | Off | 3 | On | CARE Dose4D | 9.97 | 30s | 25% | 30 |  | 0.5 |  | 2 | 0.9 |  | 2 | 128x0.6mm | [15 empty] | Off
|  | Lsp bon ps070 | [22 empty] | 2 | 2 | 81 | B60f sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Lsp | [22 empty] | 2 | 2 | 81 | B31f medium smooth + | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | Lsp cor bon | [22 empty] | 2 | 1.8 | 84 | B60f sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Lsp sag bon | [22 empty] | 2 | 1.8 | 84 | B60f sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Lsp sag | [22 empty] | 2 | 2 | 76 | B31f medium smooth + | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | Lsp bon ps070 :-D | [22 empty] | 0.75 | 0.7 | 229 | B60f sharp | Bone |  |  |  | ROUTER | GEPACS | [8 empty] | manual
13100_Tsp_wo_cst_PED (Child)
13100_Tsp_wo_cst_PED (Child) | Topogram | Topogram |  | 100 |  | 109 |  |  |  | 3 | Off | CARE Dose | 0.08 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Body |  |  |  | GEPACS
| Tsp |  | 120 | 120 | 180 | 85 |  |  | Off | 3 | On | CARE Dose4D | 9.97 | 30s | 25% | 30 |  | 0.5 |  | 2 | 0.9 |  | 2 | 128x0.6mm | [15 empty] | Off
|  | Tsp bon ps100 | [22 empty] | 2 | 2 | 81 | B60f sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Tsp | [22 empty] | 2 | 2 | 81 | B31f medium smooth + | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | Tsp cor bon | [22 empty] | 2 | 1.8 | 84 | B60f sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Tsp sag bon | [22 empty] | 2 | 1.8 | 84 | B60f sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Tsp sag | [22 empty] | 2 | 2 | 76 | B31f medium smooth + | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | Tsp bon ps070 :-D | [22 empty] | 0.75 | 0.7 | 229 | B60f sharp | Bone |  |  |  | ROUTER | GEPACS | [8 empty] | manual
SpineRoutine (Child)
SpineRoutine (Child) | Topogram | Topogram 0.6 T80f |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.07 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Body
| Spine |  | 120 | 120 | 330 | 165 |  |  | On | 3 | On | CARE Dose4D | 18.26 | 40s | 25% |  |  | 1 |  | 2 | 0.8 |  | 2 | 128x0.6mm | [15 empty] | Off
|  | Spine 2.0 B30s | [22 empty] | 2 | 2 | 81 | B30s medium smooth | Baby Spine | [13 empty] | manual
|  | Spine 0.6 B60s | [22 empty] | 0.6 | 0.4 | 401 | B60s sharp | Baby Spine | [13 empty] | manual
SpineRoutine_Vol (Child)
SpineRoutine_Vol (Child) | Topogram | Topogram 0.6 T80f |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.07 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Body
| Spine |  | 120 | 120 | 330 | 165 |  |  | On | 3 | On | CARE Dose4D | 18.26 | 40s | 25% |  |  | 1 |  | 2 | 0.8 |  | 2 | 128x0.6mm | [15 empty] | Off
|  | Spine 2.0 B30s | [22 empty] | 2 | 2 | 81 | B30s medium smooth | Baby Spine | [13 empty] | manual
|  | Spine 0.6 MPR spine multi | [22 empty] | 0.6 | 0.4 | 0 | B30s medium smooth | Baby Spine | [13 empty] | manual
|  | Spine 0.6 MPR spine multi | [22 empty] | 0.6 | 0.4 | 0 | B60s sharp | Baby Spine | [13 empty] | manual
Spine_IR (Child)
Spine_IR (Child) | Topogram | Topogram 0.6 T80f |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.07 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Body
| Spine |  | 120 | 120 | 231 | 116 |  |  | On | 3 | On | CARE Dose4D | 12.79 | 40s | 25% |  |  | 1 |  | 2 | 0.8 |  | 2 | 128x0.6mm | [15 empty] | Off
|  | Spine 2.0 I30s 3 | [22 empty] | 2 | 2 | 81 | I30s medium smooth | Baby Spine | [13 empty] | manual
|  | Spine 0.6 I50s 3 | [22 empty] | 0.6 | 0.4 | 401 | I50s medium sharp ASA | Baby Spine | [13 empty] | manual
Spine_IR_Vol (Child)
Spine_IR_Vol (Child) | Topogram | Topogram 0.6 T80f |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.07 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Body
| Spine |  | 120 | 120 | 231 | 116 |  |  | On | 3 | On | CARE Dose4D | 12.79 | 40s | 25% |  |  | 1 |  | 2 | 0.8 |  | 2 | 128x0.6mm | [15 empty] | Off
|  | Spine 2.0 I30s 3 | [22 empty] | 2 | 2 | 81 | I30s medium smooth | Baby Spine | [13 empty] | manual
|  | Spine 0.6 MPR spine multi | [22 empty] | 0.6 | 0.4 | 0 | I30s medium smooth | Baby Spine | [13 empty] | manual
|  | Spine 0.6 MPR spine multi | [22 empty] | 0.6 | 0.4 | 0 | I50s medium sharp ASA | Baby Spine | [13 empty] | manual
Pelvis
PelvisRoutine (Child)
PelvisRoutine (Child) | Topogram | Topogram 0.6 T80f |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.07 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Body
| Pelvis |  | 100 | 100 | 296 | 148 |  |  | On | 7 | On | CARE Dose4D | 10.08 | 30s | 25% |  |  | 0.5 |  | 2 | 0.8 |  | 5 | 32x1.2mm | [15 empty] | Off
|  | Pelvis 5.0 B31f | [22 empty] | 5 | 5 | 31 | B31f medium smooth + | Pelvis | [13 empty] | manual
Pelvis_IR (Child)
Pelvis_IR (Child) | Topogram | Topogram 0.6 T80f |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.07 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Body
| Pelvis |  | 100 | 100 | 207 | 104 |  |  | On | 7 | On | CARE Dose4D | 7.06 | 30s | 25% |  |  | 0.5 |  | 2 | 0.8 |  | 5 | 32x1.2mm | [15 empty] | Off
|  | Pelvis 5.0 I31f 3 | [22 empty] | 5 | 5 | 31 | I31f medium smooth | Pelvis | [13 empty] | manual
Specials
Flash_NeonateBody (Child)
Flash_NeonateBody (Child) | Topogram | Topogram 0.6 T20f |  | 80 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.03 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Fl_Neonate |  | 80 | 80 | 276 |  | 94 |  | On | 7 | On | CARE Dose4D | 1.94 | 5s | 25% |  |  | 0.285 |  | 2 | 3 |  | 5 | 128x0.6mm
|  | Fl_Neonate 5.0 B30f | [22 empty] | 5 | 5 | 31 | B30f medium smooth | Baby Abdomen
Flash_NeonateBody_70kV (Child)
Flash_NeonateBody_70kV (Child) | Topogram | Topogram 0.6 T20f |  | 70 |  | 164 |  |  |  | 3 | Off | CARE Dose | 0.03 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Fl_Neonate |  | 70 | 70 | 508 |  | 172 |  | On | 7 | On | CARE Dose4D | 1.79 | 5s | 25% |  |  | 0.285 |  | 2 | 1.55 |  | 5 | 128x0.6mm
|  | Fl_Neonate 5.0 B30f | [22 empty] | 5 | 5 | 31 | B30f medium smooth | Baby Abdomen
Flash_NeonateBody_IR (Child)
Flash_NeonateBody_IR (Child) | Topogram | Topogram 0.6 T20f |  | 80 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.03 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Fl_Neonate |  | 80 | 80 | 166 |  | 56 |  | On | 7 | On | CARE Dose4D | 1.94 | 5s | 25% |  |  | 0.285 |  | 2 | 3 |  | 5 | 128x0.6mm
|  | Fl_Neonate 5.0 I30f 3 | [22 empty] | 5 | 5 | 31 | I30f medium smooth | Baby Abdomen
NeonateBody (Child)
NeonateBody (Child) | Topogram | Topogram 0.6 T20f |  | 80 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.03 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| NeonateBody |  | 80 | 80 | 276 | 94 |  |  | On | 7 | On | CARE Dose4D | 3.76 | 5s | 25% |  |  | 0.5 |  | 2 | 1.4 |  | 5 | 32x1.2mm | [15 empty] | Off
|  | NeonateBody 5.0 B30f | [22 empty] | 5 | 5 | 31 | B30f medium smooth | Baby Abdomen | [13 empty] | manual
NeonateBody_70kV (Child)
NeonateBody_70kV (Child) | Topogram | Topogram 0.6 T20f |  | 70 |  | 164 |  |  |  | 3 | Off | CARE Dose | 0.03 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| NeonateBody |  | 70 | 70 | 507 | 172 |  |  | On | 7 | On | CARE Dose4D | 1.78 | 5s | 25% |  |  | 0.5 |  | 2 | 1.4 |  | 5 | 32x1.2mm | [15 empty] | Off
|  | NeonateBody 5.0 B30f | [22 empty] | 5 | 5 | 31 | B30f medium smooth | Baby Abdomen | [13 empty] | manual
NeonateBody_IR (Child)
NeonateBody_IR (Child) | Topogram | Topogram 0.6 T20f |  | 80 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.03 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| NeonateBody |  | 80 | 80 | 166 | 56 |  |  | On | 7 | On | CARE Dose4D | 2.69 | 5s | 25% |  |  | 0.5 |  | 2 | 1.4 |  | 5 | 32x1.2mm | [15 empty] | Off
|  | NeonateBody 5.0 I30f 3 | [22 empty] | 5 | 5 | 31 | I30f medium smooth | Baby Abdomen | [13 empty] | manual
SELECT_PROTOCOL (Child)
SELECT__PROTOCOL (Child)
UpperExtremities
14000_UppExt_sim_wo_cst_PED (Child)
14000_UppExt_sim_wo_cst_PED (Child) | Topogram | Topogram |  | 100 |  | 109 |  |  |  | 3 | Off | CARE Dose | 0.08 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Body |  |  |  | GEPACS
| UppExt PED |  | 120 | 120 | 250 | 250 |  |  | Off | 3 | Off | CARE Dose4D | 16.87 | 30s | 25% | 40 |  | 0.5 |  | 2 | 1 |  | 0.6 | 128x0.6mm | [15 empty] | Off
|  | UppExt bon pu000 :-D | [22 empty] | 0.6 | 0.6 | 771 | B70f very sharp | Bone |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
|  | UppExt bon pu000 | [22 empty] | 2 | 2 | 76 | B70f very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | UppExt :-D | [22 empty] | 0.6 | 0.6 | 251 | B26f medium smooth ASA | Bone |  |  |  | ROUTER | GEPACS | [8 empty] | manual
|  | UppExt | [22 empty] | 2 | 2 | 76 | B40f medium | Soft Tissue |  |  |  | GEPACS | [9 empty] | manual
14010_UppExt_com_wo_cst_PED (Child)
14010_UppExt_com_wo_cst_PED (Child) | Topogram | Topogram |  | 100 |  | 109 |  |  |  | 3 | Off | CARE Dose | 0.08 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Body |  |  |  | GEPACS
| UppExt PED |  | 120 | 120 | 250 | 250 |  |  | Off | 3 | Off | CARE Dose4D | 16.87 | 30s | 25% | 40 |  | 0.5 |  | 2 | 1 |  | 0.6 | 128x0.6mm | [15 empty] | Off
|  | UppExt bon pu010 :-D | [22 empty] | 0.6 | 0.6 | 771 | B70f very sharp | Bone |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
|  | UppExt bon pu010 | [22 empty] | 2 | 2 | 76 | B70f very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | UppExt :-D | [22 empty] | 0.6 | 0.6 | 251 | B26f medium smooth ASA | Bone |  |  |  | ROUTER | GEPACS | [8 empty] | manual
|  | UppExt | [22 empty] | 2 | 2 | 76 | B40f medium | Soft Tissue |  |  |  | GEPACS | [9 empty] | manual
Extremity (Child)
Extremity (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.07 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Extremity |  |  | 100 | 100 | 35 |  |  |  | 3 | On | CARE Dose | 1.16 | 6s | 25% |  |  | 1 |  | 2 | 0.8 |  | 3 | 128x0.6mm | [15 empty] | Off
|  | Extremity 3.0 B30s | [22 empty] | 3 | 3 | 34 | B30s medium smooth | Baby Extremity | [13 empty] | manual
|  | Extremity 1.0 B30s | [22 empty] | 1 | 0.7 | 143 | B30s medium smooth | Baby Extremity | [13 empty] | manual
ExtrRoutineZUHR (Child)
ExtrRoutineZUHR (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.07 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| ExtrUHR |  |  | 120 | 180 | 90 |  |  |  | 3 | On | CARE Dose | 8.88 | 20s | 25% |  |  | 1 |  | 2 | 0.85 |  | 0.4 | 16x0.3mm | [15 empty] | Off
|  | ExtrUHR 0.4 U80u | [22 empty] | 0.4 | 0.4 | 151 | U80u very sharp | Baby Extremity | [13 empty] | manual
ExtrUHR (Child)
ExtrUHR (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.07 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| ExtrUHR |  |  | 120 | 120 | 42 |  |  |  | 3 | On | CARE Dose | 4.15 | 20s | 25% |  |  | 1 |  | 2 | 0.85 |  | 0.6 | 16x0.6mm | [15 empty] | Off
|  | ExtrUHR 0.6 U80u | [22 empty] | 0.6 | 0.6 | 101 | U80u very sharp | Baby Extremity | [13 empty] | manual
ExtrZUHR_IR (Child)
ExtrZUHR_IR (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.07 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| ExtrUHR |  |  | 120 | 126 | 65 |  |  |  | 3 | On | CARE Dose | 6.46 | 20s | 25% |  |  | 1 |  | 2 | 0.85 |  | 0.4 | 16x0.3mm | [15 empty] | Off
|  | ExtrUHR 0.4 V80u 3 | [22 empty] | 0.4 | 0.4 | 151 | V80u very sharp | Baby Extremity | [13 empty] | manual
LowerExtremities
19000_LowExt_sim_wo_cst_PED (Child)
19000_LowExt_sim_wo_cst_PED (Child) | Topogram | Topogram |  | 100 |  | 109 |  |  |  | 3 | Off | CARE Dose | 0.08 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Body |  |  |  | GEPACS
| LowExt PED |  | 120 | 120 | 250 | 250 |  |  | Off | 3 | Off | CARE Dose4D | 16.87 | 30s | 25% | 40 |  | 0.5 |  | 2 | 1 |  | 0.6 | 128x0.6mm | [15 empty] | Off
|  | LowExt bon pl000 :-D | [22 empty] | 0.6 | 0.6 | 771 | B70f very sharp | Bone |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
|  | LowExt bon pl000 | [22 empty] | 2 | 2 | 76 | B70f very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | LowExt :-D | [22 empty] | 0.6 | 0.6 | 251 | B26f medium smooth ASA | Bone |  |  |  | ROUTER | GEPACS | [8 empty] | manual
|  | LowExt | [22 empty] | 2 | 2 | 76 | B40f medium | Soft Tissue |  |  |  | GEPACS | [9 empty] | manual
19010_LowExt_com_wo_cst_PED (Child)
19010_LowExt_com_wo_cst_PED (Child) | Topogram | Topogram |  | 100 |  | 109 |  |  |  | 3 | Off | CARE Dose | 0.08 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Body |  |  |  | GEPACS
| LowExt PED |  | 120 | 120 | 250 | 250 |  |  | Off | 3 | Off | CARE Dose4D | 16.87 | 30s | 25% | 40 |  | 0.5 |  | 2 | 1 |  | 0.6 | 128x0.6mm | [15 empty] | Off
|  | LowExt bon pl010 :-D | [22 empty] | 0.6 | 0.6 | 771 | B70f very sharp | Bone |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
|  | LowExt bon pl010 | [22 empty] | 2 | 2 | 76 | B70f very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | LowExt :-D | [22 empty] | 0.6 | 0.6 | 251 | B26f medium smooth ASA | Bone |  |  |  | ROUTER | GEPACS | [8 empty] | manual
|  | LowExt | [22 empty] | 2 | 2 | 76 | B40f medium | Soft Tissue |  |  |  | GEPACS | [9 empty] | manual
Extremity (Child)
Extremity (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.07 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Extremity |  |  | 100 | 100 | 35 |  |  |  | 3 | On | CARE Dose | 1.16 | 6s | 25% |  |  | 1 |  | 2 | 0.8 |  | 3 | 128x0.6mm | [15 empty] | Off
|  | Extremity 3.0 B30s | [22 empty] | 3 | 3 | 34 | B30s medium smooth | Baby Extremity | [13 empty] | manual
|  | Extremity 1.0 B30s | [22 empty] | 1 | 0.7 | 143 | B30s medium smooth | Baby Extremity | [13 empty] | manual
ExtrRoutineZUHR (Child)
ExtrRoutineZUHR (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.07 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| ExtrUHR |  |  | 120 | 180 | 90 |  |  |  | 3 | On | CARE Dose | 8.88 | 20s | 25% |  |  | 1 |  | 2 | 0.85 |  | 0.4 | 16x0.3mm | [15 empty] | Off
|  | ExtrUHR 0.4 U80u | [22 empty] | 0.4 | 0.4 | 151 | U80u very sharp | Baby Extremity | [13 empty] | manual
ExtrUHR (Child)
ExtrUHR (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.07 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| ExtrUHR |  |  | 120 | 120 | 42 |  |  |  | 3 | On | CARE Dose | 4.15 | 20s | 25% |  |  | 1 |  | 2 | 0.85 |  | 0.6 | 16x0.6mm | [15 empty] | Off
|  | ExtrUHR 0.6 U80u | [22 empty] | 0.6 | 0.6 | 101 | U80u very sharp | Baby Extremity | [13 empty] | manual
ExtrZUHR_IR (Child)
ExtrZUHR_IR (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.07 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| ExtrUHR |  |  | 120 | 126 | 65 |  |  |  | 3 | On | CARE Dose | 6.46 | 20s | 25% |  |  | 1 |  | 2 | 0.85 |  | 0.4 | 16x0.3mm | [15 empty] | Off
|  | ExtrUHR 0.4 V80u 3 | [22 empty] | 0.4 | 0.4 | 151 | V80u very sharp | Baby Extremity | [13 empty] | manual
Cardiac
Flash_Cardio (Child)
Flash_Cardio (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.07 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Fl_CorCTA |  | 80 | 80 | 594 |  | 298 |  | On | 11 | On | CARE Dose4D | 1.71 |  | 25% |  |  | 0.285 |  | 2 | 3.4 |  | 3 | 128x0.6mm
|  | Fl_CorCTA 3.0 B30f 30% | [22 empty] | 3 | 3 | 24 | B30f medium smooth | Cardiac | [14 empty] | 30%
|  | Fl_CorCTA 0.6 B26f 30% | [22 empty] | 0.6 | 0.3 | 234 | B26f medium smooth ASA | Cardiac | [14 empty] | 30%
Flash_Cardio_IR (Child)
Flash_Cardio_IR (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.07 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Fl_CorCTA |  | 80 | 80 | 416 |  | 298 |  | On | 11 | On | CARE Dose4D | 1.71 |  | 25% |  |  | 0.285 |  | 2 | 3.4 |  | 3 | 128x0.6mm
|  | Fl_CorCTA 3.0 I30f 3 30% | [22 empty] | 3 | 3 | 24 | I30f medium smooth | Cardiac | [14 empty] | 30%
|  | Fl_CorCTA 0.6 I26f 3 30% | [22 empty] | 0.6 | 0.3 | 234 | I26f medium smooth ASA | Cardiac | [14 empty] | 30%
Vascular
BodyAngio (Child)
BodyAngio (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.07 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| BodyAngio |  | 80 | 80 | 223 | 112 |  |  | On | 11 | On | CARE Dose4D | 3.51 | 5s | 25% |  |  | 0.5 |  | 2 | 1.4 |  | 3 | 128x0.6mm | [15 empty] | Off
|  | BodyAngio 3.0 B30f | [22 empty] | 3 | 3 | 67 | B30f medium smooth | Baby Abdomen | [13 empty] | manual
|  | BodyAngio 0.6 B26f | [22 empty] | 0.6 | 0.4 | 501 | B26f medium smooth ASA | CT Angio | [13 empty] | manual
BodyAngio_IR (Child)
BodyAngio_IR (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.07 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| BodyAngio |  | 80 | 80 | 156 | 78 |  |  | On | 11 | On | CARE Dose4D | 2.46 | 5s | 25% |  |  | 0.5 |  | 2 | 1.4 |  | 3 | 128x0.6mm | [15 empty] | Off
|  | BodyAngio 3.0 I30f 3 | [22 empty] | 3 | 3 | 67 | I30f medium smooth | Baby Abdomen | [13 empty] | manual
|  | BodyAngio 0.6 I26f 3 | [22 empty] | 0.6 | 0.4 | 501 | I26f medium smooth ASA | CT Angio | [13 empty] | manual
CarotidAngio (Child)
CarotidAngio (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.07 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| CarotidAngio |  | 80 | 80 | 223 | 112 |  |  | On | 11 | On | CARE Dose4D | 3.51 | 5s | 25% |  |  | 0.5 |  | 2 | 1.4 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | CarotidAngio 5.0 B30f | [22 empty] | 5 | 5 | 35 | B30f medium smooth | Baby Neck | [13 empty] | manual
|  | CarotidAngio 0.6 B26f | [22 empty] | 0.6 | 0.4 | 426 | B26f medium smooth ASA | CT Angio | [13 empty] | manual
CarotidAngio_IR (Child)
CarotidAngio_IR (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.07 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| CarotidAngio |  | 80 | 80 | 156 | 78 |  |  | On | 11 | On | CARE Dose4D | 2.46 | 5s | 25% |  |  | 0.5 |  | 2 | 1.4 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | CarotidAngio 5.0 I30f 3 | [22 empty] | 5 | 5 | 35 | I30f medium smooth | Baby Neck | [13 empty] | manual
|  | CarotidAngio 0.6 I26f 3 | [22 empty] | 0.6 | 0.4 | 426 | I26f medium smooth ASA | CT Angio | [13 empty] | manual
Flash_BodyAngio (Child)
Flash_BodyAngio (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.07 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Fl_BodyAngio |  | 80 | 80 | 224 |  | 112 |  | On | 11 | On | CARE Dose4D | 1.94 | 5s | 25% |  |  | 0.285 |  | 2 | 3 |  | 3 | 128x0.6mm
|  | Fl_BodyAngio 3.0 B30f | [22 empty] | 3 | 3 | 67 | B30f medium smooth | Baby Abdomen
|  | Fl_BodyAngio 0.6 B26f | [22 empty] | 0.6 | 0.4 | 501 | B26f medium smooth ASA | CT Angio
Flash_BodyAngio_IR (Child)
Flash_BodyAngio_IR (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.07 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Fl_BodyAngio |  | 80 | 80 | 158 |  | 80 |  | On | 11 | On | CARE Dose4D | 1.94 | 5s | 25% |  |  | 0.285 |  | 2 | 3 |  | 3 | 128x0.6mm
|  | Fl_BodyAngio 3.0 I30f 3 | [22 empty] | 3 | 3 | 67 | I30f medium smooth | Baby Abdomen
|  | Fl_BodyAngio 0.6 I26f 3 | [22 empty] | 0.6 | 0.4 | 501 | I26f medium smooth ASA | CT Angio
HeadAngio (Child)
HeadAngio (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.16 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head
| HeadAngio |  | 80 | 80 | 221 | 111 |  |  | On | 11 | On | CARE Dose4D | 8.43 | 5s | 25% |  |  | 0.5 |  | 2 | 1.4 |  | 2 | 128x0.6mm | [15 empty] | Off
|  | HeadAngio 2.0 H20f | [22 empty] | 2 | 2 | 31 | H20f smooth | Baby Brain | [13 empty] | manual
|  | HeadAngio 0.6 H10f | [22 empty] | 0.6 | 0.4 | 151 | H10f very smooth | CT Angio | [13 empty] | manual
HeadAngio_IR (Child)
HeadAngio_IR (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.16 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head
| HeadAngio |  | 80 | 80 | 155 | 78 |  |  | On | 11 | On | CARE Dose4D | 5.91 | 5s | 25% |  |  | 0.5 |  | 2 | 1.4 |  | 2 | 128x0.6mm | [15 empty] | Off
|  | HeadAngio 2.0 J30f 3 | [22 empty] | 2 | 2 | 31 | J30f medium smooth | Baby Brain | [13 empty] | manual
|  | HeadAngio 0.6 J30f 3 | [22 empty] | 0.6 | 0.4 | 151 | J30f medium smooth | CT Angio | [13 empty] | manual
HeadAngio_XCARE (Child)
HeadAngio_XCARE (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.16 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head
| HeadAngio |  | 80 | 80 | 221 | 111 |  |  | On | 11 | On | CARE Dose4D | 8.45 | 5s | 25% |  |  | 0.285 |  | 2 | 0.6 |  | 2 | 128x0.6mm | [15 empty] | Off
|  | HeadAngio 2.0 H20f | [22 empty] | 2 | 2 | 31 | H20f smooth | Baby Brain |  |  | X-CARE | [10 empty] | manual
|  | HeadAngio 0.6 H10f | [22 empty] | 0.6 | 0.4 | 151 | H10f very smooth | CT Angio |  |  | X-CARE | [10 empty] | manual
HeadAngio_XCARE_IR (Child)
HeadAngio_XCARE_IR (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.16 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head
| HeadAngio |  | 80 | 80 | 155 | 78 |  |  | On | 11 | On | CARE Dose4D | 5.93 | 5s | 25% |  |  | 0.285 |  | 2 | 0.6 |  | 2 | 128x0.6mm | [15 empty] | Off
|  | HeadAngio 2.0 J30f 3 | [22 empty] | 2 | 2 | 31 | J30f medium smooth | Baby Brain |  |  | X-CARE | [10 empty] | manual
|  | HeadAngio 0.6 J30f 3 | [22 empty] | 0.6 | 0.4 | 151 | J30f medium smooth | CT Angio |  |  | X-CARE | [10 empty] | manual
Private
ImageGently1T_Customized (Child)
ImageGently1T_Customized (Child) | Topogram | Topogram 0.6 T20f |  | 120 |  | 696 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| 50cm 1T |  | 120 | 120 | 240 | 90 |  |  | Off | 7 | On | CARE Dose4D | 13.28 | 20s | 25% |  |  | 0.5 |  | 2 | 0.8 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | 50cm 1T 5.0 B40f | [22 empty] | 5 | 5 | 13 | B40f medium | 400/40 Abd | [13 empty] | manual
| 45cm 1T |  | 120 | 120 | 240 | 90 |  |  | Off | 7 | On | CARE Dose4D | 13.28 | 20s | 25% |  |  | 0.5 |  | 2 | 0.8 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | 45cm 1T 5.0 B40f | [22 empty] | 5 | 5 | 13 | B40f medium | 400/40 Abd | [13 empty] | manual
| 40cm 1T |  | 120 | 120 | 240 | 90 |  |  | Off | 7 | On | CARE Dose4D | 13.28 | 20s | 25% |  |  | 0.5 |  | 2 | 0.8 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | 40cm 1T 5.0 B40f | [22 empty] | 5 | 5 | 13 | B40f medium | 400/40 Abd | [13 empty] | manual
| 35cm 1T |  | 120 | 120 | 240 | 90 |  |  | Off | 7 | On | CARE Dose4D | 13.28 | 20s | 25% |  |  | 0.5 |  | 2 | 0.8 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | 35cm 1T 5.0 B40f | [22 empty] | 5 | 5 | 13 | B40f medium | 400/40 Abd | [13 empty] | manual
| 30cm 1T |  | 120 | 120 | 240 | 90 |  |  | Off | 7 | On | CARE Dose4D | 13.28 | 20s | 25% |  |  | 0.5 |  | 2 | 0.8 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | 30cm 1T 5.0 B40f | [22 empty] | 5 | 5 | 13 | B40f medium | 400/40 Abd | [13 empty] | manual
| 25cm 1T |  | 120 | 120 | 240 | 90 |  |  | Off | 7 | On | CARE Dose4D | 13.28 | 20s | 25% |  |  | 0.5 |  | 2 | 0.8 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | 25cm 1T 5.0 B40f | [22 empty] | 5 | 5 | 13 | B40f medium | 400/40 Abd | [13 empty] | manual
| 20cm 1T |  | 120 | 120 | 240 | 90 |  |  | Off | 7 | On | CARE Dose4D | 13.28 | 20s | 25% |  |  | 0.5 |  | 2 | 0.8 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | 20cm 1T 5.0 B40f | [22 empty] | 5 | 5 | 13 | B40f medium | 400/40 Abd | [13 empty] | manual
| 15cm 1T |  | 120 | 120 | 240 | 90 |  |  | Off | 7 | On | CARE Dose4D | 13.28 | 20s | 25% |  |  | 0.5 |  | 2 | 0.8 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | 15cm 1T 5.0 B40f | [22 empty] | 5 | 5 | 13 | B40f medium | 400/40 Abd | [13 empty] | manual
| 10cm 1T |  | 120 | 120 | 240 | 90 |  |  | Off | 7 | On | CARE Dose4D | 13.28 | 20s | 25% |  |  | 0.5 |  | 2 | 0.8 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | 10cm 1T 5.0 B40f | [22 empty] | 5 | 5 | 13 | B40f medium | 400/40 Abd | [13 empty] | manual
ImageGently2T_Customized (Child)
ImageGently2T_Customized (Child) | Topogram | Topogram 0.6 T20f |  | 120 |  | 696 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Topogram | Topogram 0.6 T20f |  | 120 |  | 696 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| 50cm 2T |  | 120 | 120 | 240 | 90 |  |  | Off | 7 | On | CARE Dose4D | 13.28 | 20s | 25% |  |  | 0.5 |  | 2 | 0.8 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | 50cm 2T 5.0 B40f | [22 empty] | 5 | 5 | 13 | B40f medium | 400/40 Abd | [13 empty] | manual
| 45cm 2T |  | 120 | 120 | 240 | 90 |  |  | Off | 7 | On | CARE Dose4D | 13.28 | 20s | 25% |  |  | 0.5 |  | 2 | 0.8 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | 45cm 2T 5.0 B40f | [22 empty] | 5 | 5 | 13 | B40f medium | 400/40 Abd | [13 empty] | manual
| 40cm 2T |  | 120 | 120 | 240 | 90 |  |  | Off | 7 | On | CARE Dose4D | 13.28 | 20s | 25% |  |  | 0.5 |  | 2 | 0.8 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | 40cm 2T 5.0 B40f | [22 empty] | 5 | 5 | 13 | B40f medium | 400/40 Abd | [13 empty] | manual
| 35cm 2T |  | 120 | 120 | 240 | 90 |  |  | Off | 7 | On | CARE Dose4D | 13.28 | 20s | 25% |  |  | 0.5 |  | 2 | 0.8 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | 35cm 2T 5.0 B40f | [22 empty] | 5 | 5 | 13 | B40f medium | 400/40 Abd | [13 empty] | manual
| 30cm 2T |  | 120 | 120 | 240 | 90 |  |  | Off | 7 | On | CARE Dose4D | 13.28 | 20s | 25% |  |  | 0.5 |  | 2 | 0.8 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | 30cm 2T 5.0 B40f | [22 empty] | 5 | 5 | 13 | B40f medium | 400/40 Abd | [13 empty] | manual
| 25cm 2T |  | 120 | 120 | 240 | 90 |  |  | Off | 7 | On | CARE Dose4D | 13.28 | 20s | 25% |  |  | 0.5 |  | 2 | 0.8 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | 25cm 2T 5.0 B40f | [22 empty] | 5 | 5 | 13 | B40f medium | 400/40 Abd | [13 empty] | manual
| 20cm 2T |  | 120 | 120 | 240 | 90 |  |  | Off | 7 | On | CARE Dose4D | 13.28 | 20s | 25% |  |  | 0.5 |  | 2 | 0.8 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | 20cm 2T 5.0 B40f | [22 empty] | 5 | 5 | 13 | B40f medium | 400/40 Abd | [13 empty] | manual
| 15cm 2T |  | 120 | 120 | 240 | 90 |  |  | Off | 7 | On | CARE Dose4D | 13.28 | 20s | 25% |  |  | 0.5 |  | 2 | 0.8 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | 15cm 2T 5.0 B40f | [22 empty] | 5 | 5 | 13 | B40f medium | 400/40 Abd | [13 empty] | manual
| 10cm 1T |  | 120 | 120 | 240 | 90 |  |  | Off | 7 | On | CARE Dose4D | 13.28 | 20s | 25% |  |  | 0.5 |  | 2 | 0.8 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | 10cm 1T 5.0 B40f | [22 empty] | 5 | 5 | 13 | B40f medium | 400/40 Abd | [13 empty] | manual
RT
RT_Abdomen (Child)
RT_Abdomen (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.07 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Abdomen |  | 120 | 120 | 250 | 125 |  |  | Semi | 3 | On | CARE Dose4D | 14.21 | 99s | 25% |  |  | 0.5 |  | 2 | 0.6 |  | 4 | 32x1.2mm | [15 empty] | Off
|  | Abdomen 4.0 B30f | [22 empty] | 4 | 4 | 51 | B30f medium smooth | Baby Abdomen | [13 empty] | manual
RT_Head (Child)
RT_Head (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.16 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head
| Head |  | 120 | 120 | 410 | 205 |  |  | Off | 3 | On | CARE Dose4D | 63.87 | 99s | 25% |  |  | 1 |  | 2 | 0.8 |  | 3 | 32x1.2mm | [15 empty] | Off
|  | Head 3.0 H30s | [22 empty] | 3 | 3 | 41 | H30s medium smooth | Baby Brain | [13 empty] | manual
Head
1010_ang_Hea (Adult)
1010_ang_Hea (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS |  | TRACKER
| PreMonitoring |  | 120 | 100 | 111 | 100 |  |  | Off | 11 | Off | CARE Dose4D | 8.03 |  | 25% |  |  | 0.5 |  | 2 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | PreMonitoring 10.0 H40s | [22 empty] | 10 |  | 1 | H40s medium | 400/40 Abd | [13 empty] | manual
| Monitoring |  | 120 | 100 | 111 | 30 |  |  | Off | 11 | Off | CARE Dose4D | 96.4 |  | 25% |  |  | 0.5 |  | 10 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | Monitoring 10.0 H40s | [22 empty] | 10 |  | 40 | H40s medium | 400/40 Abd | [13 empty] | manual
| DE art Hea |  |  | 100 | 250 |  | 200 | 200 |  | 11 | On | CARE Dose4D | 43.02 | 30s | 25% | 80 |  | 0.5 |  | 2 | 0.8 |  | 0.75 | 32x0.6mm | [15 empty] | Off
|  | art Hea h010 :-D | [22 empty] | 0.75 | 0.7 | 358 | J40f medium | 600/175 Neuro |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
|  | art Hea h010 | [22 empty] | 2 | 2 | 126 | J40f medium | 600/175 Neuro |  |  |  | GEPACS | [9 empty] | manual
|  | art Hea cor mip | [22 empty] | 7 | 2.5 | 101 | J40f medium | Head MIP |  |  |  | GEPACS | [9 empty] | manual
|  | art Hea sag mip | [22 empty] | 7 | 2.5 | 101 | J40f medium | Head MIP |  |  |  | GEPACS | [9 empty] | manual
|  | art Hea mip | [22 empty] | 7 | 2.5 | 101 | J40f medium | Head MIP |  |  |  | GEPACS | [9 empty] | manual
|  | art Hea cor | [22 empty] | 2 | 2 | 126 | J40f medium | Head |  |  |  | GEPACS | [9 empty] | manual
|  | art Hea sag | [22 empty] | 2 | 2 | 126 | J40f medium | Head |  |  |  | GEPACS | [9 empty] | manual
|  | art Hea h010 | [22 empty] | 0.75 | 0.4 | 626 | D20f smooth | 600/175 Neuro |  |  |  | GEPACS | SYNGO_VIA | [8 empty] | manual
1015_ang_Hea_fus_nav (Adult)
1015_ang_Hea_fus_nav (Adult) | Topogram | Topogram |  | 100 |  | 137 |  |  |  | 11 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 137 |  |  |  | 11 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS |  | TRACKER
| PreMonitoring |  | 120 | 120 | 175 | 100 |  |  | Off | 11 | Off | CARE Dose4D | 13.9 |  | 25% |  |  | 0.5 |  | 2 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | PreMonitoring | [22 empty] | 10 |  | 1 | H20s smooth | 400/40 Abd | [13 empty] | manual
| Monitoring |  | 120 | 120 | 175 | 100 |  |  | Off | 11 | Off | CARE Dose4D | 556.17 |  | 25% |  |  | 0.5 |  | 10 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | Monitoring | [22 empty] | 10 |  | 40 | H20s smooth | 400/40 Abd | [13 empty] | manual
| art Hea |  | 120 | 120 | 265 | 265 |  |  | Off | 11 | Off | CARE Dose4D | 40.59 | 30s | 25% | 80 |  | 0.5 |  | 2 | 0.7 |  | 0.6 | 128x0.6mm | [15 empty] | Off
|  | art Hea FUSION h015 | [22 empty] | 0.6 | 0.6 | 201 | J40f medium | Head |  |  |  | OR_MIDIA_POOL | ROUTER | GEPACS | [7 empty] | manual
|  | art Hea bon | [22 empty] | 3 | 2 | 61 | J45f medium | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | art Hea cor | [22 empty] | 2 | 2 | 126 | J40f medium | Head |  |  |  | GEPACS | [9 empty] | manual
|  | art Hea sag | [22 empty] | 2 | 2 | 126 | J40f medium | Head |  |  |  | GEPACS | [9 empty] | manual
|  | art Hea mip | [22 empty] | 7 | 2.5 | 101 | J40f medium | Head MIP |  |  |  | GEPACS | [9 empty] | manual
|  | art Hea cor mip | [22 empty] | 7 | 2.5 | 101 | J40f medium | Head MIP |  |  |  | GEPACS | [9 empty] | manual
|  | art Hea sag mip | [22 empty] | 7 | 2.5 | 101 | J40f medium | Head MIP |  |  |  | GEPACS | [9 empty] | manual
1020_ang_Hea_int_ven (Adult)
1020_ang_Hea_int_ven (Adult) | Topogram | Topogram |  | 100 |  | 137 |  |  |  | 11 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 137 |  |  |  | 11 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS | ROEMIR001a | TRACKER
| DE ven Hea |  |  | 100 | 250 |  | 200 | 200 |  | 11 | On | CARE Dose4D | 43.02 | 30s | 25% | 80 |  | 0.5 |  | 45 | 0.8 |  | 0.75 | 32x0.6mm | [15 empty] | Off
|  | ven Hea h020 :-D | [22 empty] | 0.75 | 0.7 | 358 | J40f medium | 600/175 Neuro |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
|  | ven Hea h020 | [22 empty] | 2 | 2 | 126 | J40f medium | 600/175 Neuro |  |  |  | GEPACS | [9 empty] | manual
|  | ven Hea cor mip | [22 empty] | 7 | 2.5 | 101 | J40f medium | Head MIP |  |  |  | GEPACS | [9 empty] | manual
|  | ven Hea sag mip | [22 empty] | 7 | 2.5 | 101 | J40f medium | Head MIP |  |  |  | GEPACS | [9 empty] | manual
|  | ven Hea mip | [22 empty] | 7 | 2.5 | 101 | J40f medium | Head MIP |  |  |  | GEPACS | [9 empty] | manual
|  | ven Hea cor | [22 empty] | 2 | 2 | 126 | J40f medium | Head |  |  |  | GEPACS | [9 empty] | manual
|  | ven Hea sag | [22 empty] | 2 | 2 | 126 | J40f medium | Head |  |  |  | GEPACS | [9 empty] | manual
|  | ven Hea h020 | [22 empty] | 0.75 | 0.4 | 626 | D20f smooth | 600/175 Neuro |  |  |  | MMWP1 | [9 empty] | manual
1030_ang_HeaNec (Adult)
1030_ang_HeaNec (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| PreMonitoring |  | 120 | 100 | 111 | 100 |  |  | Off | 11 | Off | CARE Dose4D | 8.03 |  | 25% |  |  | 0.5 |  | 2 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | PreMonitoring 10.0 H40s | [22 empty] | 10 |  | 1 | H40s medium | 400/40 Abd | [13 empty] | manual
| Monitoring |  | 120 | 100 | 111 | 30 |  |  | Off | 11 | Off | CARE Dose4D | 96.4 |  | 25% |  |  | 0.5 |  | 10 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | Monitoring 10.0 H40s | [22 empty] | 10 |  | 40 | H40s medium | 400/40 Abd | [13 empty] | manual
| DE art HeaNec |  |  | 100 | 250 |  | 200 | 200 |  | 11 | On | CARE Dose4D | 43.02 | 30s | 25% | 80 |  | 0.5 |  | 2 | 0.8 |  | 0.75 | 32x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | art HeaNec h030 :-D | [22 empty] | 0.75 | 0.7 | 358 | J40f medium | 600/175 Neuro |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
|  | art HeaNec h030 | [22 empty] | 2 | 2 | 126 | J40f medium | 600/175 Neuro |  |  |  | GEPACS | [9 empty] | manual
|  | art HeaNec cor mip | [22 empty] | 7 | 2.5 | 101 | J40f medium | Head MIP |  |  |  | GEPACS | [9 empty] | manual
|  | art HeaNec sag mip | [22 empty] | 7 | 2.5 | 101 | J40f medium | Head MIP |  |  |  | GEPACS | [9 empty] | manual
|  | art HeaNec mip | [22 empty] | 7 | 2.5 | 101 | J40f medium | Head MIP |  |  |  | GEPACS | [9 empty] | manual
|  | art HeaNec cor | [22 empty] | 2 | 2 | 126 | J40f medium | Head |  |  |  | GEPACS | [9 empty] | manual
|  | art HeaNec sag | [22 empty] | 2 | 2 | 126 | J40f medium | Head |  |  |  | GEPACS | [9 empty] | manual
|  | art HeaNec h030 #PP | [22 empty] | 0.75 | 0.4 | 626 | D20f smooth | 600/175 Neuro |  |  |  | MMWP1 | [9 empty] | manual
1040_ang_HeaNec_cer_per_N03_01A (Adult)
1040_ang_HeaNec_cer_per_N03_01A (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS | ROEMIR001a | TRACKER
| Hea |  | 120 | 120 | 390 | 290 |  |  | Off | 3 | Off | CARE Dose4D | 44.37 | 15s | 25% | 80 |  | 0.5 |  | 2 | 0.6 |  | 4 | 128x0.6mm | [15 empty] | Off
|  | Hea h150 | [22 empty] | 4 | 4 | 31 | J40f medium | Head Soft |  |  | X-CARE | GEPACS | [9 empty] | manual
|  | Hea bon | [22 empty] | 4 | 4 | 31 | H60f sharp FR | Bone |  |  | X-CARE | GEPACS | [9 empty] | manual
|  | Hea cor | [22 empty] | 3 | 3 | 84 | J40f medium | Head Soft |  |  | X-CARE | GEPACS | [9 empty] | manual
|  | Hea sag | [22 empty] | 3 | 3 | 84 | J40f medium | Head Soft |  |  | X-CARE | GEPACS | [9 empty] | manual
|  | Hea h150 :-D | [22 empty] | 1 | 1 | 121 | J40f medium | Head Soft |  |  | X-CARE | ROUTER | GEPACS | [8 empty] | manual
| per Hea |  | 120 | 80 |  | 200 |  |  | Off | 11 | Off | CARE Dose | 218.08 |  | 25% | 600 |  | 0.285 |  | 2 | 0.5 |  | 5 | 32x1.2mm
|  | per Hea h040 | [22 empty] | 5 | 5 | 416 | H20f smooth | 600/175 Neuro CTA
| PreMonitoring |  | 120 | 100 | 111 | 100 |  |  | Off | 11 | Off | CARE Dose4D | 8.03 |  | 25% |  |  | 0.5 |  | 2 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | PreMonitoring 10.0 H40s | [22 empty] | 10 |  | 1 | H40s medium | 400/40 Abd | [13 empty] | manual
| Monitoring |  | 120 | 100 | 111 | 30 |  |  | Off | 11 | Off | CARE Dose4D | 96.4 |  | 25% |  |  | 0.5 |  | 10 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | Monitoring 10.0 H40s | [22 empty] | 10 |  | 40 | H40s medium | 400/40 Abd | [13 empty] | manual
| DE ang HeaNec |  |  | 100 | 250 |  | 200 | 200 |  | 11 | On | CARE Dose4D | 43.02 | 30s | 25% | 80 |  | 0.5 |  | 2 | 0.8 |  | 0.75 | 32x0.6mm | [15 empty] | Off
|  | art HeaNec h030 :-D | [22 empty] | 0.75 | 0.7 | 358 | J40f medium | 600/175 Neuro |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
|  | art HeaNec h030 | [22 empty] | 2 | 2 | 126 | J40f medium | 600/175 Neuro |  |  |  | GEPACS | [9 empty] | manual
|  | art HeaNec cor mip | [22 empty] | 7 | 2.5 | 101 | J40f medium | Head MIP |  |  |  | GEPACS | [9 empty] | manual
|  | art HeaNec sag mip | [22 empty] | 7 | 2.5 | 101 | J40f medium | Head MIP |  |  |  | GEPACS | [9 empty] | manual
|  | art HeaNec mip | [22 empty] | 7 | 2.5 | 101 | J40f medium | Head MIP |  |  |  | GEPACS | [9 empty] | manual
|  | art HeaNec cor | [22 empty] | 2 | 2 | 126 | J40f medium | Head |  |  |  | GEPACS | [9 empty] | manual
|  | art HeaNec sag | [22 empty] | 2 | 2 | 126 | J40f medium | Head |  |  |  | GEPACS | [9 empty] | manual
|  | art HeaNec h030 | [22 empty] | 0.75 | 0.4 | 626 | D20f smooth | 600/175 Neuro |  |  |  | MMWP1 | [9 empty] | manual
1050_ang_HeaNec_thu (Adult)
1050_ang_HeaNec_thu (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS | ROEMIR001a | TRACKER
| per Hea |  | 120 | 80 |  | 200 |  |  | Off | 11 | Off | CARE Dose | 218.08 |  | 25% | 600 |  | 0.285 |  | 2 | 0.5 |  | 5 | 32x1.2mm
|  | per Hea h040 | [22 empty] | 5 | 5 | 416 | H20f smooth | 600/175 Neuro CTA
| PreMonitoring |  | 120 | 100 | 111 | 100 |  |  | Off | 11 | Off | CARE Dose4D | 8.03 |  | 25% |  |  | 0.5 |  | 2 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | PreMonitoring 10.0 H40s | [22 empty] | 10 |  | 1 | H40s medium | 400/40 Abd | [13 empty] | manual
| Monitoring |  | 120 | 100 | 111 | 30 |  |  | Off | 11 | Off | CARE Dose4D | 96.4 |  | 25% |  |  | 0.5 |  | 10 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | Monitoring 10.0 H40s | [22 empty] | 10 |  | 40 | H40s medium | 400/40 Abd | [13 empty] | manual
| DE art HeaNec |  |  | 100 | 250 |  | 200 | 200 |  | 11 | On | CARE Dose4D | 43.02 | 30s | 25% | 80 |  | 0.5 |  | 2 | 0.8 |  | 0.75 | 32x0.6mm | [15 empty] | Off
|  | art HeaNec h030 :-D | [22 empty] | 0.75 | 0.7 | 358 | J40f medium | 600/175 Neuro |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
|  | art HeaNec h030 | [22 empty] | 2 | 2 | 126 | J40f medium | 600/175 Neuro |  |  |  | GEPACS | [9 empty] | manual
|  | art HeaNec cor mip | [22 empty] | 7 | 2.5 | 101 | J40f medium | Head MIP |  |  |  | GEPACS | [9 empty] | manual
|  | art HeaNec sag mip | [22 empty] | 7 | 2.5 | 101 | J40f medium | Head MIP |  |  |  | GEPACS | [9 empty] | manual
|  | art HeaNec mip | [22 empty] | 7 | 2.5 | 101 | J40f medium | Head MIP |  |  |  | GEPACS | [9 empty] | manual
|  | art HeaNec cor | [22 empty] | 2 | 2 | 126 | J40f medium | Head |  |  |  | GEPACS | [9 empty] | manual
|  | art HeaNec sag | [22 empty] | 2 | 2 | 126 | J40f medium | Head |  |  |  | GEPACS | [9 empty] | manual
|  | art HeaNec h030 | [22 empty] | 0.75 | 0.4 | 626 | D20f smooth | 600/175 Neuro |  |  |  | MMWP1 | [9 empty] | manual
| 3m Hea |  | 120 | 120 | 390 | 290 |  |  | Off | 3 | Off | CARE Dose4D | 44.37 | 15s | 25% | 80 |  | 0.5 |  | 180 | 0.6 |  | 4 | 128x0.6mm | [15 empty] | Off
|  | 3m Hea h100 | [22 empty] | 4 | 4 | 31 | J40f medium | Head Soft |  |  | X-CARE | GEPACS | [9 empty] | manual
|  | 3m Hea bon | [22 empty] | 4 | 4 | 31 | H60f sharp FR | Bone |  |  | X-CARE | GEPACS | [9 empty] | manual
|  | 3m Hea cor | [22 empty] | 3 | 3 | 84 | J40f medium | Head Soft |  |  | X-CARE | GEPACS | [9 empty] | manual
|  | 3m Hea sag | [22 empty] | 3 | 3 | 84 | J40f medium | Head Soft |  |  | X-CARE | GEPACS | [9 empty] | manual
|  | 3m Hea h100 :-D | [22 empty] | 1 | 1 | 121 | J40f medium | Head Soft |  |  | X-CARE | ROUTER | GEPACS | [8 empty] | manual
1055_ang_HeaNecChe_2 (Adult)
1055_ang_HeaNecChe_2 (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| PreMonitoring |  | 120 | 120 | 180 | 50 |  |  | Off | 3 | Off | CARE Dose4D | 3.08 |  | 25% |  |  | 0.5 |  | 2 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | PreMonitoring | [22 empty] | 10 |  | 1 | B40s medium | 400/40 Abd | [13 empty] | manual
| Monitoring |  | 120 | 120 | 180 | 50 |  |  | Off | 3 | Off | CARE Dose4D | 123.28 |  | 25% |  |  | 0.5 |  | 10 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | Monitoring | [22 empty] | 10 |  | 40 | B40s medium | 400/40 Abd | [13 empty] | manual
| art HeaNec... |  | 120 | 120 | 180 | 110 |  |  | On | 8 | On | CARE Dose4D | 12.14 | 30s | 25% | 80 |  | 0.5 |  | 4 | 1 |  | 0.75 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | art HeaNecChe h055 :-D | [22 empty] | 0.75 | 0.7 | 412 | I30f medium smooth | Mediastinum |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
1056_ang_HeaNecCheAbd (Adult)
1056_ang_HeaNecCheAbd (Adult) | Topogram | Topogram |  | 100 |  | 393 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | TRACKER
| Topogram | Topogram |  | 100 |  | 393 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| PreMonitoring |  | 120 | 120 | 180 | 50 |  |  | Off | 3 | Off | CARE Dose4D | 3.08 |  | 25% |  |  | 0.5 |  | 2 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | PreMonitoring | [22 empty] | 10 |  | 1 | B40s medium | 400/40 Abd | [13 empty] | manual
| Monitoring |  | 120 | 120 | 180 | 50 |  |  | Off | 3 | Off | CARE Dose4D | 123.28 |  | 25% |  |  | 0.5 |  | 10 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | Monitoring | [22 empty] | 10 |  | 40 | B40s medium | 400/40 Abd | [13 empty] | manual
| art HeaNec... |  | 120 | 120 | 180 | 110 |  |  | On | 8 | On | CARE Dose4D | 12.14 | 30s | 25% | 80 |  | 0.5 |  | 4 | 1 |  | 0.75 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | art HeaNecCheAbd h056 :-D | [22 empty] | 0.75 | 0.7 | 412 | I30f medium smooth | Mediastinum |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
1057_ang_HeaNecCheAbdPel (Adult)
1057_ang_HeaNecCheAbdPel (Adult) | Topogram | Topogram |  | 100 |  | 521 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | TRACKER
| Topogram | Topogram |  | 100 |  | 521 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| PreMonitoring |  | 120 | 120 | 180 | 50 |  |  | Off | 3 | Off | CARE Dose4D | 3.08 |  | 25% |  |  | 0.5 |  | 2 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | PreMonitoring | [22 empty] | 10 |  | 1 | B40s medium | 400/40 Abd | [13 empty] | manual
| Monitoring |  | 120 | 120 | 180 | 50 |  |  | Off | 3 | Off | CARE Dose4D | 123.28 |  | 25% |  |  | 0.5 |  | 10 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | Monitoring | [22 empty] | 10 |  | 40 | B40s medium | 400/40 Abd | [13 empty] | manual
| art HeaNec... |  | 120 | 120 | 180 | 110 |  |  | On | 8 | On | CARE Dose4D | 12.14 | 30s | 25% | 80 |  | 0.5 |  | 4 | 1 |  | 0.75 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | art HeaNecCheAbdPel h057 :-D | [22 empty] | 0.75 | 0.7 | 412 | I30f medium smooth | Mediastinum |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
1060_ang_Nec (Adult)
1060_ang_Nec (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| PreMonitoring |  | 120 | 100 | 111 | 100 |  |  | Off | 11 | Off | CARE Dose4D | 8.03 |  | 25% |  |  | 0.5 |  | 2 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | PreMonitoring 10.0 H40s | [22 empty] | 10 |  | 1 | H40s medium | 400/40 Abd | [13 empty] | manual
| Monitoring |  | 120 | 100 | 111 | 30 |  |  | Off | 11 | Off | CARE Dose4D | 96.4 |  | 25% |  |  | 0.5 |  | 10 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | Monitoring 10.0 H40s | [22 empty] | 10 |  | 40 | H40s medium | 400/40 Abd | [13 empty] | manual
| DE art Nec |  |  | 100 | 250 |  | 200 | 200 |  | 11 | On | CARE Dose4D | 43.02 | 30s | 25% | 80 |  | 0.5 |  | 2 | 0.8 |  | 0.75 | 32x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | art Nec h060 :-D | [22 empty] | 0.75 | 0.7 | 358 | J40f medium | 600/175 Neuro |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
|  | art Nec h060 | [22 empty] | 2 | 2 | 126 | J40f medium | 600/175 Neuro |  |  |  | GEPACS | [9 empty] | manual
|  | art Nec cor mip | [22 empty] | 7 | 2.5 | 101 | J40f medium | Head MIP |  |  |  | GEPACS | [9 empty] | manual
|  | art Nec sag mip | [22 empty] | 7 | 2.5 | 101 | J40f medium | Head MIP |  |  |  | GEPACS | [9 empty] | manual
|  | art Nec mip | [22 empty] | 7 | 2.5 | 101 | J40f medium | Head MIP |  |  |  | GEPACS | [9 empty] | manual
|  | art Nec cor | [22 empty] | 2 | 2 | 126 | J40f medium | Head |  |  |  | GEPACS | [9 empty] | manual
|  | art Nec sag | [22 empty] | 2 | 2 | 126 | J40f medium | Head |  |  |  | GEPACS | [9 empty] | manual
|  | art Nec h060 | [22 empty] | 0.75 | 0.4 | 626 | D20f smooth | 600/175 Neuro |  |  |  | MMWP1 | [9 empty] | manual
1070_ang_Nec_4_pha_par (Adult)
1070_ang_Nec_4_pha_par (Adult) | Topogram | Topogram |  | 100 |  | 137 |  |  |  | 11 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 137 |  |  |  | 11 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS | ROEMIR001a | TRACKER
| Nec |  | 120 | 120 | 210 | 265 |  |  | Off | 11 | On | CARE Dose4D | 32.13 | 30s | 25% | 80 |  | 0.5 |  | 2 | 0.6 |  | 3 | 128x0.6mm | [15 empty] | Off
|  | Nec h070 | [22 empty] | 3 | 2 | 61 | J40f medium | Head |  |  |  | GEPACS | [9 empty] | manual
|  | art Nec cor mip | [22 empty] | 7 | 2.5 | 101 | J40f medium | Head MIP |  |  |  | GEPACS | [9 empty] | manual
|  | art Nec sag mip | [22 empty] | 7 | 2.5 | 101 | J40f medium | Head MIP |  |  |  | GEPACS | [9 empty] | manual
|  | Nec mip | [22 empty] | 3 | 2 | 61 | J40f medium | Head MIP |  |  |  | GEPACS | [9 empty] | manual
|  | Nec h070 :-D | [22 empty] | 0.6 | 0.6 | 201 | J40f medium | Head |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
| PreMonitoring |  | 120 | 120 | 175 | 100 |  |  | Off | 11 | Off | CARE Dose4D | 13.9 |  | 25% |  |  | 0.5 |  | 2 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | PreMonitoring | [22 empty] | 10 |  | 1 | H20s smooth | 400/40 Abd | [13 empty] | manual
| Monitoring |  | 120 | 120 | 175 | 100 |  |  | Off | 11 | Off | CARE Dose4D | 556.17 |  | 25% |  |  | 0.5 |  | 10 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | Monitoring | [22 empty] | 10 |  | 40 | H20s smooth | 400/40 Abd | [13 empty] | manual
| art Nec |  | 120 | 120 | 210 | 265 |  |  | Off | 11 | On | CARE Dose4D | 32.13 | 30s | 25% | 80 |  | 0.5 |  | 7 | 0.6 |  | 3 | 128x0.6mm | [15 empty] | Off
|  | art Nec h070 | [22 empty] | 3 | 2 | 61 | J40f medium | Head |  |  |  | GEPACS | [9 empty] | manual
|  | art Nec cor mip | [22 empty] | 7 | 2.5 | 101 | J40f medium | Head MIP |  |  |  | GEPACS | [9 empty] | manual
|  | art Nec sag mip | [22 empty] | 7 | 2.5 | 101 | J40f medium | Head MIP |  |  |  | GEPACS | [9 empty] | manual
|  | art Nec mip | [22 empty] | 3 | 2 | 61 | J40f medium | Head MIP |  |  |  | GEPACS | [9 empty] | manual
|  | art Nec h070 :-D | [22 empty] | 0.6 | 0.6 | 201 | J40f medium | Head |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
| ven Nec |  | 120 | 120 | 210 | 265 |  |  | Off | 11 | On | CARE Dose4D | 32.13 | 30s | 25% | 80 |  | 0.5 |  | 30 | 0.6 |  | 3 | 128x0.6mm | [15 empty] | Off
|  | ven Nec h070 | [22 empty] | 3 | 2 | 61 | J40f medium | Head |  |  |  | GEPACS | [9 empty] | manual
|  | ven Nec cor mip | [22 empty] | 7 | 2.5 | 101 | J40f medium | Head MIP |  |  |  | GEPACS | [9 empty] | manual
|  | ven Nec sag mip | [22 empty] | 7 | 2.5 | 101 | J40f medium | Head MIP |  |  |  | GEPACS | [9 empty] | manual
|  | ven Nec mip | [22 empty] | 3 | 2 | 61 | J40f medium | Head MIP |  |  |  | GEPACS | [9 empty] | manual
|  | ven Nec h070 :-D | [22 empty] | 0.6 | 0.6 | 201 | J40f medium | Head |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
| 90s Nec |  | 120 | 120 | 210 | 265 |  |  | Off | 11 | On | CARE Dose4D | 32.13 | 30s | 25% | 80 |  | 0.5 |  | 2 | 0.6 |  | 3 | 128x0.6mm | [15 empty] | Off
|  | 90s Nec h070 | [22 empty] | 3 | 2 | 61 | J40f medium | Head |  |  |  | GEPACS | [9 empty] | manual
|  | 90s Nec cor mip | [22 empty] | 7 | 2.5 | 101 | J40f medium | Head MIP |  |  |  | GEPACS | [9 empty] | manual
|  | 90s Nec sag mip | [22 empty] | 7 | 2.5 | 101 | J40f medium | Head MIP |  |  |  | GEPACS | [9 empty] | manual
|  | 90s Nec mip | [22 empty] | 3 | 2 | 61 | J40f medium | Head MIP |  |  |  | GEPACS | [9 empty] | manual
|  | 90s Nec h070 :-D | [22 empty] | 0.6 | 0.6 | 201 | J40f medium | Head |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
1080_Fac_bon (Adult)
1080_Fac_bon (Adult) | Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS | ROEMIR001a | TRACKER
| Fac |  | 120 | 120 | 390 | 350 |  |  | Off | 3 | Off | CARE Dose4D | 53.55 | 30s | 25% | 80 |  | 1 |  | 2 | 0.6 |  | 3 | 128x0.6mm | [15 empty] | Off
|  | ven Fac bon h080 | [22 empty] | 3 | 2 | 61 | J70h very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | ven Fac cor bon | [22 empty] | 2 | 2 | 76 | J70h very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | ven Fac sag bon | [22 empty] | 2 | 2 | 76 | J70h very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | ven Fac | [22 empty] | 3 | 2 | 61 | J30s medium smooth | Head |  |  |  | GEPACS | [9 empty] | manual
|  | ven Fac bon h080 :-D | [22 empty] | 0.6 | 0.6 | 201 | J70h very sharp | Bone |  |  |  | ROUTER | GEPACS | [8 empty] | manual
1090_Fac_bon_wo_cst (Adult)
1090_Fac_bon_wo_cst (Adult) | Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS | ROEMIR001a | TRACKER
| Fac |  | 120 | 120 | 390 | 350 |  |  | Off | 3 | Off | CARE Dose4D | 53.55 | 30s | 25% | 80 |  | 1 |  | 2 | 0.6 |  | 3 | 128x0.6mm | [15 empty] | Off
|  | Fac bon h090 | [22 empty] | 3 | 2 | 61 | J70h very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Fac cor bon | [22 empty] | 2 | 2 | 76 | J70h very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Fac sag bon | [22 empty] | 2 | 2 | 76 | J70h very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Fac | [22 empty] | 3 | 2 | 61 | J30s medium smooth | Head |  |  |  | GEPACS | [9 empty] | manual
|  | Fac bon h090 :-D | [22 empty] | 0.6 | 0.6 | 201 | J70h very sharp | Bone |  |  |  | ROUTER | GEPACS | [8 empty] | manual
1100_Hea (Adult)
1100_Hea (Adult) | Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.26 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.26 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS |  | TRACKER
| 3m Hea |  | 120 | 120 | 390 | 290 |  |  | Off | 3 | Off | CARE Dose4D | 44.37 | 15s | 25% | 80 |  | 0.5 |  | 180 | 0.6 |  | 4 | 128x0.6mm | [15 empty] | Off
|  | 3m Hea h100 | [22 empty] | 4 | 4 | 31 | J40f medium | Head Soft |  |  | X-CARE | GEPACS | [9 empty] | manual
|  | 3m Hea bon | [22 empty] | 4 | 4 | 31 | H60f sharp FR | Bone |  |  | X-CARE | GEPACS | [9 empty] | manual
|  | 3m Hea cor | [22 empty] | 3 | 3 | 84 | J40f medium | Head Soft |  |  | X-CARE | GEPACS | [9 empty] | manual
|  | 3m Hea sag | [22 empty] | 3 | 3 | 84 | J40f medium | Head Soft |  |  | X-CARE | GEPACS | [9 empty] | manual
|  | 3m Hea h100 :-D | [22 empty] | 1 | 1 | 121 | J40f medium | Head Soft |  |  | X-CARE | ROUTER | GEPACS | [8 empty] | manual
1105_Hea_cis (Adult)
1105_Hea_cis (Adult) | Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS | ROEMIR001a | TRACKER
| Hea |  | 120 | 120 | 230 | 375 |  |  | Off | 3 | Off | CARE Dose4D | 82.2 | 30s | 25% | 80 |  | 1 |  | 4 | 0.8 |  | 0.6 | 16x0.3mm | [15 empty] | Off
|  | Hea h105 :-D | [22 empty] | 0.6 | 0.3 | 334 | U70u sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Hea bon | [22 empty] | 2 | 2 | 51 | U70u sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Hea h105 | [22 empty] | 0.6 | 0.3 | 334 | U30u medium smooth | Head |  |  |  | GEPACS | [9 empty] | manual
|  | Hea cor | [22 empty] | 1 | 0.8 | 276 | U40u medium | Head |  |  |  | GEPACS | [9 empty] | manual
|  | Hea sag | [22 empty] | 1 | 0.8 | 276 | U40u medium | Head |  |  |  | ROUTER | GEPACS | [8 empty] | manual
| cis Hea |  | 120 | 120 | 230 | 375 |  |  | Off | 3 | Off | CARE Dose4D | 82.2 | 30s | 25% | 80 |  | 1 |  | 2 | 0.8 |  | 0.6 | 16x0.3mm | [15 empty] | Off
|  | cis Hea bon :-D | [22 empty] | 0.6 | 0.3 | 334 | U70u sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | cis Hea h105 | [22 empty] | 2 | 2 | 51 | U30u medium smooth | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | cis Hea h105 | [22 empty] | 0.6 | 0.3 | 334 | U30u medium smooth | Head |  |  |  | GEPACS | [9 empty] | manual
|  | cis Hea cor | [22 empty] | 1 | 0.8 | 276 | U40u medium | Head |  |  |  | GEPACS | [9 empty] | manual
|  | cis Hea sag | [22 empty] | 1 | 0.8 | 276 | U40u medium | Head |  |  |  | ROUTER | GEPACS | [8 empty] | manual
1130_Hea_sth (Adult)
1130_Hea_sth (Adult) | Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS |  | TRACKER
| 3m Hea |  | 120 | 120 | 390 | 350 |  |  | Off | 3 | Off | CARE Dose4D | 53.55 | 30s | 25% | 80 |  | 0.5 |  | 2 | 0.6 |  | 1 | 128x0.6mm | [15 empty] | Off
|  | 3m Hea h130 | [22 empty] | 1 | 1 | 121 | J37f medium smooth | Head Soft |  |  |  | GEPACS | COMPASS | [8 empty] | manual
|  | 3m Hea bon | [22 empty] | 1 | 1 | 121 | H60f sharp FR | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | 3m Hea h130 | [22 empty] | 1 | 1 | 121 | J37f medium smooth | Head Soft |  |  |  | GEPACS | [9 empty] | manual
|  | 3m Hea bon :-D | [22 empty] | 0.6 | 0.6 | 201 | H60f sharp FR | Bone |  |  |  | ROUTER | GEPACS | [8 empty] | manual
1132_Hea_sth_wo_cst (Adult)
1132_Hea_sth_wo_cst (Adult) | Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS |  | TRACKER
| Hea |  | 120 | 120 | 390 | 350 |  |  | Off | 3 | Off | CARE Dose4D | 53.55 | 30s | 25% | 80 |  | 0.5 |  | 2 | 0.6 |  | 1 | 128x0.6mm | [15 empty] | Off
|  | Hea h132 :-D | [22 empty] | 1 | 1 | 121 | J37f medium smooth | Head Soft |  |  |  | ROUTER | GEPACS | COMPASS | [7 empty] | manual
|  | Hea bon | [22 empty] | 1 | 1 | 121 | H60f sharp FR | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Hea h132 | [22 empty] | 1 | 1 | 121 | J37f medium smooth | Head Soft |  |  |  | GEPACS | [9 empty] | manual
|  | Hea bon | [22 empty] | 0.6 | 0.6 | 201 | H60f sharp FR | Bone |  |  |  | GEPACS | [9 empty] | manual
1138_Hea_ste_fra (Adult)
1138_Hea_ste_fra (Adult) | Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS |  | TRACKER
| 3m Hea |  | 120 | 120 | 390 | 350 |  |  | Off | 3 | Off | CARE Dose4D | 53.55 | 30s | 25% | 80 |  | 0.5 |  | 2 | 0.6 |  | 1 | 128x0.6mm | [15 empty] | Off
|  | 3m Hea h138 :-D | [22 empty] | 1 | 1 | 121 | J37f medium smooth | Head Soft |  |  |  | ROUTER | GEPACS | COMPASS | [7 empty] | manual
|  | 3m Hea bon | [22 empty] | 1 | 1 | 121 | H60f sharp FR | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | 3m Hea h138 | [22 empty] | 1 | 1 | 121 | J37f medium smooth | Head Soft |  |  |  | GEPACS | [9 empty] | manual
|  | 3m Hea bon | [22 empty] | 0.6 | 0.6 | 201 | H60f sharp FR | Bone |  |  |  | GEPACS | [9 empty] | manual
1140_Hea_ste_fra_wo_cst (Adult)
1140_Hea_ste_fra_wo_cst (Adult) | Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS |  | TRACKER
| Hea |  | 120 | 120 | 390 | 350 |  |  | Off | 3 | Off | CARE Dose4D | 53.55 | 30s | 25% | 80 |  | 0.5 |  | 2 | 0.6 |  | 1 | 128x0.6mm | [15 empty] | Off
|  | Hea h140 :-D | [22 empty] | 1 | 1 | 121 | J37f medium smooth | Head Soft |  |  |  | ROUTER | GEPACS | COMPASS | [7 empty] | manual
|  | Hea bon | [22 empty] | 1 | 1 | 121 | H60f sharp FR | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Hea h140 | [22 empty] | 1 | 1 | 121 | J37f medium smooth | Head Soft |  |  |  | GEPACS | [9 empty] | manual
|  | Hea bon | [22 empty] | 0.6 | 0.6 | 201 | H60f sharp FR | Bone |  |  |  | GEPACS | [9 empty] | manual
1150_Hea_wo_cst_N09_02A (Adult)
1150_Hea_wo_cst_N09_02A (Adult) | Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.26 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.26 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS |  | TRACKER
| Hea |  | 120 | 120 | 390 | 290 |  |  | Off | 3 | Off | CARE Dose4D | 44.37 | 15s | 25% | 80 |  | 0.5 |  | 2 | 0.6 |  | 4 | 128x0.6mm | [15 empty] | Off
|  | Hea h150 | [22 empty] | 4 | 4 | 31 | J40f medium | Head Soft |  |  | X-CARE | GEPACS | [9 empty] | manual
|  | Hea bon | [22 empty] | 4 | 4 | 31 | H60f sharp FR | Bone |  |  | X-CARE | GEPACS | [9 empty] | manual
|  | Hea cor | [22 empty] | 3 | 3 | 84 | J40f medium | Head Soft |  |  | X-CARE | GEPACS | [9 empty] | manual
|  | Hea sag | [22 empty] | 3 | 3 | 84 | J40f medium | Head Soft |  |  | X-CARE | GEPACS | [9 empty] | manual
|  | Hea h150 :-D | [22 empty] | 1 | 1 | 121 | J40f medium | Head Soft |  |  | X-CARE | ROUTER | GEPACS | [8 empty] | manual
1163_HeaCsp (Adult)
1163_HeaCsp (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| ven Csp |  | 120 | 140 | 280 | 280 |  |  | Off | 3 | On | CARE Dose4D | 27.54 | 30s | 25% | 50 |  | 1 |  | 70 | 0.9 |  | 2 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven Csp bon s020 | [22 empty] | 2 | 2 | 81 | B70s very sharp | Bone |  |  |  | GE_PACS | [9 empty] | manual
|  | ven Csp cor bon | [22 empty] | 2 | 2 | 76 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | ven Csp sag bon | [22 empty] | 2 | 2 | 76 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | ven Csp | [22 empty] | 2 | 2 | 81 | B30s medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | ven Csp sag | [22 empty] | 2 | 2 | 76 | B30s medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | ven Csp bon s020 :-D | [22 empty] | 0.6 | 0.6 | 267 | B70s very sharp | Bone |  |  |  | ROUTER | GEPACS | [8 empty] | manual
| 3m Hea |  | 120 | 120 | 390 | 290 |  |  | Off | 3 | Off | CARE Dose4D | 44.37 | 15s | 25% | 80 |  | 0.5 |  | 2 | 0.6 |  | 4 | 128x0.6mm | [15 empty] | Off
|  | 3m Hea h100 | [22 empty] | 4 | 4 | 31 | J40f medium | Head Soft |  |  | X-CARE | GEPACS | [9 empty] | manual
|  | 3m Hea bon | [22 empty] | 4 | 4 | 31 | H60f sharp FR | Bone |  |  | X-CARE | GEPACS | [9 empty] | manual
|  | 3m Hea cor | [22 empty] | 3 | 3 | 84 | J40f medium | Head Soft |  |  | X-CARE | GEPACS | [9 empty] | manual
|  | 3m Hea sag | [22 empty] | 3 | 3 | 84 | J40f medium | Head Soft |  |  | X-CARE | GEPACS | [9 empty] | manual
|  | 3m Hea h100 :-D | [22 empty] | 1 | 1 | 121 | J40f medium | Head Soft |  |  | X-CARE | ROUTER | GEPACS | [8 empty] | manual
1167_HeaCsp_wo_cst (Adult)
1167_HeaCsp_wo_cst (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Csp |  | 120 | 140 | 280 | 280 |  |  | Off | 3 | On | CARE Dose4D | 27.54 | 30s | 25% | 50 |  | 1 |  | 3 | 0.9 |  | 2 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | Csp bon s040 | [22 empty] | 2 | 2 | 81 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Csp cor bon | [22 empty] | 2 | 2 | 76 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Csp sag bon | [22 empty] | 2 | 2 | 76 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Csp sag | [22 empty] | 2 | 2 | 76 | B30s medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | Csp | [22 empty] | 2 | 2 | 81 | B30s medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | Csp bon s040 :-D | [22 empty] | 0.6 | 0.6 | 267 | B70s very sharp | Bone |  |  |  | ROUTER | GEPACS | [8 empty] | manual
| Hea |  | 120 | 120 | 390 | 290 |  |  | Off | 3 | Off | CARE Dose4D | 44.37 | 15s | 25% | 80 |  | 0.5 |  | 2 | 0.6 |  | 4 | 128x0.6mm | [15 empty] | Off
|  | Hea h150 | [22 empty] | 4 | 4 | 31 | J37f medium smooth | Head Soft |  |  | X-CARE | GEPACS | [9 empty] | manual
|  | Hea bon | [22 empty] | 4 | 4 | 31 | H60f sharp FR | Bone |  |  | X-CARE | GEPACS | [9 empty] | manual
|  | Hea cor | [22 empty] | 3 | 3 | 84 | J40f medium | Head Soft |  |  | X-CARE | GEPACS | [9 empty] | manual
|  | Hea sag | [22 empty] | 3 | 3 | 84 | J40f medium | Head Soft |  |  | X-CARE | GEPACS | [9 empty] | manual
|  | Hea h150 :-D | [22 empty] | 1 | 1 | 251 | J40f medium | Head Soft |  |  | X-CARE | ROUTER | GEPACS | [8 empty] | manual
1170_HeaFac_bon (Adult)
1170_HeaFac_bon (Adult) | Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS | ROEMIR001a | TRACKER
| Fac |  | 120 | 120 | 390 | 350 |  |  | Off | 3 | Off | CARE Dose4D | 53.55 | 30s | 25% | 80 |  | 1 |  | 2 | 0.6 |  | 3 | 128x0.6mm | [15 empty] | Off
|  | ven Fac bon h080 | [22 empty] | 3 | 2 | 61 | H60s sharp FR | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | ven Fac cor bon | [22 empty] | 2 | 2 | 76 | H60s sharp FR | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | ven Fac sag bon | [22 empty] | 2 | 2 | 76 | H60s sharp FR | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | ven Fac | [22 empty] | 3 | 2 | 61 | J30s medium smooth | Head |  |  |  | GEPACS | [9 empty] | manual
|  | ven Fac bon h080 :-D | [22 empty] | 0.6 | 0.6 | 201 | H60s sharp FR | Bone |  |  |  | ROUTER | GEPACS | [8 empty] | manual
| 3m Hea |  | 120 | 120 | 390 | 290 |  |  | Off | 3 | Off | CARE Dose4D | 44.37 | 15s | 25% | 80 |  | 0.5 |  | 2 | 0.6 |  | 4 | 128x0.6mm | [15 empty] | Off
|  | 3m Hea h100 | [22 empty] | 4 | 4 | 31 | J40f medium | Head Soft |  |  | X-CARE | GEPACS | [9 empty] | manual
|  | 3m Hea bon | [22 empty] | 4 | 4 | 31 | H60f sharp FR | Bone |  |  | X-CARE | GEPACS | [9 empty] | manual
|  | 3m Hea cor | [22 empty] | 3 | 3 | 84 | J40f medium | Head Soft |  |  | X-CARE | GEPACS | [9 empty] | manual
|  | 3m Hea sag | [22 empty] | 3 | 3 | 84 | J40f medium | Head Soft |  |  | X-CARE | GEPACS | [9 empty] | manual
|  | 3m Hea h100 :-D | [22 empty] | 1 | 1 | 121 | J40f medium | Head Soft |  |  | X-CARE | ROUTER | GEPACS | [8 empty] | manual
1180_HeaFac_bon_wo_cst (Adult)
1180_HeaFac_bon_wo_cst (Adult) | Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS | ROEMIR001a | TRACKER
| Fac |  | 120 | 120 | 390 | 350 |  |  | Off | 3 | Off | CARE Dose4D | 53.55 | 30s | 25% | 80 |  | 1 |  | 2 | 0.6 |  | 3 | 128x0.6mm | [15 empty] | Off
|  | Fac bon h090 | [22 empty] | 3 | 2 | 61 | H60s sharp FR | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Fac cor bon | [22 empty] | 2 | 2 | 76 | H60s sharp FR | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Fac sag bon | [22 empty] | 2 | 2 | 76 | H60s sharp FR | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Fac | [22 empty] | 3 | 2 | 61 | J37s medium smooth | Head |  |  |  | GEPACS | [9 empty] | manual
|  | Fac bon h090 :-D | [22 empty] | 0.6 | 0.6 | 201 | H60s sharp FR | Bone |  |  |  | ROUTER | GEPACS | [8 empty] | manual
| Hea |  | 120 | 120 | 390 | 290 |  |  | Off | 3 | Off | CARE Dose4D | 44.37 | 15s | 25% | 80 |  | 0.5 |  | 2 | 0.6 |  | 4 | 128x0.6mm | [15 empty] | Off
|  | Hea h150 | [22 empty] | 4 | 4 | 31 | J37f medium smooth | Head Soft |  |  | X-CARE | GEPACS | [9 empty] | manual
|  | Hea bon | [22 empty] | 4 | 4 | 31 | H60f sharp FR | Bone |  |  | X-CARE | GEPACS | [9 empty] | manual
|  | Hea cor | [22 empty] | 3 | 3 | 84 | J40f medium | Head Soft |  |  | X-CARE | GEPACS | [9 empty] | manual
|  | Hea sag | [22 empty] | 3 | 3 | 84 | J40f medium | Head Soft |  |  | X-CARE | GEPACS | [9 empty] | manual
|  | Hea h150 :-D | [22 empty] | 1 | 1 | 251 | J40f medium | Head Soft |  |  | X-CARE | ROUTER | GEPACS | [8 empty] | manual
1183_HeaNec (Adult)
1183_HeaNec (Adult) | Topogram | Topogram |  | 100 |  | 137 |  |  |  | 11 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 137 |  |  |  | 11 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS | ROEMIR001a | TRACKER
| ven Nec |  | 120 | 120 | 240 | 265 |  |  | Off | 11 | On | CARE Dose4D | 36.76 | 30s | 25% | 80 |  | 0.5 |  | 90 | 0.7 |  | 3 | 128x0.6mm | [15 empty] | Off
|  | ven Nec h260 | [22 empty] | 3 | 2 | 61 | J40f medium | Head |  |  |  | GEPACS | [9 empty] | manual
|  | ven Nec cor | [22 empty] | 2 | 2 | 126 | J40f medium | Head |  |  |  | GEPACS | [9 empty] | manual
|  | ven Nec sag | [22 empty] | 2 | 2 | 126 | J40f medium | Head |  |  |  | GEPACS | [9 empty] | manual
|  | ven Nec h260 :-D | [22 empty] | 0.6 | 0.6 | 201 | J40f medium | Head |  |  |  | ROUTER | GEPACS | MCH_HDCT_AW | [7 empty] | manual
| 3m Hea |  | 120 | 120 | 390 | 290 |  |  | Off | 3 | Off | CARE Dose4D | 44.37 | 15s | 25% | 80 |  | 0.5 |  | 85 | 0.6 |  | 4 | 128x0.6mm | [15 empty] | Off
|  | 3m Hea h100 | [22 empty] | 4 | 4 | 31 | J40f medium | Head Soft |  |  | X-CARE | GEPACS | [9 empty] | manual
|  | 3m Hea bon | [22 empty] | 4 | 4 | 31 | H60f sharp FR | Bone |  |  | X-CARE | GEPACS | [9 empty] | manual
|  | 3m Hea cor | [22 empty] | 3 | 3 | 84 | J40f medium | Head Soft |  |  | X-CARE | GEPACS | [9 empty] | manual
|  | 3m Hea sag | [22 empty] | 3 | 3 | 84 | J40f medium | Head Soft |  |  | X-CARE | GEPACS | [9 empty] | manual
|  | 3m Hea h100 :-D | [22 empty] | 1 | 1 | 121 | J40f medium | Head Soft |  |  | X-CARE | ROUTER | GEPACS | [8 empty] | manual
1187_HeaNec_wo_cst (Adult)
1187_HeaNec_wo_cst (Adult) | Topogram | Topogram |  | 100 |  | 137 |  |  |  | 11 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 137 |  |  |  | 11 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS | ROEMIR001a | TRACKER
| Nec |  | 120 | 120 | 240 | 265 |  |  | Off | 11 | On | CARE Dose4D | 36.76 | 30s | 25% | 80 |  | 0.5 |  | 3 | 0.7 |  | 3 | 128x0.6mm | [15 empty] | Off
|  | Nec h270 | [22 empty] | 3 | 2 | 61 | J40f medium | Head |  |  |  | GEPACS | [9 empty] | manual
|  | Nec cor | [22 empty] | 2 | 2 | 126 | J40f medium | Head |  |  |  | GEPACS | [9 empty] | manual
|  | Nec sag | [22 empty] | 2 | 2 | 126 | J40f medium | Head |  |  |  | GEPACS | [9 empty] | manual
|  | Nec h270 :-D | [22 empty] | 0.6 | 0.6 | 201 | J40f medium | Head |  |  |  | ROUTER | GEPACS | MCH_HDCT_AW | [7 empty] | manual
| Hea |  | 120 | 120 | 390 | 290 |  |  | Off | 3 | Off | CARE Dose4D | 44.37 | 15s | 25% | 80 |  | 0.5 |  | 2 | 0.6 |  | 4 | 128x0.6mm | [15 empty] | Off
|  | Hea h150 | [22 empty] | 4 | 4 | 31 | J37f medium smooth | Head Soft |  |  | X-CARE | GEPACS | [9 empty] | manual
|  | Hea bon | [22 empty] | 4 | 4 | 31 | H60f sharp FR | Bone |  |  | X-CARE | GEPACS | [9 empty] | manual
|  | Hea cor | [22 empty] | 3 | 3 | 84 | J40f medium | Head Soft |  |  | X-CARE | GEPACS | [9 empty] | manual
|  | Hea sag | [22 empty] | 3 | 3 | 84 | J40f medium | Head Soft |  |  | X-CARE | GEPACS | [9 empty] | manual
|  | Hea h150 :-D | [22 empty] | 1 | 1 | 251 | J40f medium | Head Soft |  |  | X-CARE | ROUTER | GEPACS | [8 empty] | manual
1190_HeaOrb (Adult)
1190_HeaOrb (Adult) | Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS | ROEMIR001a | TRACKER
| ven Orb |  | 120 | 120 | 390 | 350 |  |  | Off | 3 | Off | CARE Dose4D | 53.55 | 30s | 25% | 80 |  | 1 |  | 2 | 0.6 |  | 3 | 128x0.6mm | [15 empty] | Off
|  | ven Orb h300 | [22 empty] | 3 | 2 | 61 | J40s medium | Head |  |  |  | GE_PACS | [9 empty] | manual
|  | ven Orb cor | [22 empty] | 2 | 2 | 76 | J40s medium | Head |  |  |  | GEPACS | [9 empty] | manual
|  | ven Orb sag | [22 empty] | 2 | 2 | 76 | J40s medium | Head |  |  |  | GEPACS | [9 empty] | manual
|  | ven Orb cor bon | [22 empty] | 2 | 2 | 91 | J70h very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | ven Orb h300 :-D | [22 empty] | 0.6 | 0.6 | 201 | J40s medium | Head |  |  |  | ROUTER | GEPACS | [8 empty] | manual
| 3m Hea |  | 120 | 120 | 390 | 290 |  |  | Off | 3 | Off | CARE Dose4D | 44.37 | 15s | 25% | 80 |  | 0.5 |  | 2 | 0.6 |  | 4 | 128x0.6mm | [15 empty] | Off
|  | 3m Hea h100 | [22 empty] | 4 | 4 | 31 | J40f medium | Head Soft |  |  | X-CARE | GEPACS | [9 empty] | manual
|  | 3m Hea bon | [22 empty] | 4 | 4 | 31 | H60f sharp FR | Bone |  |  | X-CARE | GEPACS | [9 empty] | manual
|  | 3m Hea cor | [22 empty] | 3 | 3 | 84 | J40f medium | Head Soft |  |  | X-CARE | GEPACS | [9 empty] | manual
|  | 3m Hea sag | [22 empty] | 3 | 3 | 84 | J40f medium | Head Soft |  |  | X-CARE | GEPACS | [9 empty] | manual
|  | 3m Hea h100 :-D | [22 empty] | 1 | 1 | 121 | J40f medium | Head Soft |  |  | X-CARE | ROUTER | GEPACS | [8 empty] | manual
1200_HeaOrb_wo_cst (Adult)
1200_HeaOrb_wo_cst (Adult) | Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS | ROEMIR001a | TRACKER
| Orb |  | 120 | 120 | 390 | 350 |  |  | Off | 3 | Off | CARE Dose4D | 53.55 | 30s | 25% | 80 |  | 1 |  | 2 | 0.6 |  | 3 | 128x0.6mm | [15 empty] | Off
|  | Orb h310 | [22 empty] | 3 | 2 | 61 | J40s medium | Head |  |  |  | GE_PACS | [9 empty] | manual
|  | Orb cor | [22 empty] | 2 | 2 | 76 | J40s medium | Head |  |  |  | GEPACS | [9 empty] | manual
|  | Orb sag | [22 empty] | 2 | 2 | 76 | J40s medium | Head |  |  |  | GEPACS | [9 empty] | manual
|  | Orb cor bon | [22 empty] | 2 | 2 | 91 | J70h very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Orb h310 :-D | [22 empty] | 0.6 | 0.6 | 201 | J40s medium | Head |  |  |  | ROUTER | GEPACS | [8 empty] | manual
| Hea |  | 120 | 120 | 390 | 290 |  |  | Off | 3 | Off | CARE Dose4D | 44.37 | 15s | 25% | 80 |  | 0.5 |  | 2 | 0.6 |  | 4 | 128x0.6mm | [15 empty] | Off
|  | Hea h150 | [22 empty] | 4 | 4 | 31 | J37f medium smooth | Head Soft |  |  | X-CARE | GEPACS | [9 empty] | manual
|  | Hea bon | [22 empty] | 4 | 4 | 31 | H60f sharp FR | Bone |  |  | X-CARE | GEPACS | [9 empty] | manual
|  | Hea cor | [22 empty] | 3 | 3 | 84 | J40f medium | Head Soft |  |  | X-CARE | GEPACS | [9 empty] | manual
|  | Hea sag | [22 empty] | 3 | 3 | 84 | J40f medium | Head Soft |  |  | X-CARE | GEPACS | [9 empty] | manual
|  | Hea h150 :-D | [22 empty] | 1 | 1 | 251 | J40f medium | Head Soft |  |  | X-CARE | ROUTER | GEPACS | [8 empty] | manual
1210_HeaSin_fus_nav (Adult)
1210_HeaSin_fus_nav (Adult) | Topogram | Topogram 0.6 T20s |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS |  | TRACKER
| Topogram | Topogram 0.6 T20s |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS |  | TRACKER
| ven Sin |  | 120 | 120 | 390 | 350 |  |  | Off | 3 | Off | CARE Dose4D | 53.49 | 30s | 25% | 80 |  | 1 |  | 2 | 0.8 |  | 0.6 | 128x0.6mm | [15 empty] | Off
|  | ven Sin FUSION | [22 empty] | 0.6 | 0.6 | 201 | J30s medium smooth | Soft Tissue |  |  |  | OR_MIDIA_POOL | ROUTER | GEPACS | [7 empty] | manual
|  | ven Sin bon h350 :-D | [22 empty] | 0.6 | 0.6 | 201 | J70h very sharp | Bone |  |  |  | ROUTER | GEPACS | [8 empty] | manual
|  | ven Sin cor bon | [22 empty] | 1 | 1 | 151 | J70h very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | ven Sin sag bon | [22 empty] | 1 | 1 | 221 | J70h very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
| 3m Hea |  | 120 | 120 | 390 | 290 |  |  | Off | 3 | Off | CARE Dose4D | 44.37 | 15s | 25% | 80 |  | 0.5 |  | 2 | 0.6 |  | 4 | 128x0.6mm | [15 empty] | Off
|  | 3m Hea h100 | [22 empty] | 4 | 4 | 31 | J40f medium | Head Soft |  |  | X-CARE | GEPACS | [9 empty] | manual
|  | 3m Hea bon | [22 empty] | 4 | 4 | 31 | H60f sharp FR | Bone |  |  | X-CARE | GEPACS | [9 empty] | manual
|  | 3m Hea cor | [22 empty] | 3 | 3 | 84 | J40f medium | Head Soft |  |  | X-CARE | GEPACS | [9 empty] | manual
|  | 3m Hea sag | [22 empty] | 3 | 3 | 84 | J40f medium | Head Soft |  |  | X-CARE | GEPACS | [9 empty] | manual
|  | 3m Hea h100 :-D | [22 empty] | 1 | 1 | 121 | J40f medium | Head Soft |  |  | X-CARE | ROUTER | GEPACS | [8 empty] | manual
1220_HeaSin_fus_nav_wo_cst (Adult)
1220_HeaSin_fus_nav_wo_cst (Adult) | Topogram | Topogram 0.6 T20s |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS |  | TRACKER
| Topogram | Topogram 0.6 T20s |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS |  | TRACKER
| Sin |  | 120 | 120 | 390 | 175 |  |  | Off | 3 | Off | CARE Dose4D | 26.74 | 30s | 25% | 80 |  | 1 |  | 2 | 0.8 |  | 0.6 | 128x0.6mm | [15 empty] | Off
|  | Sin FUSION | [22 empty] | 0.6 | 0.6 | 201 | J30s medium smooth | Soft Tissue |  |  |  | OR_MIDIA_POOL | ROUTER | GEPACS | [7 empty] | manual
|  | Sin bon h360 :-D | [22 empty] | 0.6 | 0.6 | 201 | J70h very sharp | Bone |  |  |  | ROUTER | GEPACS | [8 empty] | manual
|  | Sin cor bon | [22 empty] | 1 | 1 | 151 | J70h very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Sin sag bon | [22 empty] | 1 | 1 | 221 | J70h very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
| Hea |  | 120 | 120 | 390 | 290 |  |  | Off | 3 | Off | CARE Dose4D | 44.37 | 15s | 25% | 80 |  | 0.5 |  | 2 | 0.6 |  | 4 | 128x0.6mm | [15 empty] | Off
|  | Hea h150 | [22 empty] | 4 | 4 | 31 | J37f medium smooth | Head Soft |  |  | X-CARE | GEPACS | [9 empty] | manual
|  | Hea bon | [22 empty] | 4 | 4 | 31 | H60f sharp FR | Bone |  |  | X-CARE | GEPACS | [9 empty] | manual
|  | Hea cor | [22 empty] | 3 | 3 | 84 | J40f medium | Head Soft |  |  | X-CARE | GEPACS | [9 empty] | manual
|  | Hea sag | [22 empty] | 3 | 3 | 84 | J40f medium | Head Soft |  |  | X-CARE | GEPACS | [9 empty] | manual
|  | Hea h150 :-D | [22 empty] | 1 | 1 | 251 | J40f medium | Head Soft |  |  | X-CARE | ROUTER | GEPACS | [8 empty] | manual
1230_Int_aud_can (Adult)
1230_Int_aud_can (Adult) | Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS |  | TRACKER
| ven Tem bon |  | 120 | 120 | 230 | 300 |  |  | Off | 3 | Off | CARE Dose4D | 65.76 | 30s | 25% | 80 |  | 1 |  | 2 | 0.8 |  | 2 | 16x0.3mm | [15 empty] | Off
|  | ven Tem h230 | [22 empty] | 2 | 2 | 51 | U40u medium | Head |  |  |  | GEPACS | [9 empty] | manual
|  | ven R Tem bon | [22 empty] | 0.5 | 0.3 | 334 | U70u sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | ven L Tem bon | [22 empty] | 0.5 | 0.3 | 334 | U70u sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | ven Tem bon h230 | [22 empty] | 1 | 1 | 101 | V80u very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | ven R Tem cor bon | [22 empty] | 0.5 | 0.3 | 334 | U70u sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | ven L Tem cor bon | [22 empty] | 0.5 | 0.3 | 334 | U70u sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | ven Tem bon h230 :-D | [22 empty] | 0.5 | 0.3 | 334 | V80u very sharp | Bone |  |  |  | ROUTER | GEPACS | [8 empty] | manual
1240_Man (Adult)
1240_Man (Adult) | Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS | ROEMIR001a | TRACKER
| ven Man |  | 120 | 120 | 390 | 350 |  |  | Off | 3 | Off | CARE Dose4D | 53.55 | 30s | 25% | 80 |  | 1 |  | 2 | 0.6 |  | 3 | 128x0.6mm | [15 empty] | Off
|  | ven Man bon h240 | [22 empty] | 3 | 2 | 61 | H60s sharp FR | Bone |  |  |  | GE_PACS | [9 empty] | manual
|  | ven Man cor bon | [22 empty] | 2 | 2 | 76 | H60s sharp FR | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | ven Man | [22 empty] | 3 | 2 | 61 | J37s medium smooth | Head |  |  |  | GEPACS | [9 empty] | manual
|  | ven Man bon h240 :-D | [22 empty] | 0.6 | 0.6 | 201 | H60s sharp FR | Bone |  |  |  | ROUTER | GEPACS | [8 empty] | manual
1250_Man_wo_cst (Adult)
1250_Man_wo_cst (Adult) | Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS | ROEMIR001a | TRACKER
| Man |  | 120 | 120 | 390 | 350 |  |  | Off | 3 | Off | CARE Dose4D | 53.55 | 30s | 25% | 80 |  | 1 |  | 2 | 0.6 |  | 3 | 128x0.6mm | [15 empty] | Off
|  | Man bon h250 | [22 empty] | 3 | 2 | 61 | H60s sharp FR | Bone |  |  |  | GE_PACS | [9 empty] | manual
|  | Man cor bon | [22 empty] | 2 | 2 | 76 | H60s sharp FR | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Man | [22 empty] | 3 | 2 | 61 | J37s medium smooth | Head |  |  |  | GEPACS | [9 empty] | manual
|  | Man bon h250 :-D | [22 empty] | 0.6 | 0.6 | 201 | H60s sharp FR | Bone |  |  |  | ROUTER | GEPACS | [8 empty] | manual
1260_Nec (Adult)
1260_Nec (Adult) | Topogram | Topogram |  | 100 |  | 137 |  |  |  | 11 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 137 |  |  |  | 11 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS | ROEMIR001a | TRACKER
| ven Nec |  | 120 | 120 | 240 | 265 |  |  | Off | 11 | On | CARE Dose4D | 36.76 | 30s | 25% | 80 |  | 0.5 |  | 90 | 0.7 |  | 3 | 128x0.6mm | [15 empty] | Off
|  | ven Nec h260 | [22 empty] | 3 | 2 | 61 | J40f medium | Head |  |  |  | GEPACS | [9 empty] | manual
|  | ven Nec cor | [22 empty] | 2 | 2 | 126 | J40f medium | Head |  |  |  | GEPACS | [9 empty] | manual
|  | ven Nec sag | [22 empty] | 2 | 2 | 126 | J40f medium | Head |  |  |  | GEPACS | [9 empty] | manual
|  | ven Nec h260 :-D | [22 empty] | 0.6 | 0.6 | 201 | J40f medium | Head |  |  |  | ROUTER | GEPACS | MCH_HDCT_AW | [7 empty] | manual
1270_Nec_wo_cst (Adult)
1270_Nec_wo_cst (Adult) | Topogram | Topogram |  | 100 |  | 137 |  |  |  | 11 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 137 |  |  |  | 11 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS | ROEMIR001a | TRACKER
| Nec |  | 120 | 120 | 240 | 265 |  |  | Off | 11 | On | CARE Dose4D | 36.76 | 30s | 25% | 80 |  | 0.5 |  | 3 | 0.7 |  | 3 | 128x0.6mm | [15 empty] | Off
|  | Nec h270 | [22 empty] | 3 | 2 | 61 | J40f medium | Head |  |  |  | GEPACS | [9 empty] | manual
|  | Nec cor | [22 empty] | 2 | 2 | 126 | J40f medium | Head |  |  |  | GEPACS | [9 empty] | manual
|  | Nec sag | [22 empty] | 2 | 2 | 126 | J40f medium | Head |  |  |  | GEPACS | [9 empty] | manual
|  | Nec h270 :-D | [22 empty] | 0.6 | 0.6 | 201 | J40f medium | Head |  |  |  | ROUTER | GEPACS | MCH_HDCT_AW | [7 empty] | manual
1275_NecChe_voc_cor_par (Adult)
1275_NecChe_voc_cor_par (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 11 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 11 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS |  | TRACKER
| ven NecChe |  | 120 | 120 | 240 | 265 |  |  | Off | 11 | On | CARE Dose4D | 36.76 | 30s | 25% | 80 |  | 0.5 |  | 90 | 0.7 |  | 3 | 128x0.6mm | [15 empty] | Off
|  | ven NecChe h275 | [22 empty] | 3 | 2 | 61 | J40f medium | Head |  |  |  | GEPACS | [9 empty] | manual
|  | ven NecChe cor | [22 empty] | 2 | 2 | 126 | J40f medium | Head |  |  |  | GEPACS | [9 empty] | manual
|  | ven NecChe sag | [22 empty] | 2 | 2 | 126 | J40f medium | Head |  |  |  | GEPACS | [9 empty] | manual
|  | ven NecChe h275 :-D | [22 empty] | 0.6 | 0.6 | 201 | J40f medium | Head |  |  |  | ROUTER | GEPACS | MCH_HDCT_AW | [7 empty] | manual
1280_NecChe_tra_ste (Adult)
1280_NecChe_tra_ste (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 11 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 11 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS | ROEMIR001a | TRACKER
| ven NecChe |  | 120 | 120 | 240 | 265 |  |  | Off | 11 | On | CARE Dose4D | 36.76 | 30s | 25% | 80 |  | 0.5 |  | 50 | 0.7 |  | 3 | 128x0.6mm | [15 empty] | Off
|  | ven NecChe h280 | [22 empty] | 3 | 2 | 61 | J30f medium smooth | Head |  |  |  | GEPACS | [9 empty] | manual
|  | ven NecChe cor | [22 empty] | 2 | 2 | 126 | J30f medium smooth | Head |  |  |  | GEPACS | [9 empty] | manual
|  | ven NecChe sag | [22 empty] | 2 | 2 | 126 | J30f medium smooth | Head |  |  |  | GEPACS | [9 empty] | manual
|  | ven NecChe h280 :-D | [22 empty] | 0.6 | 0.6 | 201 | J30f medium smooth | Head |  |  |  | ROUTER | GEPACS | MCH_HDCT_AW | [7 empty] | manual
1290_NecChe_tra_ste_wo_cst (Adult)
1290_NecChe_tra_ste_wo_cst (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 11 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 11 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS | ROEMIR001a | TRACKER
| NecChe |  | 120 | 120 | 240 | 265 |  |  | Off | 11 | On | CARE Dose4D | 36.76 | 30s | 25% | 80 |  | 0.5 |  | 3 | 0.7 |  | 3 | 128x0.6mm | [15 empty] | Off
|  | NecChe h290 | [22 empty] | 3 | 2 | 61 | J30f medium smooth | Head |  |  |  | GEPACS | [9 empty] | manual
|  | NecChe cor | [22 empty] | 2 | 2 | 126 | J30f medium smooth | Head |  |  |  | GEPACS | [9 empty] | manual
|  | NecChe sag | [22 empty] | 2 | 2 | 126 | J30f medium smooth | Head |  |  |  | GEPACS | [9 empty] | manual
|  | NecChe h290 :-D | [22 empty] | 0.6 | 0.6 | 201 | J30f medium smooth | Head |  |  |  | ROUTER | GEPACS | MCH_HDCT_AW | [7 empty] | manual
1295_NecChe_voc_cor_par_wo_cst (Adult)
1295_NecChe_voc_cor_par_wo_cst (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 11 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 11 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS |  | TRACKER
| NecChe |  | 120 | 120 | 240 | 265 |  |  | Off | 11 | On | CARE Dose4D | 36.76 | 30s | 25% | 80 |  | 0.5 |  | 3 | 0.7 |  | 3 | 128x0.6mm | [15 empty] | Off
|  | NecChe h295 | [22 empty] | 3 | 2 | 61 | J30f medium smooth | Head |  |  |  | GEPACS | [9 empty] | manual
|  | NecChe cor | [22 empty] | 2 | 2 | 126 | J30f medium smooth | Head |  |  |  | GEPACS | [9 empty] | manual
|  | NecChe sag | [22 empty] | 2 | 2 | 126 | J30f medium smooth | Head |  |  |  | GEPACS | [9 empty] | manual
|  | NecChe h295 :-D | [22 empty] | 0.6 | 0.6 | 201 | J30f medium smooth | Head |  |  |  | ROUTER | GEPACS | MCH_HDCT_AW | [7 empty] | manual
1300_Orb (Adult)
1300_Orb (Adult) | Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS | ROEMIR001a | TRACKER
| ven Orb |  | 120 | 120 | 390 | 350 |  |  | Off | 3 | Off | CARE Dose4D | 53.55 | 30s | 25% | 80 |  | 1 |  | 2 | 0.6 |  | 3 | 128x0.6mm | [15 empty] | Off
|  | ven Orb h300 | [22 empty] | 3 | 2 | 61 | J40s medium | Head |  |  |  | GE_PACS | [9 empty] | manual
|  | ven Orb cor | [22 empty] | 2 | 2 | 76 | J40s medium | Head |  |  |  | GEPACS | [9 empty] | manual
|  | ven Orb sag | [22 empty] | 2 | 2 | 76 | J40s medium | Head |  |  |  | GEPACS | [9 empty] | manual
|  | ven Orb cor bon | [22 empty] | 2 | 2 | 91 | J70h very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | ven Orb h300 :-D | [22 empty] | 0.6 | 0.6 | 201 | J40s medium | Head |  |  |  | ROUTER | GEPACS | [8 empty] | manual
1310_Orb_wo_cst (Adult)
1310_Orb_wo_cst (Adult) | Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS | ROEMIR001a | TRACKER
| Orb |  | 120 | 120 | 390 | 350 |  |  | Off | 3 | Off | CARE Dose4D | 53.55 | 30s | 25% | 80 |  | 1 |  | 2 | 0.6 |  | 3 | 128x0.6mm | [15 empty] | Off
|  | Orb h310 | [22 empty] | 3 | 2 | 61 | J40s medium | Head |  |  |  | GE_PACS | [9 empty] | manual
|  | Orb cor | [22 empty] | 2 | 2 | 76 | J40s medium | Head |  |  |  | GEPACS | [9 empty] | manual
|  | Orb sag | [22 empty] | 2 | 2 | 76 | J40s medium | Head |  |  |  | GEPACS | [9 empty] | manual
|  | Orb cor bon | [22 empty] | 2 | 2 | 91 | J70h very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Orb h310 :-D | [22 empty] | 0.6 | 0.6 | 201 | J40s medium | Head |  |  |  | ROUTER | GEPACS | [8 empty] | manual
1320_Sel (Adult)
1320_Sel (Adult) | Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS | ROEMIR001a | TRACKER
| ven Sel |  | 120 | 120 | 390 | 350 |  |  | Off | 3 | Off | CARE Dose4D | 53.55 | 30s | 25% | 80 |  | 1 |  | 2 | 0.6 |  | 3 | 128x0.6mm | [15 empty] | Off
|  | ven Sel h320 | [22 empty] | 3 | 2 | 61 | J40s medium | Head |  |  |  | GE_PACS | [9 empty] | manual
|  | ven Sel cor | [22 empty] | 2 | 2 | 76 | J40s medium | Head |  |  |  | GEPACS | [9 empty] | manual
|  | ven Sel sag | [22 empty] | 2 | 2 | 76 | J40s medium | Head |  |  |  | GEPACS | [9 empty] | manual
|  | ven Sel bon | [22 empty] | 3 | 2 | 61 | J70h very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | ven Sel h320 :-D | [22 empty] | 0.6 | 0.6 | 201 | J40s medium | Head |  |  |  | ROUTER | GEPACS | [8 empty] | manual
1330_Sel_wo_cst (Adult)
1330_Sel_wo_cst (Adult) | Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS | ROEMIR001a | TRACKER
| Sel |  | 120 | 120 | 390 | 350 |  |  | Off | 3 | Off | CARE Dose4D | 53.55 | 30s | 25% | 80 |  | 1 |  | 2 | 0.6 |  | 3 | 128x0.6mm | [15 empty] | Off
|  | Sel h330 | [22 empty] | 3 | 2 | 61 | J40s medium | Head |  |  |  | GE_PACS | [9 empty] | manual
|  | Sel cor | [22 empty] | 2 | 2 | 76 | J40s medium | Head |  |  |  | GEPACS | [9 empty] | manual
|  | Sel sag | [22 empty] | 2 | 2 | 76 | J40s medium | Head |  |  |  | GEPACS | [9 empty] | manual
|  | Sel bon | [22 empty] | 3 | 2 | 61 | J70h very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Sel h330 :-D | [22 empty] | 0.6 | 0.6 | 201 | J40s medium | Head |  |  |  | ROUTER | GEPACS | [8 empty] | manual
1350_Sin_fus_nav (Adult)
1350_Sin_fus_nav (Adult) | Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS |  | TRACKER
| ven Sin |  | 120 | 120 | 390 | 175 |  |  | Off | 3 | Off | CARE Dose4D | 26.74 | 30s | 25% | 80 |  | 1 |  | 2 | 0.8 |  | 0.6 | 128x0.6mm | [15 empty] | Off
|  | ven Sin FUSION | [22 empty] | 0.6 | 0.6 | 201 | J30s medium smooth | Soft Tissue |  |  |  | OR_MIDIA_POOL | ROUTER | GEPACS | [7 empty] | manual
|  | ven Sin bon h350 :-D | [22 empty] | 0.6 | 0.6 | 201 | J70h very sharp | Bone |  |  |  | ROUTER | GEPACS | [8 empty] | manual
|  | ven Sin cor bon | [22 empty] | 1 | 1 | 151 | J70h very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | ven Sin sag bon | [22 empty] | 1 | 1 | 221 | J70h very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
1360_Sin_fus_nav_wo_cst (Adult)
1360_Sin_fus_nav_wo_cst (Adult) | Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS |  | TRACKER
| Sin |  | 120 | 120 | 390 | 175 |  |  | Off | 3 | Off | CARE Dose4D | 26.74 | 30s | 25% | 80 |  | 1 |  | 2 | 0.8 |  | 0.6 | 128x0.6mm | [15 empty] | Off
|  | Sin FUSION | [22 empty] | 0.6 | 0.6 | 201 | J30s medium smooth | Soft Tissue |  |  |  | OR_MIDIA_POOL | ROUTER | GEPACS | [7 empty] | manual
|  | Sin bon h360 :-D | [22 empty] | 0.6 | 0.6 | 201 | J70h very sharp | Bone |  |  |  | ROUTER | GEPACS | [8 empty] | manual
|  | Sin cor bon | [22 empty] | 2 | 2 | 76 | J70h very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Sin sag bon | [22 empty] | 2 | 2 | 111 | J70h very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
1390_Tem_bon_sim_wo_cst (Adult)
1390_Tem_bon_sim_wo_cst (Adult) | Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS |  | TRACKER
| Tem bon |  | 120 | 120 | 230 | 300 |  |  | Off | 3 | Off | CARE Dose4D | 65.76 | 30s | 25% | 80 |  | 1 |  | 2 | 0.8 |  | 2 | 16x0.3mm | [15 empty] | Off
|  | Tem h390 | [22 empty] | 2 | 2 | 51 | U40u medium | Head |  |  |  | GEPACS | [9 empty] | manual
|  | R Tem bon | [22 empty] | 0.5 | 0.3 | 334 | U70u sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | L Tem bon | [22 empty] | 0.5 | 0.3 | 334 | U70u sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Tem bon h390 | [22 empty] | 1 | 1 | 101 | U70u sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | R Tem cor bon | [22 empty] | 0.5 | 0.3 | 267 | U70u sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | L Tem cor bon | [22 empty] | 0.5 | 0.3 | 267 | U70u sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Tem bon h390 :-D | [22 empty] | 0.5 | 0.3 | 334 | V80u very sharp | Bone |  |  |  | MCH3DLabAW | [9 empty] | manual
1400_Tem_bon_com_wo_cst (Adult)
1400_Tem_bon_com_wo_cst (Adult) | Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS |  | TRACKER
| Tem bon |  | 120 | 120 | 230 | 300 |  |  | Off | 3 | Off | CARE Dose4D | 65.76 | 30s | 25% | 80 |  | 1 |  | 2 | 0.8 |  | 2 | 16x0.3mm | [15 empty] | Off
|  | Tem h400 | [22 empty] | 2 | 2 | 51 | U40u medium | Head |  |  |  | GEPACS | [9 empty] | manual
|  | R Tem bon | [22 empty] | 0.5 | 0.3 | 334 | U70u sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | L Tem bon | [22 empty] | 0.5 | 0.3 | 334 | U70u sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Tem bon h400 | [22 empty] | 1 | 1 | 101 | U70u sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | R Tem cor bon | [22 empty] | 0.5 | 0.3 | 267 | U70u sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | L Tem cor bon | [22 empty] | 0.5 | 0.3 | 267 | U70u sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Tem bon h400 :-D | [22 empty] | 0.5 | 0.3 | 334 | V80u very sharp | Bone |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
1405_TMJ_con_wo_cst (Adult)
1405_TMJ_con_wo_cst (Adult) | Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS |  | TRACKER
| TMJ bon |  | 120 | 120 | 135 | 350 |  |  | Off | 3 | On | CARE Dose4D | 20.63 | 30s | 25% | 80 |  | 1 |  | 2 | 0.8 |  | 0.6 | 128x0.6mm | [15 empty] | Off
|  | TMJ bon h405 | [22 empty] | 0.6 | 0.6 | 201 | J70h very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | TMJ sag bon | [22 empty] | 3 | 3 | 67 | J70h very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | TMJ h405 :-D | [22 empty] | 0.6 | 0.6 | 334 | J40s medium | Head |  |  |  | ROUTER | GEPACS | [8 empty] | manual
Dental (Adult)
Dental (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.29 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS | ROEMIR001a | TRACKER
| Dental |  | 120 | 120 | 90 | 90 |  |  | On | 5 | On | CARE Dose4D | 12.69 | 7s | 25% |  |  | 1 |  | 2 | 0.8 |  | 0.75 | 128x0.6mm | [15 empty] | Off
|  | Dental 0.75 H60s | [22 empty] | 0.75 | 0.5 | 101 | H60s sharp FR | Dental | [13 empty] | manual
HeadIntervention (Adult)
HeadIntervention (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.29 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS | ROEMIR001a | TRACKER
| i-Spiral |  | 120 | 120 | 420 | 420 |  |  | On | 3 | On | CARE Dose4D | 60.49 | 5s | 25% |  |  | 0.5 |  | 2 | 0.55 |  | 5 | 32x1.2mm | [15 empty] | Off
|  | i-Spiral 5.0 H32f | [22 empty] | 5 | 3.5 | 35 | H32f med. smooth FR | Base Orbita | [13 empty] | manual
HeadNeuro (Adult)
HeadNeuro (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.29 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS | ROEMIR001a | TRACKER
| Head |  | 120 | 120 | 320 | 320 |  |  | On | 3 | On | CARE Dose4D | 58.93 | 77s | 25% |  |  | 1 |  | 2 | 0.55 |  | 5 | 20x0.6mm | [15 empty] | Off
|  | Head 5.0 H31s | [22 empty] | 5 | 5 | 25 | H31s medium smooth + | Base Orbita | [13 empty] | manual
HeadNeuroSeq (Adult)
HeadNeuroSeq (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.29 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS | ROEMIR001a | TRACKER
| HeadSeq |  | 120 | 120 | 320 | 320 |  |  | On | 3 | On | CARE Dose4D | 58.7 |  | 25% |  |  | 1 |  | 2 | 0 | 14 | 4.8 | 12x1.2mm | [15 empty] | Off
|  | HeadSeq 4.8 H31s | [22 empty] | 4.8 |  | 27 | H31s medium smooth + | Base Orbita | [13 empty] | manual
HeadNeuro_XCARE (Adult)
HeadNeuro_XCARE (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.29 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS | ROEMIR001a | TRACKER
| Head |  | 120 | 120 | 320 | 320 |  |  | On | 3 | On | CARE Dose4D | 60.06 | 40s | 25% |  |  | 1 |  | 2 | 0.55 |  | 5 | 40x0.6mm | [15 empty] | Off
|  | Head 5.0 H31s | [22 empty] | 5 | 5 | 25 | H31s medium smooth + | Base Orbita |  |  | X-CARE | [10 empty] | manual
HeadRoutine (Adult)
HeadRoutine (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.29 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS | ROEMIR001a | TRACKER
| Head |  | 120 | 120 | 390 | 390 |  |  | On | 3 | On | CARE Dose4D | 59.76 | 15s | 25% |  |  | 1 |  | 2 | 0.55 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | Head 5.0 H31s | [22 empty] | 5 | 5 | 25 | H31s medium smooth + | Base Orbita | [13 empty] | manual
HeadRoutineSeq (Adult)
HeadRoutineSeq (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.29 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS | ROEMIR001a | TRACKER
| HeadSeq |  | 120 | 120 | 340 | 340 |  |  | On | 3 | On | CARE Dose4D | 58.91 |  | 25% |  |  | 1 |  | 2 | 0 | 34 | 5 | 32x1.2mm | [15 empty] | Off
|  | HeadSeq 5.0 H31s | [22 empty] | 5 |  | 27 | H31s medium smooth + | Base Orbita | [13 empty] | manual
HeadSeq_XCARE (Adult)
HeadSeq_XCARE (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.29 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS | ROEMIR001a | TRACKER
| HeadSeq |  | 120 | 120 | 340 | 340 |  |  | On | 3 | On | CARE Dose4D | 58.91 |  | 25% |  |  | 1 |  | 2 | 0 | 34 | 5 | 32x1.2mm | [15 empty] | Off
|  | HeadSeq 5.0 H31s | [22 empty] | 5 |  | 27 | H31s medium smooth + | Base Orbita |  |  | X-CARE | [10 empty] | manual
Head_IR (Adult)
Head_IR (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.29 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head
| Head |  | 120 | 120 | 273 | 273 |  |  | On | 3 | On | CARE Dose4D | 41.97 | 15s | 25% |  |  | 1 |  | 2 | 0.55 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | Head 5.0 J30s 3 | [22 empty] | 5 | 5 | 25 | J30s medium smooth | Base Orbita | [13 empty] | manual
Head_XCARE (Adult)
Head_XCARE (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.29 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS | ROEMIR001a | TRACKER
| Head |  | 120 | 120 | 390 | 390 |  |  | On | 3 | On | CARE Dose4D | 59.76 | 15s | 25% |  |  | 1 |  | 2 | 0.55 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | Head 5.0 H31s | [22 empty] | 5 | 5 | 25 | H31s medium smooth + | Base Orbita |  |  | X-CARE | [10 empty] | manual
Head_XCARE_IR (Adult)
Head_XCARE_IR (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.29 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head
| Head |  | 120 | 120 | 273 | 273 |  |  | On | 3 | On | CARE Dose4D | 41.97 | 15s | 25% |  |  | 1 |  | 2 | 0.55 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | Head 5.0 J30s 3 | [22 empty] | 5 | 5 | 25 | J30s medium smooth | Base Orbita |  |  | X-CARE | [10 empty] | manual
InnerEar (Adult)
InnerEar (Adult) | Topogram | Topogram 0.6 T80f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.29 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Head |  |  |  | GEPACS | ROEMIR001a | TRACKER
| InnerEar |  | 120 | 120 | 180 | 180 |  |  | On | 5 | On | CARE Dose4D | 27.51 | 7s | 25% |  |  | 1 |  | 2 | 0.8 |  | 2 | 128x0.6mm | [15 empty] | Off
|  | InnerEar 2.0 H60s | [22 empty] | 2 | 2 | 21 | H60s sharp FR | Inner Ear | [13 empty] | manual
|  | R InnerEar 0.6 H60s | [22 empty] | 0.6 | 0.4 | 101 | H60s sharp FR | Inner Ear | [13 empty] | manual
|  | L InnerEar 0.6 H60s | [22 empty] | 0.6 | 0.4 | 101 | H60s sharp FR | Inner Ear | [13 empty] | manual
InnerEarUHR (Adult)
InnerEarUHR (Adult) | Topogram | Topogram 0.6 T80f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.29 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Head |  |  |  | GEPACS | ROEMIR001a | TRACKER
| InnerEarUHR |  | 120 | 120 | 180 | 180 |  |  | On | 5 | On | CARE Dose4D | 39.82 | 24s | 25% |  |  | 1 |  | 2 | 0.85 |  | 2 | 16x0.6mm | [15 empty] | Off
|  | InnerEarUHR 2.0 U75u | [22 empty] | 2 | 2 | 21 | U75u very sharp ASA | Inner Ear | [13 empty] | manual
|  | R InnerEarUHR 0.6 U75u | [22 empty] | 0.6 | 0.4 | 101 | U75u very sharp ASA | Inner Ear | [13 empty] | manual
|  | L InnerEarUHR 0.6 U75u | [22 empty] | 0.6 | 0.4 | 101 | U75u very sharp ASA | Inner Ear | [13 empty] | manual
InnerEarUHRSeq (Adult)
InnerEarUHRSeq (Adult) | Topogram | Topogram 0.6 T80f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.29 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Head |  |  |  | GEPACS | ROEMIR001a | TRACKER
| InnEarUHRSeq |  | 120 | 120 | 180 | 180 |  |  | On | 5 | On | CARE Dose4D | 41.65 |  | 25% |  |  | 1 |  | 2 | 0 | 4 | 2 | 16x0.6mm | [15 empty] | Off
|  | InnEarUHRSeq 2.0 U75u | [22 empty] | 2 |  | 20 | U75u very sharp ASA | Inner Ear | [13 empty] | manual
|  | InnEarUHRSeq 0.6 U75u | [22 empty] | 0.6 |  | 101 | U75u very sharp ASA | Inner Ear | [13 empty] | manual
InnerEarZUHR (Adult)
InnerEarZUHR (Adult) | Topogram | Topogram 0.6 T80f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.29 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Head |  |  |  | GEPACS | ROEMIR001a | TRACKER
| InnerEarUHR |  | 120 | 120 | 230 | 230 |  |  | On | 5 | On | CARE Dose4D | 51.02 | 24s | 25% |  |  | 1 |  | 2 | 0.85 |  | 2 | 16x0.3mm | [15 empty] | Off
|  | InnerEarUHR 2.0 U75u | [22 empty] | 2 | 2 | 21 | U75u very sharp ASA | Inner Ear | [13 empty] | manual
|  | R InnerEarUHR 0.4 U75u | [22 empty] | 0.4 | 0.2 | 201 | U75u very sharp ASA | Inner Ear | [13 empty] | manual
|  | L InnerEarUHR 0.4 U75u | [22 empty] | 0.4 | 0.2 | 201 | U75u very sharp ASA | Inner Ear | [13 empty] | manual
InnerEarZUHR_IR (Adult)
InnerEarZUHR_IR (Adult) | Topogram | Topogram 0.6 T80f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.29 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Head
| InnerEarUHR |  | 120 | 120 | 161 | 161 |  |  | Semi | 5 | On | CARE Dose4D | 35.66 | 24s | 25% |  |  | 1 |  | 2 | 0.85 |  | 2 | 16x0.3mm | [15 empty] | Off
|  | InnerEarUHR 2.0 V80u 3 | [22 empty] | 2 | 2 | 21 | V80u very sharp | Inner Ear | [13 empty] | manual
|  | R InnerEarUHR 0.4 V80u 3 | [22 empty] | 0.4 | 0.2 | 201 | V80u very sharp | Inner Ear | [13 empty] | manual
|  | L InnerEarUHR 0.4 V80u 3 | [22 empty] | 0.4 | 0.2 | 201 | V80u very sharp | Inner Ear | [13 empty] | manual
Orbit (Adult)
Orbit (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.29 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS | ROEMIR001a | TRACKER
| Orbit |  | 120 | 120 | 125 | 125 |  |  | On | 5 | On | CARE Dose4D | 17.62 | 5s | 25% |  |  | 1 |  | 2 | 0.8 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | Orbit 5.0 H60s | [22 empty] | 5 | 5 | 11 | H60s sharp FR | Base Orbita | [13 empty] | manual
|  | Orbit 5.0 H30s | [22 empty] | 5 | 5 | 11 | H30s medium smooth | Base Orbita | [13 empty] | manual
|  | Orbit 0.6 H60s | [22 empty] | 0.6 | 0.4 | 126 | H60s sharp FR | Base Orbita | [13 empty] | manual
Sinus (Adult)
Sinus (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.29 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS | ROEMIR001a | TRACKER
| Sinus |  | 120 | 120 | 50 | 50 |  |  | On | 5 | On | CARE Dose4D | 7.05 | 6s | 25% |  |  | 1 |  | 2 | 0.8 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | Sinus 5.0 H60s | [22 empty] | 5 | 5 | 17 | H60s sharp FR | Sinuses | [13 empty] | manual
|  | Sinus 5.0 H30s | [22 empty] | 5 | 5 | 17 | H30s medium smooth | Sinuses | [13 empty] | manual
|  | Sinus 1.0 H60s | [22 empty] | 1 | 0.7 | 115 | H60s sharp FR | Sinuses | [13 empty] | manual
|  | Sinus 1.0 H30s | [22 empty] | 1 | 0.7 | 115 | H30s medium smooth | Sinuses | [13 empty] | manual
Sinus_Vol (Adult)
Sinus_Vol (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.29 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head
| Sinus |  | 120 | 120 | 50 | 50 |  |  | On | 5 | On | CARE Dose4D | 7.05 | 10s | 25% |  |  | 1 |  | 2 | 0.8 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | Sinus 5.0 H60s | [22 empty] | 5 | 5 | 17 | H60s sharp FR | Sinuses | [13 empty] | manual
|  | Sinus 3.0 MPR cor | [22 empty] | 3 | 3 | 37 | H30s medium smooth | Sinuses | [13 empty] | manual
|  | Sinus 3.0 MPR cor | [22 empty] | 3 | 3 | 37 | H60s sharp FR | Sinuses | [13 empty] | manual
Neck
NeckRoutine (Adult)
NeckRoutine (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Neck |  | 120 | 120 | 165 | 165 |  |  | On | 7 | On | CARE Dose4D | 11.14 | 9s | 25% |  |  | 1 |  | 2 | 0.8 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | Neck 5.0 B31s | [22 empty] | 5 | 5 | 37 | B31s medium smooth + | Larynx | [13 empty] | manual
|  | Neck 1.0 B20s | [22 empty] | 1 | 0.7 | 258 | B20s smooth | Larynx | [13 empty] | manual
NeckThorAbd (Adult)
NeckThorAbd (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 365 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Neck |  | 120 | 120 | 165 | 165 |  |  | On | 7 | On | CARE Dose4D | 11.14 | 7s | 25% |  |  | 1 |  | 2 | 0.8 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | Neck 5.0 B31s | [22 empty] | 5 | 5 | 29 | B31s medium smooth + | Larynx | [13 empty] | manual
|  | Neck 1.0 B20s | [22 empty] | 1 | 0.7 | 201 | B20s smooth | Larynx | [13 empty] | manual
| ThorAbd |  | 120 | 120 | 210 | 210 |  |  | On | 7 | On | CARE Dose4D | 14.18 | 40s | 25% |  |  | 0.5 |  | 2 | 0.8 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | Tho ThorAbd 5.0 B31f | [22 empty] | 5 | 5 | 41 | B31f medium smooth + | Mediastinum | [13 empty] | manual
|  | Lun ThorAbd 5.0 B80f | [22 empty] | 5 | 5 | 41 | B80f ultra sharp | Lung | [13 empty] | manual
|  | Abd ThorAbd 5.0 B30f | [22 empty] | 5 | 5 | 81 | B30f medium smooth | Abdomen | [13 empty] | manual
|  | MPR ThorAbd 1.0 B30f | [22 empty] | 1 | 0.7 | 858 | B30f medium smooth | Mediastinum | [13 empty] | manual
NeckThorax (Adult)
NeckThorax (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Neck |  | 120 | 120 | 165 | 165 |  |  | On | 7 | On | CARE Dose4D | 11.14 | 7s | 25% |  |  | 1 |  | 2 | 0.8 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | Neck 5.0 B31s | [22 empty] | 5 | 5 | 29 | B31s medium smooth + | Larynx | [13 empty] | manual
|  | Neck 1.0 B20s | [22 empty] | 1 | 0.7 | 201 | B20s smooth | Larynx | [13 empty] | manual
| Thorax |  | 120 | 120 | 110 | 110 |  |  | On | 7 | On | CARE Dose4D | 7.43 | 7s | 25% |  |  | 0.5 |  | 2 | 0.8 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | Thorax 5.0 B31f | [22 empty] | 5 | 5 | 61 | B31f medium smooth + | Mediastinum | [13 empty] | manual
|  | Thorax 5.0 B70f | [22 empty] | 5 | 5 | 61 | B70f very sharp | Lung | [13 empty] | manual
|  | Thorax 1.0 B80f | [22 empty] | 1 | 0.7 | 429 | B80f ultra sharp | Lung | [13 empty] | manual
Neck_IR (Adult)
Neck_IR (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Neck |  | 120 | 120 | 116 | 116 |  |  | On | 7 | On | CARE Dose4D | 7.85 | 9s | 25% |  |  | 1 |  | 2 | 0.8 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | Neck 5.0 I31s 3 | [22 empty] | 5 | 5 | 37 | I31s medium smooth | Larynx | [13 empty] | manual
|  | Neck 1.0 I26s 3 | [22 empty] | 1 | 0.7 | 258 | I26s medium smooth ASA | Larynx | [13 empty] | manual
Shoulder
ShoulderRoutine (Adult)
ShoulderRoutine (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Shoulder |  | 120 | 120 | 165 | 165 |  |  | On | 5 | On | CARE Dose4D | 11.14 | 8s | 25% |  |  | 1 |  | 2 | 0.8 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | Shoulder 5.0 B60s | [22 empty] | 5 | 5 | 31 | B60s sharp | Bone | [13 empty] | manual
|  | Shoulder 0.6 B31s | [22 empty] | 0.6 | 0.4 | 376 | B31s medium smooth + | Shoulder | [13 empty] | manual
|  | Shoulder 0.6 B60s | [22 empty] | 0.6 | 0.4 | 376 | B60s sharp | Bone | [13 empty] | manual
Thorax
5010_ang_Che_aor (Adult)
5010_ang_Che_aor (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| PreMonitoring |  | 100 | 120 | 146 | 40 |  |  | Off | 8 | Off | CARE Dose4D | 2.04 |  | 25% |  |  | 0.285 |  | 2 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | PreMonitoring 10.0 B30f | [22 empty] | 10 |  | 1 | B30f medium smooth | Mediastinum | [13 empty] | manual
| Monitoring |  | 100 | 120 | 146 | 40 |  |  | Off | 8 | Off | CARE Dose4D | 61.08 |  | 25% |  |  | 0.285 |  | 5 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | Monitoring 10.0 B30f | [22 empty] | 10 |  | 30 | B30f medium smooth | Mediastinum | [13 empty] | manual
| art Che Flash |  | 100 | 100 | 146 |  | 100 |  | On | 8 | On | CARE Dose4D | 4.84 | 2s | 25% | 50 |  | 0.285 |  | 10 | 2.2 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4
|  | art Che | [22 empty] | 3 | 2 | 126 | I41f medium | Mediastinum |  |  |  | GEPACS
|  | art Che sag | [22 empty] | 3 | 3 | 101 | I41f medium | Bone |  |  |  | GEPACS
|  | art Che cor mip | [22 empty] | 10 | 10 | 31 | I41f medium | Chest MIP |  |  |  | GEPACS
|  | art Che obl sag | [22 empty] | 2 | 2 | 151 | I41f medium | CT Angio |  |  |  | GEPACS
|  | art Che c010 :-D | [22 empty] | 1 | 1 | 251 | I26f medium smooth ASA | Mediastinum |  |  |  | ROUTER | GEPACS
| art Che Helical |  | 120 | 120 | 180 | 110 |  |  | On | 8 | On | CARE Dose4D | 12.14 | 30s | 25% | 50 |  | 0.5 |  | 10 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | art Che | [22 empty] | 3 | 2 | 145 | I30f medium smooth | Mediastinum |  |  |  | GEPACS | [9 empty] | manual
|  | art Che sag | [22 empty] | 3 | 3 | 127 | I30f medium smooth | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | art Che obl sag | [22 empty] | 3 | 3 | 127 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | art Che cor mip | [22 empty] | 10 | 10 | 37 | I30f medium smooth | Chest MIP |  |  |  | GEPACS | [9 empty] | manual
|  | art Che c010 :-D | [22 empty] | 1 | 1 | 289 | I30f medium smooth | Abdomen |  |  |  | ROUTER | GEPACS | [8 empty] | manual
5020_ang_CheAbd_aor (Adult)
5020_ang_CheAbd_aor (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| PreMonitoring |  | 100 | 120 | 146 | 40 |  |  | Off | 8 | Off | CARE Dose4D | 2.04 |  | 25% |  |  | 0.285 |  | 2 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | PreMonitoring 10.0 B30f | [22 empty] | 10 |  | 1 | B30f medium smooth | Mediastinum | [13 empty] | manual
| Monitoring |  | 100 | 120 | 146 | 40 |  |  | Off | 8 | Off | CARE Dose4D | 61.08 |  | 25% |  |  | 0.285 |  | 5 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | Monitoring 10.0 B30f | [22 empty] | 10 |  | 30 | B30f medium smooth | Mediastinum | [13 empty] | manual
| art CA Flash |  | 100 | 100 | 146 |  | 100 |  | On | 8 | On | CARE Dose4D | 4.84 | 2s | 25% | 50 |  | 0.285 |  | 10 | 2.2 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4
|  | art CheAbd | [22 empty] | 3 | 2 | 126 | I41f medium | Mediastinum |  |  |  | GEPACS
|  | art CheAbd sag | [22 empty] | 3 | 3 | 101 | I41f medium | Bone |  |  |  | GEPACS
|  | art CheAbd cor mip | [22 empty] | 10 | 10 | 31 | I41f medium | Chest MIP |  |  |  | GEPACS
|  | art CheAbd obl sag | [22 empty] | 2 | 2 | 151 | I41f medium | CT Angio |  |  |  | GEPACS
|  | art CheAbd c020 :-D | [22 empty] | 1 | 1 | 251 | I26f medium smooth ASA | Mediastinum |  |  |  | ROUTER | GEPACS
| art CA Helical |  | 120 | 120 | 180 | 110 |  |  | On | 8 | On | CARE Dose4D | 12.14 | 30s | 25% | 50 |  | 0.5 |  | 10 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | art CA Helical c020 | [22 empty] | 3 | 2 | 145 | I30f medium smooth | Mediastinum |  |  |  | GEPACS | [9 empty] | manual
|  | art CheAbd sag | [22 empty] | 3 | 3 | 127 | I30f medium smooth | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | art CA Helical obl sag | [22 empty] | 3 | 3 | 127 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | art CheAbd cor mip | [22 empty] | 10 | 10 | 37 | I30f medium smooth | Chest MIP |  |  |  | GEPACS | [9 empty] | manual
|  | art CA Helical c020 :-D | [22 empty] | 1 | 1 | 289 | I30f medium smooth | Abdomen |  |  |  | ROUTER | GEPACS | [8 empty] | manual
5030_ang_CheAbdPel_aor (Adult)
5030_ang_CheAbdPel_aor (Adult) | Topogram | Topogram |  | 100 |  | 393 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 393 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| PreMonitoring |  | 100 | 120 | 146 | 40 |  |  | Off | 8 | Off | CARE Dose4D | 2.04 |  | 25% |  |  | 0.285 |  | 2 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | PreMonitoring 10.0 B30f | [22 empty] | 10 |  | 1 | B30f medium smooth | Mediastinum | [13 empty] | manual
| Monitoring |  | 100 | 120 | 146 | 40 |  |  | Off | 8 | Off | CARE Dose4D | 61.08 |  | 25% |  |  | 0.285 |  | 5 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | Monitoring 10.0 B30f | [22 empty] | 10 |  | 30 | B30f medium smooth | Mediastinum | [13 empty] | manual
| art CAP Flash |  | 100 | 100 | 146 |  | 100 |  | On | 8 | On | CARE Dose4D | 4.84 | 2s | 25% | 50 |  | 0.285 |  | 10 | 2.2 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4
|  | art CheAbdPel | [22 empty] | 3 | 2 | 126 | I41f medium | Mediastinum |  |  |  | GEPACS
|  | art CheAbdPel sag | [22 empty] | 3 | 3 | 101 | I41f medium | Bone |  |  |  | GEPACS
|  | art CheAbdPel cor mip | [22 empty] | 10 | 10 | 31 | I41f medium | Chest MIP |  |  |  | GEPACS
|  | art CheAbdPel obl sag | [22 empty] | 2 | 2 | 151 | I41f medium | CT Angio |  |  |  | GEPACS
|  | art CheAbdPel c030 :-D | [22 empty] | 1 | 1 | 251 | I26f medium smooth ASA | Mediastinum |  |  |  | ROUTER | GEPACS
| art CAP Helical |  | 120 | 120 | 180 | 110 |  |  | On | 8 | On | CARE Dose4D | 12.14 | 30s | 25% | 50 |  | 0.5 |  | 10 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | art CheAbdPel | [22 empty] | 3 | 2 | 145 | I30f medium smooth | Mediastinum |  |  |  | GEPACS | [9 empty] | manual
|  | art CheAbdPel sag | [22 empty] | 3 | 3 | 127 | I30f medium smooth | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | art CheAbdPel obl sag | [22 empty] | 3 | 3 | 127 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | art CheAbdPel cor mip | [22 empty] | 24 | 5 | 73 | I30f medium smooth | Chest MIP |  |  |  | GEPACS | [9 empty] | manual
|  | art CheAbdPel c030 :-D | [22 empty] | 1 | 1 | 289 | I30f medium smooth | Abdomen |  |  |  | ROUTER | GEPACS | [8 empty] | manual
5035_ang_Che_ro_dis (Adult)
5035_ang_Che_ro_dis (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Che c035 |  | 120 | 120 | 150 | 110 |  |  | On | 7 | On | CARE Dose4D | 10.12 | 30s | 25% | 50 |  | 0.5 |  | 4 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | Che | [22 empty] | 3 | 2 | 145 | I41f medium | Chest |  |  |  | GEPACS | [9 empty] | manual
|  | Che sag | [22 empty] | 3 | 2 | 191 | I41f medium | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Che mip | [22 empty] | 10 | 3 | 97 | I41f medium | Lung |  |  |  | GEPACS | [9 empty] | manual
|  | Che c130 3dt | [22 empty] | 1 | 1 | 289 | I50f medium sharp ASA | Lung |  |  |  | ROUTER | GEPACS | [8 empty] | manual
| PreMonitoring |  | 120 | 120 | 180 | 50 |  |  | Off | 3 | Off | CARE Dose4D | 3.08 |  | 25% |  |  | 0.5 |  | 2 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | PreMonitoring | [22 empty] | 10 |  | 1 | B40s medium | 400/40 Abd | [13 empty] | manual
| Monitoring |  | 120 | 120 | 180 | 50 |  |  | Off | 3 | Off | CARE Dose4D | 123.28 |  | 25% |  |  | 0.5 |  | 10 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | Monitoring | [22 empty] | 10 |  | 40 | B40s medium | 400/40 Abd | [13 empty] | manual
| art Che |  | 120 | 120 | 180 | 110 |  |  | On | 8 | On | CARE Dose4D | 12.14 | 30s | 25% | 50 |  | 0.5 |  | 4 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | art Che c035 | [22 empty] | 3 | 2 | 145 | I30f medium smooth | 500/60 Vascular |  |  |  | GEPACS | [9 empty] | manual
|  | art Che cor | [22 empty] | 3 | 3 | 127 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | art Che sag | [22 empty] | 3 | 3 | 127 | I30f medium smooth | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | art Che obl sag | [22 empty] | 3 | 3 | 127 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | art Che cor mip | [22 empty] | 10 | 10 | 37 | I30f medium smooth | Chest MIP |  |  |  | GEPACS | [9 empty] | manual
|  | art Che c035 :-D | [22 empty] | 0.6 | 0.6 | 481 | I30f medium smooth | Abdomen |  |  |  | ROUTER | GEPACS | [8 empty] | manual
5040_ang_CheAbd_ro_dis (Adult)
5040_ang_CheAbd_ro_dis (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Che c040 |  | 120 | 120 | 150 | 110 |  |  | On | 7 | On | CARE Dose4D | 10.12 | 30s | 25% | 50 |  | 0.5 |  | 4 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | Che | [22 empty] | 3 | 2 | 145 | I41f medium | Chest |  |  |  | GEPACS | [9 empty] | manual
|  | Che sag | [22 empty] | 3 | 2 | 191 | I41f medium | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Che mip | [22 empty] | 10 | 3 | 97 | I41f medium | Lung |  |  |  | GEPACS | [9 empty] | manual
|  | Che c130 3dt | [22 empty] | 1 | 1 | 289 | I50f medium sharp ASA | Lung |  |  |  | ROUTER | GEPACS | [8 empty] | manual
| PreMonitoring |  | 120 | 120 | 180 | 50 |  |  | Off | 3 | Off | CARE Dose4D | 3.08 |  | 25% |  |  | 0.5 |  | 2 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | PreMonitoring | [22 empty] | 10 |  | 1 | B40s medium | 400/40 Abd | [13 empty] | manual
| Monitoring |  | 120 | 120 | 180 | 50 |  |  | Off | 3 | Off | CARE Dose4D | 123.28 |  | 25% |  |  | 0.5 |  | 10 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | Monitoring | [22 empty] | 10 |  | 40 | B40s medium | 400/40 Abd | [13 empty] | manual
| art CheAbd |  | 120 | 120 | 180 | 110 |  |  | On | 8 | On | CARE Dose4D | 12.14 | 30s | 25% | 50 |  | 0.5 |  | 4 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | art CheAbd c040 | [22 empty] | 3 | 2 | 145 | I30f medium smooth | 500/60 Vascular |  |  |  | GEPACS | [9 empty] | manual
|  | art CheAbd cor | [22 empty] | 3 | 3 | 127 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | art CheAbd sag | [22 empty] | 3 | 3 | 127 | I30f medium smooth | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | art CheAbd obl sag | [22 empty] | 3 | 3 | 127 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | art CheAbd cor mip | [22 empty] | 10 | 10 | 37 | I30f medium smooth | Chest MIP |  |  |  | GEPACS | [9 empty] | manual
|  | art CheAbd c040 :-D | [22 empty] | 0.6 | 0.6 | 481 | I30f medium smooth | Abdomen |  |  |  | ROUTER | GEPACS | [8 empty] | manual
5050_ang_CheAbdPel_ro_dis (Adult)
5050_ang_CheAbdPel_ro_dis (Adult) | Topogram | Topogram |  | 100 |  | 393 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 393 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Che c050 |  | 120 | 120 | 150 | 110 |  |  | On | 7 | On | CARE Dose4D | 10.12 | 30s | 25% | 50 |  | 0.5 |  | 4 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | Che | [22 empty] | 3 | 2 | 145 | I41f medium | Chest |  |  |  | GEPACS | [9 empty] | manual
|  | Che sag | [22 empty] | 3 | 2 | 191 | I41f medium | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Che mip | [22 empty] | 10 | 3 | 97 | I41f medium | Lung |  |  |  | GEPACS | [9 empty] | manual
|  | Che c130 3dt | [22 empty] | 1 | 1 | 289 | I50f medium sharp ASA | Lung |  |  |  | ROUTER | GEPACS | [8 empty] | manual
| PreMonitoring |  | 120 | 120 | 180 | 50 |  |  | Off | 3 | Off | CARE Dose4D | 3.08 |  | 25% |  |  | 0.5 |  | 2 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | PreMonitoring | [22 empty] | 10 |  | 1 | B40s medium | 400/40 Abd | [13 empty] | manual
| Monitoring |  | 120 | 120 | 180 | 50 |  |  | Off | 3 | Off | CARE Dose4D | 123.28 |  | 25% |  |  | 0.5 |  | 10 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | Monitoring | [22 empty] | 10 |  | 40 | B40s medium | 400/40 Abd | [13 empty] | manual
| art CheAbdPel |  | 120 | 120 | 180 | 110 |  |  | On | 8 | On | CARE Dose4D | 12.14 | 30s | 25% | 50 |  | 0.5 |  | 4 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | art CheAbdPel c050 | [22 empty] | 3 | 2 | 145 | I30f medium smooth | 500/60 Vascular |  |  |  | GEPACS | [9 empty] | manual
|  | art CheAbdPel cor | [22 empty] | 3 | 3 | 127 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | art CheAbdPel sag | [22 empty] | 3 | 3 | 127 | I30f medium smooth | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | art CheAbdPel obl sag | [22 empty] | 3 | 3 | 127 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | art CheAbdPel cor mip | [22 empty] | 10 | 10 | 37 | I30f medium smooth | Chest MIP |  |  |  | GEPACS | [9 empty] | manual
|  | art CheAbdPel c040 :-D | [22 empty] | 0.6 | 0.6 | 481 | I30f medium smooth | Abdomen |  |  |  | ROUTER | GEPACS | [8 empty] | manual
5052_ang_Che_pos_ste (Adult)
5052_ang_Che_pos_ste (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Che c052 |  | 120 | 120 | 150 | 110 |  |  | On | 7 | On | CARE Dose4D | 10.12 | 30s | 25% | 50 |  | 0.5 |  | 4 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | Che c052 | [22 empty] | 3 | 2 | 145 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | Che c052 :-D | [22 empty] | 0.6 | 0.6 | 481 | I30f medium smooth | 1500/600 Lung |  |  |  | ROUTER | GEPACS | [8 empty] | manual
| PreMonitoring |  | 120 | 120 | 180 | 50 |  |  | Off | 3 | Off | CARE Dose4D | 3.08 |  | 25% |  |  | 0.5 |  | 2 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | PreMonitoring | [22 empty] | 10 |  | 1 | B40s medium | 400/40 Abd | [13 empty] | manual
| Monitoring |  | 120 | 120 | 180 | 50 |  |  | Off | 3 | Off | CARE Dose4D | 123.28 |  | 25% |  |  | 0.5 |  | 10 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | Monitoring | [22 empty] | 10 |  | 40 | B40s medium | 400/40 Abd | [13 empty] | manual
| art Che c052 |  | 120 | 120 | 180 | 110 |  |  | On | 8 | On | CARE Dose4D | 12.14 | 30s | 25% | 50 |  | 0.5 |  | 4 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | art Che c052 | [22 empty] | 3 | 2 | 145 | I30f medium smooth | 500/60 Vascular |  |  |  | GEPACS | [9 empty] | manual
|  | art Che cor | [22 empty] | 3 | 3 | 127 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | art Che sag | [22 empty] | 3 | 3 | 127 | I30f medium smooth | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | art Che cor mip | [22 empty] | 10 | 10 | 37 | I30f medium smooth | Chest MIP |  |  |  | GEPACS | [9 empty] | manual
|  | art Che c052 :-D | [22 empty] | 0.6 | 0.6 | 481 | I30f medium smooth | Abdomen |  |  |  | ROUTER | GEPACS | [8 empty] | manual
| ven Che c052 |  | 120 | 120 | 180 | 110 |  |  | On | 7 | On | CARE Dose4D | 12.14 | 30s | 25% | 50 |  | 0.5 |  | 40 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven Che c052 c052 | [22 empty] | 3 | 2 | 145 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | ven Che cor | [22 empty] | 3 | 3 | 127 | I30f medium smooth | 1500/600 Lung | [13 empty] | manual
|  | ven Che c052 :-D | [22 empty] | 0.6 | 0.6 | 481 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | [8 empty] | manual
5055_ang_CheAbd_pos_ste (Adult)
5055_ang_CheAbd_pos_ste (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Che c055 |  | 120 | 120 | 150 | 110 |  |  | On | 7 | On | CARE Dose4D | 10.12 | 30s | 25% | 50 |  | 0.5 |  | 4 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | Che c055 | [22 empty] | 3 | 2 | 145 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | Che c055 :-D | [22 empty] | 0.6 | 0.6 | 481 | I30f medium smooth | 1500/600 Lung |  |  |  | ROUTER | GEPACS | [8 empty] | manual
| PreMonitoring |  | 120 | 120 | 180 | 50 |  |  | Off | 3 | Off | CARE Dose4D | 3.08 |  | 25% |  |  | 0.5 |  | 2 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | PreMonitoring | [22 empty] | 10 |  | 1 | B40s medium | 400/40 Abd | [13 empty] | manual
| Monitoring |  | 120 | 120 | 180 | 50 |  |  | Off | 3 | Off | CARE Dose4D | 123.28 |  | 25% |  |  | 0.5 |  | 10 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | Monitoring | [22 empty] | 10 |  | 40 | B40s medium | 400/40 Abd | [13 empty] | manual
| art CheAbd |  | 120 | 120 | 180 | 110 |  |  | On | 8 | On | CARE Dose4D | 12.14 | 30s | 25% | 50 |  | 0.5 |  | 4 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | art CheAbd c055 | [22 empty] | 3 | 2 | 145 | I30f medium smooth | 500/60 Vascular |  |  |  | GEPACS | [9 empty] | manual
|  | art CheAbd cor | [22 empty] | 3 | 3 | 127 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | art CheAbd sag | [22 empty] | 3 | 3 | 127 | I30f medium smooth | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | art CheAbd cor mip | [22 empty] | 10 | 10 | 37 | I30f medium smooth | Chest MIP |  |  |  | GEPACS | [9 empty] | manual
|  | art CheAbd c055 :-D | [22 empty] | 0.6 | 0.6 | 481 | I30f medium smooth | Abdomen |  |  |  | ROUTER | GEPACS | [8 empty] | manual
| ven Che |  | 120 | 120 | 180 | 110 |  |  | On | 7 | On | CARE Dose4D | 12.14 | 30s | 25% | 50 |  | 0.5 |  | 40 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven Che c055 | [22 empty] | 3 | 2 | 145 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | ven Che cor | [22 empty] | 3 | 3 | 127 | I30f medium smooth | 1500/600 Lung | [13 empty] | manual
|  | ven Che c055 :-D | [22 empty] | 0.6 | 0.6 | 481 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | [8 empty] | manual
5058_ang_CheAbdPel_pos_ste (Adult)
5058_ang_CheAbdPel_pos_ste (Adult) | Topogram | Topogram |  | 100 |  | 393 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 393 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Che c058 |  | 120 | 120 | 150 | 110 |  |  | On | 7 | On | CARE Dose4D | 10.12 | 30s | 25% | 50 |  | 0.5 |  | 4 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | Che c058 | [22 empty] | 3 | 2 | 145 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | Che c058 :-D | [22 empty] | 0.6 | 0.6 | 481 | I30f medium smooth | 1500/600 Lung |  |  |  | ROUTER | GEPACS | [8 empty] | manual
| PreMonitoring |  | 120 | 120 | 180 | 50 |  |  | Off | 3 | Off | CARE Dose4D | 3.08 |  | 25% |  |  | 0.5 |  | 2 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | PreMonitoring | [22 empty] | 10 |  | 1 | B40s medium | 400/40 Abd | [13 empty] | manual
| Monitoring |  | 120 | 120 | 180 | 50 |  |  | Off | 3 | Off | CARE Dose4D | 123.28 |  | 25% |  |  | 0.5 |  | 10 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | Monitoring | [22 empty] | 10 |  | 40 | B40s medium | 400/40 Abd | [13 empty] | manual
| art CheAbdPel |  | 120 | 120 | 180 | 110 |  |  | On | 8 | On | CARE Dose4D | 12.14 | 30s | 25% | 50 |  | 0.5 |  | 4 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | art CheAbdPel c058 | [22 empty] | 3 | 2 | 145 | I30f medium smooth | 500/60 Vascular |  |  |  | GEPACS | [9 empty] | manual
|  | art CheAbdPel cor | [22 empty] | 3 | 3 | 127 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | art CheAbdPel sag | [22 empty] | 3 | 3 | 127 | I30f medium smooth | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | art CheAbdPel cor mip | [22 empty] | 10 | 10 | 37 | I30f medium smooth | Chest MIP |  |  |  | GEPACS | [9 empty] | manual
|  | art CheAbdPel c058 :-D | [22 empty] | 0.6 | 0.6 | 481 | I30f medium smooth | Abdomen |  |  |  | ROUTER | GEPACS | [8 empty] | manual
| ven Che |  | 120 | 120 | 180 | 110 |  |  | On | 7 | On | CARE Dose4D | 12.14 | 30s | 25% | 50 |  | 0.5 |  | 40 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven Che c058 | [22 empty] | 3 | 2 | 145 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | ven Che cor | [22 empty] | 3 | 3 | 127 | I30f medium smooth | 1500/600 Lung | [13 empty] | manual
|  | ven Che c058 :-D | [22 empty] | 0.6 | 0.6 | 481 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | [8 empty] | manual
5060_ang_Che_pe (Adult)
5060_ang_Che_pe (Adult) | Topogram | Topogram 0.6 T20f |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 3 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 06fa8454-4bdd-4eb6-b75d-9ae4e311a76d |  |  | GEPACS |  | TRACKER
| Topogram | Topogram 0.6 T20f |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 06fa8454-4bdd-4eb6-b75d-9ae4e311a76d |  |  | GEPACS |  | TRACKER
| PreMonitoring |  | 100 | 120 | 120 | 40 |  |  | Off | 9 | Off | CARE Dose4D | 2.04 |  | 25% |  |  | 0.285 |  | 2 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | PreMonitoring 10.0 B30f | [22 empty] | 10 |  | 1 | B30f medium smooth | Mediastinum | [13 empty] | manual
| Monitoring |  | 100 | 120 | 120 | 40 |  |  | Off | 9 | Off | CARE Dose4D | 61.08 |  | 25% |  |  | 0.285 |  | 10 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | Monitoring 10.0 B30f | [22 empty] | 10 |  | 30 | B30f medium smooth | Mediastinum | [13 empty] | manual
| art Che Flash |  | 100 | 100 | 146 |  | 100 |  | On | 8 | On | CARE Dose4D | 4.84 | 2s | 25% | 50 |  | 0.285 |  | 6 | 2.2 |  | 3 | 128x0.6mm | [5 empty] | 06fa8454-4bdd-4eb6-b75d-9ae4e311a76d
|  | art Che | [22 empty] | 3 | 2 | 126 | I41f medium | Chest |  |  |  | GEPACS
|  | art Che sag | [22 empty] | 3 | 2 | 151 | I41f medium | Bone |  |  |  | GEPACS
|  | art Che cor | [22 empty] | 3 | 2 | 151 | I41f medium | Bone |  |  |  | GEPACS
|  | art Che mip | [22 empty] | 10 | 3 | 101 | I41f medium | Lung |  |  |  | GEPACS
|  | art Che c060 3dt | [22 empty] | 1 | 1 | 251 | I26f medium smooth ASA | Chest |  |  |  | ROUTER | GEPACS
| Topogram | Topogram 0.6 T20f |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram 0.6 T20f |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| art Che Helical |  | 120 | 120 | 150 | 110 |  |  | On | 8 | On | CARE Dose4D | 10.12 | 30s | 25% | 50 |  | 0.5 |  | 6 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | art Che | [22 empty] | 3 | 2 | 145 | I41f medium | Chest |  |  |  | GEPACS | [9 empty] | manual
|  | art Che sag | [22 empty] | 3 | 2 | 191 | I41f medium | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | art Che cor | [22 empty] | 3 | 2 | 191 | I41f medium | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | art Che mip | [22 empty] | 10 | 3 | 121 | I41f medium | Lung |  |  |  | GEPACS | [9 empty] | manual
|  | art Che c060 3dt | [22 empty] | 1 | 1 | 289 | I26f medium smooth ASA | Chest |  |  |  | ROUTER | GEPACS | [8 empty] | manual
5063_ang_Che_pe_ld (Adult)
5063_ang_Che_pe_ld (Adult) | Topogram | Topogram 0.6 T20f |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 3 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 06fa8454-4bdd-4eb6-b75d-9ae4e311a76d |  |  | GEPACS |  | TRACKER
| Topogram | Topogram 0.6 T20f |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 06fa8454-4bdd-4eb6-b75d-9ae4e311a76d |  |  | GEPACS |  | TRACKER
| PreMonitoring |  | 100 | 120 | 120 | 40 |  |  | Off | 9 | Off | CARE Dose4D | 2.04 |  | 25% |  |  | 0.285 |  | 2 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | PreMonitoring 10.0 B30f | [22 empty] | 10 |  | 1 | B30f medium smooth | Mediastinum | [13 empty] | manual
| Monitoring |  | 100 | 120 | 120 | 40 |  |  | Off | 9 | Off | CARE Dose4D | 61.08 |  | 25% |  |  | 0.285 |  | 10 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | Monitoring 10.0 B30f | [22 empty] | 10 |  | 30 | B30f medium smooth | Mediastinum | [13 empty] | manual
| art Che Flash |  | 100 | 100 | 146 |  | 100 |  | On | 9 | On | CARE Dose4D | 4.84 | 2s | 25% | 50 |  | 0.285 |  | 2 | 2.2 |  | 3 | 128x0.6mm | [5 empty] | 06fa8454-4bdd-4eb6-b75d-9ae4e311a76d
|  | art Che | [22 empty] | 3 | 2 | 126 | I41f medium | Chest |  |  |  | GEPACS
|  | art Che sag | [22 empty] | 3 | 2 | 151 | I41f medium | Bone |  |  |  | GEPACS
|  | art Che cor | [22 empty] | 3 | 2 | 151 | I41f medium | Bone |  |  |  | GEPACS
|  | art Che mip | [22 empty] | 10 | 3 | 101 | I41f medium | Lung |  |  |  | GEPACS
|  | art Che c063 3dt | [22 empty] | 1 | 1 | 251 | I26f medium smooth ASA | Chest |  |  |  | ROUTER | GEPACS
| Topogram | Topogram 0.6 T20f |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram 0.6 T20f |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| art Che Helical |  | 120 | 120 | 150 | 110 |  |  | On | 7 | On | CARE Dose4D | 10.12 | 30s | 25% | 50 |  | 0.5 |  | 6 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | art Che | [22 empty] | 3 | 2 | 145 | I41f medium | Chest |  |  |  | GEPACS | [9 empty] | manual
|  | art Che sag | [22 empty] | 3 | 2 | 191 | I41f medium | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | art Che cor | [22 empty] | 3 | 2 | 191 | I41f medium | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | art Che mip | [22 empty] | 10 | 3 | 121 | I41f medium | Lung |  |  |  | GEPACS | [9 empty] | manual
|  | art Che c063 3dt | [22 empty] | 1 | 1 | 289 | I26f medium smooth ASA | Chest |  |  |  | ROUTER | GEPACS | [8 empty] | manual
5080_ang_Che_pe_Abd (Adult)
5080_ang_Che_pe_Abd (Adult) | Topogram | Topogram 0.6 T20f |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 3 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 06fa8454-4bdd-4eb6-b75d-9ae4e311a76d |  |  | GEPACS |  | TRACKER
| Topogram | Topogram 0.6 T20f |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 06fa8454-4bdd-4eb6-b75d-9ae4e311a76d |  |  | GEPACS |  | TRACKER
| PreMonitoring |  | 100 | 120 | 120 | 40 |  |  | Off | 9 | Off | CARE Dose4D | 2.04 |  | 25% |  |  | 0.285 |  | 2 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | PreMonitoring 10.0 B30f | [22 empty] | 10 |  | 1 | B30f medium smooth | Mediastinum | [13 empty] | manual
| Monitoring |  | 100 | 120 | 120 | 40 |  |  | Off | 9 | Off | CARE Dose4D | 61.08 |  | 25% |  |  | 0.285 |  | 5 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | Monitoring 10.0 B30f | [22 empty] | 10 |  | 30 | B30f medium smooth | Mediastinum | [13 empty] | manual
| art Che Flash |  | 100 | 100 | 146 |  | 100 |  | On | 9 | On | CARE Dose4D | 4.84 | 2s | 25% | 50 |  | 0.285 |  | 2 | 2.2 |  | 3 | 128x0.6mm | [5 empty] | 06fa8454-4bdd-4eb6-b75d-9ae4e311a76d
|  | art Che | [22 empty] | 3 | 2 | 126 | I41f medium | Chest |  |  |  | GEPACS
|  | art Che sag | [22 empty] | 3 | 2 | 151 | I41f medium | Bone |  |  |  | GEPACS
|  | art Che cor | [22 empty] | 3 | 2 | 151 | I41f medium | Bone |  |  |  | GEPACS
|  | art Che mip | [22 empty] | 10 | 3 | 101 | I41f medium | Lung |  |  |  | GEPACS
|  | art Che c060 3dt | [22 empty] | 1 | 1 | 251 | I26f medium smooth ASA | Chest |  |  |  | ROUTER | GEPACS
| Topogram | Topogram 0.6 T20f |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram 0.6 T20f |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| art Che Helical |  | 120 | 120 | 150 | 110 |  |  | On | 7 | On | CARE Dose4D | 10.12 | 30s | 25% | 50 |  | 0.5 |  | 2 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | art Che | [22 empty] | 3 | 2 | 145 | I41f medium | Chest |  |  |  | GEPACS | [9 empty] | manual
|  | art Che sag | [22 empty] | 3 | 2 | 191 | I41f medium | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | art Che cor | [22 empty] | 3 | 2 | 191 | I41f medium | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | art Che mip | [22 empty] | 10 | 3 | 121 | I41f medium | Lung |  |  |  | GEPACS | [9 empty] | manual
|  | art Che c060 3dt | [22 empty] | 1 | 1 | 289 | I26f medium smooth ASA | Chest |  |  |  | ROUTER | GEPACS | [8 empty] | manual
| ven Abd |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 60 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven Abd a010 | [22 empty] | 3 | 3 | 131 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | ven Abd cor | [22 empty] | 3 | 3 | 67 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven Abd a010 :-D | [22 empty] | 0.6 | 0.6 | 653 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
5085_ang_Che_pe_ld_Abd (Adult)
5085_ang_Che_pe_ld_Abd (Adult) | Topogram | Topogram 0.6 T20f |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 3 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 06fa8454-4bdd-4eb6-b75d-9ae4e311a76d |  |  | GEPACS |  | TRACKER
| Topogram | Topogram 0.6 T20f |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 06fa8454-4bdd-4eb6-b75d-9ae4e311a76d |  |  | GEPACS |  | TRACKER
| PreMonitoring |  | 100 | 120 | 120 | 40 |  |  | Off | 9 | Off | CARE Dose4D | 2.04 |  | 25% |  |  | 0.285 |  | 2 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | PreMonitoring 10.0 B30f | [22 empty] | 10 |  | 1 | B30f medium smooth | Mediastinum | [13 empty] | manual
| Monitoring |  | 100 | 120 | 120 | 40 |  |  | Off | 9 | Off | CARE Dose4D | 61.08 |  | 25% |  |  | 0.285 |  | 5 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | Monitoring 10.0 B30f | [22 empty] | 10 |  | 30 | B30f medium smooth | Mediastinum | [13 empty] | manual
| art Che Flash |  | 100 | 100 | 146 |  | 100 |  | On | 9 | On | CARE Dose4D | 4.84 | 2s | 25% | 50 |  | 0.285 |  | 2 | 2.2 |  | 3 | 128x0.6mm | [5 empty] | 06fa8454-4bdd-4eb6-b75d-9ae4e311a76d
|  | art Che | [22 empty] | 3 | 2 | 126 | I41f medium | Chest |  |  |  | GEPACS
|  | art Che sag | [22 empty] | 3 | 2 | 151 | I41f medium | Bone |  |  |  | GEPACS
|  | art Che cor | [22 empty] | 3 | 2 | 151 | I41f medium | Bone |  |  |  | GEPACS
|  | art Che mip | [22 empty] | 10 | 3 | 101 | I41f medium | Lung |  |  |  | GEPACS
|  | art Che c063 3dt | [22 empty] | 1 | 1 | 251 | I26f medium smooth ASA | Chest |  |  |  | ROUTER | GEPACS
| Topogram | Topogram 0.6 T20f |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram 0.6 T20f |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| art Che Helical |  | 120 | 120 | 150 | 110 |  |  | On | 7 | On | CARE Dose4D | 10.12 | 30s | 25% | 50 |  | 0.5 |  | 2 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | art Che | [22 empty] | 3 | 2 | 145 | I41f medium | Chest |  |  |  | GEPACS | [9 empty] | manual
|  | art Che sag | [22 empty] | 3 | 2 | 191 | I41f medium | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | art Che cor | [22 empty] | 3 | 2 | 191 | I41f medium | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | art Che mip | [22 empty] | 10 | 3 | 121 | I41f medium | Lung |  |  |  | GEPACS | [9 empty] | manual
|  | art Che c063 3dt | [22 empty] | 1 | 1 | 289 | I26f medium smooth ASA | Chest |  |  |  | ROUTER | GEPACS | [8 empty] | manual
| ven Abd |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 60 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven Abd a010 | [22 empty] | 3 | 3 | 131 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | ven Abd cor | [22 empty] | 3 | 3 | 67 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven Abd a010 :-D | [22 empty] | 0.6 | 0.6 | 653 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
5090_ang_Che_pe_AbdPel (Adult)
5090_ang_Che_pe_AbdPel (Adult) | Topogram | Topogram |  | 100 |  | 393 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 06fa8454-4bdd-4eb6-b75d-9ae4e311a76d |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 393 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 06fa8454-4bdd-4eb6-b75d-9ae4e311a76d |  |  | GEPACS |  | TRACKER
| PreMonitoring |  | 100 | 120 | 120 | 40 |  |  | Off | 9 | Off | CARE Dose4D | 2.04 |  | 25% |  |  | 0.285 |  | 2 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | PreMonitoring 10.0 B30f | [22 empty] | 10 |  | 1 | B30f medium smooth | Mediastinum | [13 empty] | manual
| Monitoring |  | 100 | 120 | 120 | 40 |  |  | Off | 9 | Off | CARE Dose4D | 61.08 |  | 25% |  |  | 0.285 |  | 10 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | Monitoring 10.0 B30f | [22 empty] | 10 |  | 30 | B30f medium smooth | Mediastinum | [13 empty] | manual
| art Che Flash |  | 100 | 100 | 146 |  | 100 |  | On | 9 | On | CARE Dose4D | 4.84 | 2s | 25% | 50 |  | 0.285 |  | 2 | 2.2 |  | 3 | 128x0.6mm | [5 empty] | 06fa8454-4bdd-4eb6-b75d-9ae4e311a76d
|  | art Che | [22 empty] | 3 | 2 | 169 | I41f medium | Chest |  |  |  | GEPACS
|  | art Che sag | [22 empty] | 3 | 2 | 151 | I41f medium | Bone |  |  |  | GEPACS
|  | art Che cor | [22 empty] | 3 | 2 | 151 | I41f medium | Bone |  |  |  | GEPACS
|  | art Che mip | [22 empty] | 10 | 3 | 84 | I41f medium | Lung |  |  |  | GEPACS
|  | art Che c060 3dt | [22 empty] | 1 | 1 | 251 | I26f medium smooth ASA | Chest |  |  |  | ROUTER | GEPACS
| Topogram | Topogram |  | 100 |  | 393 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 393 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| art Che Helical |  | 120 | 120 | 150 | 110 |  |  | On | 7 | On | CARE Dose4D | 10.12 | 30s | 25% | 50 |  | 0.5 |  | 2 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | art Che | [22 empty] | 3 | 2 | 145 | I41f medium | Chest |  |  |  | GEPACS | [9 empty] | manual
|  | art Che sag | [22 empty] | 3 | 2 | 191 | I41f medium | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | art Che cor | [22 empty] | 3 | 2 | 191 | I41f medium | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | art Che mip | [22 empty] | 10 | 3 | 121 | I41f medium | Lung |  |  |  | GEPACS | [9 empty] | manual
|  | art Che c060 3dt | [22 empty] | 1 | 1 | 289 | I26f medium smooth ASA | Chest |  |  |  | ROUTER | GEPACS | [8 empty] | manual
| ven AbdPel |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 60 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven AbdPel a030 | [22 empty] | 3 | 3 | 131 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel cor | [22 empty] | 3 | 3 | 67 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel sag | [22 empty] | 3 | 3 | 121 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel a030 :-D | [22 empty] | 0.6 | 0.6 | 653 | I30f medium smooth | Abdomen |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
5095_ang_Che_pe_ld_AbdPel (Adult)
5095_ang_Che_pe_ld_AbdPel (Adult) | Topogram | Topogram |  | 100 |  | 393 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 06fa8454-4bdd-4eb6-b75d-9ae4e311a76d |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 393 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 06fa8454-4bdd-4eb6-b75d-9ae4e311a76d |  |  | GEPACS |  | TRACKER
| PreMonitoring |  | 100 | 120 | 120 | 40 |  |  | Off | 9 | Off | CARE Dose4D | 2.04 |  | 25% |  |  | 0.285 |  | 2 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | PreMonitoring 10.0 B30f | [22 empty] | 10 |  | 1 | B30f medium smooth | Mediastinum | [13 empty] | manual
| Monitoring |  | 100 | 120 | 120 | 40 |  |  | Off | 9 | Off | CARE Dose4D | 61.08 |  | 25% |  |  | 0.285 |  | 10 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | Monitoring 10.0 B30f | [22 empty] | 10 |  | 30 | B30f medium smooth | Mediastinum | [13 empty] | manual
| art Che Flash |  | 100 | 100 | 146 |  | 100 |  | On | 9 | On | CARE Dose4D | 4.84 | 2s | 25% | 50 |  | 0.285 |  | 2 | 2.2 |  | 3 | 128x0.6mm | [5 empty] | 06fa8454-4bdd-4eb6-b75d-9ae4e311a76d
|  | art Che | [22 empty] | 3 | 2 | 231 | I41f medium | Chest |  |  |  | GEPACS
|  | art Che sag | [22 empty] | 3 | 2 | 151 | I41f medium | Bone |  |  |  | GEPACS
|  | art Che cor | [22 empty] | 3 | 2 | 151 | I41f medium | Bone |  |  |  | GEPACS
|  | art Che mip | [22 empty] | 10 | 3 | 84 | I41f medium | Lung |  |  |  | GEPACS
|  | art Che c063 3dt | [22 empty] | 1 | 1 | 251 | I26f medium smooth ASA | Chest |  |  |  | ROUTER | GEPACS
| Topogram | Topogram |  | 100 |  | 393 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 393 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| art Che Helical |  | 120 | 120 | 150 | 110 |  |  | On | 7 | On | CARE Dose4D | 10.12 | 30s | 25% | 50 |  | 0.5 |  | 2 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | art Che | [22 empty] | 3 | 2 | 145 | I41f medium | Chest |  |  |  | GEPACS | [9 empty] | manual
|  | art Che sag | [22 empty] | 3 | 2 | 191 | I41f medium | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | art Che cor | [22 empty] | 3 | 2 | 191 | I41f medium | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | art Che mip | [22 empty] | 10 | 3 | 121 | I41f medium | Lung |  |  |  | GEPACS | [9 empty] | manual
|  | art Che c063 3dt | [22 empty] | 1 | 1 | 289 | I26f medium smooth ASA | Chest |  |  |  | ROUTER | GEPACS | [8 empty] | manual
| ven AbdPel |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 60 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven AbdPel a030 | [22 empty] | 3 | 3 | 131 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel cor | [22 empty] | 3 | 3 | 67 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel sag | [22 empty] | 3 | 3 | 121 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel a030 :-D | [22 empty] | 0.6 | 0.6 | 653 | I30f medium smooth | Abdomen |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
5100_ang_Hrt_cor_art (Adult)
5100_ang_Hrt_cor_art (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | ff800867-8571-4a15-997f-1d4ffbea68b7 |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | ff800867-8571-4a15-997f-1d4ffbea68b7 |  |  | GEPACS | ROEMIR001a | TRACKER
| CaScSeq |  | 120 | 120 | 100 |  | 80 |  | Semi | 3 | On | CARE Dose4D | 3.75 |  | 25% | 50 |  | 0.285 |  | 2 | 0 | 34 | 3 | 32x1.2mm | [5 empty] | ff800867-8571-4a15-997f-1d4ffbea68b7 | [7 empty] | 70% | 70% | Off | 70% | 70%
|  | Hrt nfov c100 | [22 empty] | 3 |  | 103 | I30f medium smooth | Mediastinum |  |  |  | GEPACS | MCH3DLabAW | [8 empty] | manual | 70%
|  | Hrt | [22 empty] | 3 |  | 103 | I30f medium smooth | Mediastinum |  |  |  | GEPACS | [9 empty] | manual | 70%
| CorCTA_Flash |  | 120 | 100 | 290 |  | 290 |  | On | 9 | On | CARE Dose4D | 3.43 |  | 25% | 50 |  | 0.285 |  | 21 | 3.4 |  | 0.75 | 128x0.6mm | [5 empty] | ff800867-8571-4a15-997f-1d4ffbea68b7
|  | art Hrt c100 :- | [22 empty] | 0.75 | 0.5 | 398 | I26f medium smooth ASA | Cardiac |  |  |  | GEPACS | MCH3DLabAW | SYNGO_VIA | [8 empty] | 60%
| Cor_Seq |  | 100 | 100 | 290 |  | 290 |  | Semi | 11 | On | CARE Dose4D | 8.19 |  | 25% | 50 |  | 0.285 |  | 21 | 0 | 34 | 0.75 | 128x0.6mm | [5 empty] | ff800867-8571-4a15-997f-1d4ffbea68b7 | [7 empty] | 65% | 70% | Normal | 70% | 70%
|  | art Hrt c100 :- | [22 empty] | 0.75 |  | 413 | I26f medium smooth ASA | Cardiac |  | Best Diastolic |  | GEPACS | MCH3DLabAW | SYNGO_VIA | [7 empty] | BestDiast | 70%
5103_ang_Hrt_cor_art_byp (Adult)
5103_ang_Hrt_cor_art_byp (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | ff800867-8571-4a15-997f-1d4ffbea68b7 |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | ff800867-8571-4a15-997f-1d4ffbea68b7 |  |  | GEPACS | ROEMIR001a | TRACKER
| CaScSeq |  | 120 | 120 | 100 |  | 80 |  | Semi | 3 | On | CARE Dose4D | 3.75 |  | 25% | 50 |  | 0.285 |  | 2 | 0 | 34 | 3 | 32x1.2mm | [5 empty] | ff800867-8571-4a15-997f-1d4ffbea68b7 | [7 empty] | 70% | 70% | Off | 70% | 70%
|  | Hrt nfov c100 | [22 empty] | 3 |  | 103 | I30f medium smooth | Mediastinum |  |  |  | GEPACS | MCH3DLabAW | [8 empty] | manual | 70%
|  | Hrt | [22 empty] | 3 |  | 103 | I30f medium smooth | Mediastinum |  |  |  | GEPACS | [9 empty] | manual | 70%
| CorCTA_Flash |  | 120 | 100 | 290 |  | 290 |  | On | 9 | On | CARE Dose4D | 3.43 |  | 25% | 50 |  | 0.285 |  | 21 | 3.4 |  | 0.75 | 128x0.6mm | [5 empty] | ff800867-8571-4a15-997f-1d4ffbea68b7
|  | art Hrt c103 :- | [22 empty] | 0.75 | 0.5 | 398 | I26f medium smooth ASA | Cardiac |  |  |  | GEPACS | MCH3DLabAW | SYNGO_VIA | [8 empty] | 60%
| Cor_Seq |  | 100 | 100 | 290 |  | 290 |  | Semi | 11 | On | CARE Dose4D | 5.47 |  | 25% | 50 |  | 0.285 |  | 21 | 0 | 34 | 0.75 | 128x0.6mm | [5 empty] | ff800867-8571-4a15-997f-1d4ffbea68b7 | [7 empty] | 65% | 70% | Normal | 70% | 70%
|  | art Hrt c103 :- | [22 empty] | 0.75 |  | 413 | I26f medium smooth ASA | Cardiac |  | Best Diastolic |  | GEPACS | MCH3DLabAW | SYNGO_VIA | [7 empty] | BestDiast | 70%
5110_ang_Hrt_cor_art_ret_gat (Adult)
5110_ang_Hrt_cor_art_ret_gat (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | ff800867-8571-4a15-997f-1d4ffbea68b7 |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | ff800867-8571-4a15-997f-1d4ffbea68b7 |  |  | GEPACS | ROEMIR001a | TRACKER
| CaScSeq |  | 120 | 120 | 100 |  | 80 |  | Semi | 3 | On | CARE Dose4D | 3.75 |  | 25% | 50 |  | 0.285 |  | 2 | 0 | 34 | 3 | 32x1.2mm | [5 empty] | ff800867-8571-4a15-997f-1d4ffbea68b7 | [7 empty] | 70% | 70% | Off | 70% | 70%
|  | Hrt nfov c100 | [22 empty] | 3 |  | 103 | I30f medium smooth | Mediastinum |  |  |  | GEPACS | MCH3DLabAW | [8 empty] | manual | 70%
|  | Hrt | [22 empty] | 3 |  | 103 | I30f medium smooth | Mediastinum |  |  |  | GEPACS | [9 empty] | manual | 70%
| Cor_Spiral |  | 120 | 120 | 290 |  | 220 |  | On | 9 | On | CARE Dose4D | 13.99 |  | 25% | 150 |  | 0.285 |  | 21 | 0.23 |  | 0.75 | 128x0.6mm | [5 empty] | ff800867-8571-4a15-997f-1d4ffbea68b7 | [9 empty] | Advance | 70% | 70%
|  | art Hrt c110 :- | [22 empty] | 0.75 | 0.5 | 401 | I26f medium smooth ASA | Cardiac |  | Best Diastolic |  | MCH3DLabAW | GEPACS | SYNGO_VIA | [7 empty] | BestDiast | 70%
5113_ang_Hrt_cor_art_byp_ret_gat (Adult)
5113_ang_Hrt_cor_art_byp_ret_gat (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | ff800867-8571-4a15-997f-1d4ffbea68b7 |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | ff800867-8571-4a15-997f-1d4ffbea68b7 |  |  | GEPACS | ROEMIR001a | TRACKER
| CaScSeq |  | 120 | 120 | 100 |  | 80 |  | Semi | 3 | On | CARE Dose4D | 3.75 |  | 25% | 50 |  | 0.285 |  | 2 | 0 | 34 | 3 | 32x1.2mm | [5 empty] | ff800867-8571-4a15-997f-1d4ffbea68b7 | [7 empty] | 70% | 70% | Off | 70% | 70%
|  | Hrt nfov c100 | [22 empty] | 3 |  | 103 | I30f medium smooth | Mediastinum |  |  |  | GEPACS | MCH3DLabAW | [8 empty] | manual | 70%
|  | Hrt | [22 empty] | 3 |  | 103 | I30f medium smooth | Mediastinum |  |  |  | GEPACS | [9 empty] | manual | 70%
| Cor_Spiral |  | 120 | 120 | 290 |  | 220 |  | On | 9 | On | CARE Dose4D | 19.54 |  | 25% | 150 |  | 0.285 |  | 21 | 0.23 |  | 0.75 | 128x0.6mm | [5 empty] | ff800867-8571-4a15-997f-1d4ffbea68b7 | [9 empty] | Advance | 70% | 70%
|  | art Hrt c113 :- | [22 empty] | 0.75 | 0.5 | 401 | I26f medium smooth ASA | Cardiac |  | Best Diastolic |  | MCH3DLabAW | GEPACS | SYNGO_VIA | [7 empty] | BestDiast | 70%
5116_ang_Hrt_val_ret_gat (Adult)
5116_ang_Hrt_val_ret_gat (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | ff800867-8571-4a15-997f-1d4ffbea68b7 |  |  |  |  | TRACKER
| Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | ff800867-8571-4a15-997f-1d4ffbea68b7 |  |  | GEPACS |  | TRACKER
| PreMonitoring |  | 120 | 120 | 290 | 40 |  |  | Off | 9 | Off | CARE Dose4D | 2.04 |  | 25% |  |  | 0.285 |  | 2 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | PreMonitoring 10.0 B30f | [22 empty] | 10 |  | 1 | B30f medium smooth | Cardiac | [13 empty] | manual
| Monitoring |  | 120 | 120 | 290 | 40 |  |  | Off | 9 | Off | CARE Dose4D | 61.08 |  | 25% |  |  | 0.285 |  | 10 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | Monitoring 10.0 B30f | [22 empty] | 10 |  | 30 | B30f medium smooth | Cardiac | [13 empty] | manual
| Retro_CorValve |  | 120 | 120 | 360 |  | 320 |  | Off | 9 | On | CARE Dose4D | 117.67 |  | 25% | 150 |  | 0.285 |  | 2 | 0.17 |  | 0.6 | 128x0.6mm | [5 empty] | ff800867-8571-4a15-997f-1d4ffbea68b7 | [9 empty] | Off | 70% | 70%
|  | art Hrt c116 :- | [22 empty] | 0.6 | 0.6 | 417 | I26f medium smooth ASA | Cardiac |  | Best Diastolic |  | GEPACS | MCH3DLabAW | TERARECON | [7 empty] | BestDiast | 70%
|  | art Hrt c116 :- | [22 empty] | 0.6 | 0.6 | 417 | I26f medium smooth ASA | Cardiac |  | Best Systolic |  | GEPACS | MCH3DLabAW | TERARECON | [7 empty] | BestSyst | 55%
|  | art Hrt function 0% | [22 empty] | 0.6 | 0.6 | 417 | I26f medium smooth ASA | Cardiac |  | Best Diastolic |  | TERARECON | [9 empty] | manual |  | 0%
5117_ang_Hrt_cor_art_byp_ret_gat_AbdPelThi (Adult)
5117_ang_Hrt_cor_art_byp_ret_gat_AbdPelThi (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 275 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | ff800867-8571-4a15-997f-1d4ffbea68b7 |  |  |  |  | TRACKER
| Topogram | Topogram 0.6 T20f |  | 120 |  | 275 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | ff800867-8571-4a15-997f-1d4ffbea68b7 |  |  | GEPACS |  | TRACKER
| Cacs |  | 120 | 120 | 100 |  | 80 |  | Semi | 3 | On | CARE Dose4D | 32.77 |  | 25% | 50 |  | 0.285 |  | 2 | 0.17 |  | 2 | 128x0.6mm | [5 empty] | ff800867-8571-4a15-997f-1d4ffbea68b7 | [9 empty] | Advance | 70% | 70%
|  | Hrt nfov c100 | [22 empty] | 2 | 2 | 126 | B35f HeartView medium CaSco | Mediastinum |  |  |  | GEPACS | SYNGO_VIA | MCH3DLabAW | [7 empty] | BestDiast | 70%
|  | Hrt | [22 empty] | 2 | 2 | 126 | B35f HeartView medium CaSco | Mediastinum |  |  |  | GEPACS | [9 empty] | BestDiast | 70%
| PreMonitoring |  | 120 | 120 | 290 | 40 |  |  | Off | 9 | Off | CARE Dose4D | 2.04 |  | 25% |  |  | 0.285 |  | 2 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | PreMonitoring 10.0 B30f | [22 empty] | 10 |  | 1 | B30f medium smooth | Cardiac | [13 empty] | manual
| Monitoring |  | 120 | 120 | 290 | 40 |  |  | Off | 9 | Off | CARE Dose4D | 61.08 |  | 25% |  |  | 0.285 |  | 10 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | Monitoring 10.0 B30f | [22 empty] | 10 |  | 30 | B30f medium smooth | Cardiac | [13 empty] | manual
| Retro_CorCTA |  | 120 | 120 | 240 |  | 320 |  | On | 9 | On | CARE Dose4D | 78.57 |  | 25% | 150 |  | 0.285 |  | 2 | 0.17 |  | 0.75 | 128x0.6mm | [5 empty] | ff800867-8571-4a15-997f-1d4ffbea68b7 | [9 empty] | Off | 70% | 70%
|  | art Hrt c113 :- | [22 empty] | 0.75 | 0.75 | 334 | I26f medium smooth ASA | Cardiac |  | Best Diastolic |  | GEPACS | MCH3DLabAW | TERARECON | [7 empty] | BestDiast | 70%
|  | art Hrt c113 :- | [22 empty] | 0.75 | 0.75 | 334 | I26f medium smooth ASA | Cardiac |  | Best Systolic |  | GEPACS | MCH3DLabAW | TERARECON | [7 empty] | BestSyst | 55%
|  | art Hrt c113 25% :- | [22 empty] | 0.75 | 0.75 | 334 | I26f medium smooth ASA | Cardiac |  | Best Diastolic |  | TERARECON | [9 empty] | manual | 25%
| art AbdPelThi |  | 120 | 120 | 180 | 120 |  |  | On | 9 | On | CARE Dose4D | 9.96 | 3s | 25% |  |  | 0.285 |  | 8 | 1 |  | 1.5 | 128x0.6mm | [5 empty] | ff800867-8571-4a15-997f-1d4ffbea68b7 | [9 empty] | Off
|  | art AbdPelThi a410 | [22 empty] | 1.5 | 0.8 | 313 | I26f medium smooth ASA | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | art AbdPelThi a410 | [22 empty] | 2 | 2 | 126 | I26f medium smooth ASA | CT Angio |  |  |  | GEPACS | [9 empty] | manual
|  | art AbdPelThi cor | [22 empty] | 2 | 2 | 176 | I26f medium smooth ASA | CT Angio |  |  |  | GEPACS | [9 empty] | manual
|  | art AbdPelThi sag | [22 empty] | 2 | 2 | 176 | I26f medium smooth ASA | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | art AbdPelThi cor mip | [22 empty] | 24 | 5 | 71 | I26f medium smooth ASA | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | art AbdPelThi a410 | [22 empty] | 0.6 | 0.6 | 417 | I26f medium smooth ASA | Abdomen |  |  |  | GEPACS | MCH3DLabAW | MCHNeuroBodyAW | [7 empty] | manual
| ven Abd a410 |  | 120 | 120 | 180 | 120 |  |  | On | 7 | On | CARE Dose4D | 9.95 | 3s | 25% |  |  | 0.5 |  | 18 | 1 |  | 3 | 128x0.6mm | [5 empty] | ff800867-8571-4a15-997f-1d4ffbea68b7 | [9 empty] | Off
|  | ven Abd a410 | [22 empty] | 3 | 3 | 84 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | ven Abd a410 | [22 empty] | 0.6 | 0.6 | 417 | I30f medium smooth | CT Angio |  |  |  | MCHNeuroBodyAW | [9 empty] | manual
|  | ven Abd cor | [22 empty] | 3 | 3 | 117 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
5125_ang_Che_pro_gat (Adult)
5125_ang_Che_pro_gat (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| PreMonitoring |  | 100 | 120 | 290 | 40 |  |  | Off | 11 | Off | CARE Dose4D | 2.04 |  | 25% |  |  | 0.285 |  | 2 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | PreMonitoring 10.0 B30f | [22 empty] | 10 |  | 1 | B30f medium smooth | Cardiac | [13 empty] | manual
| Monitoring |  | 100 | 120 | 290 | 40 |  |  | Off | 11 | Off | CARE Dose4D | 61.08 |  | 25% |  |  | 0.285 |  | 10 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | Monitoring 10.0 B30f | [22 empty] | 10 |  | 30 | B30f medium smooth | Cardiac | [13 empty] | manual
| art Che |  | 100 | 100 | 290 |  | 290 |  | Semi | 11 | On | CARE Dose4D | 6.32 |  | 25% | 50 |  | 0.285 |  | 10 | 0 | 34 | 0.75 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [7 empty] | 70% | 70% | Off | 70% | 70%
|  | art Che c125 :-D | [22 empty] | 0.75 |  | 413 | I30f medium smooth | Cardiac |  | Best Diastolic |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual | 70%
|  | art Che c125 | [22 empty] | 2 |  | 103 | I30f medium smooth | Cardiac |  |  |  | GEPACS | [9 empty] | manual | 70%
5126_ang_Che_ret_gat (Adult)
5126_ang_Che_ret_gat (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | ff800867-8571-4a15-997f-1d4ffbea68b7 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | ff800867-8571-4a15-997f-1d4ffbea68b7 |  |  | GEPACS |  | TRACKER
| PreMonitoring |  | 120 | 120 | 290 | 40 |  |  | Off | 9 | Off | CARE Dose4D | 2.04 |  | 25% |  |  | 0.285 |  | 2 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | PreMonitoring | [22 empty] | 10 |  | 1 | B30f medium smooth | Cardiac | [13 empty] | manual
| Monitoring |  | 120 | 120 | 290 | 40 |  |  | Off | 9 | Off | CARE Dose4D | 61.08 |  | 25% |  |  | 0.285 |  | 10 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | Monitoring | [22 empty] | 10 |  | 30 | B30f medium smooth | Cardiac | [13 empty] | manual
| Cor_Spiral |  | 100 | 100 | 290 |  | 220 |  | Semi | 11 | On | CARE Dose4D | 25.15 |  | 25% | 150 |  | 0.285 |  | 21 | 0.17 |  | 0.75 | 128x0.6mm | [5 empty] | ff800867-8571-4a15-997f-1d4ffbea68b7 | [9 empty] | Advance | 70% | 70%
|  | art Che c126 | [22 empty] | 0.75 | 0.5 | 401 | I30f medium smooth | Cardiac |  | Best Diastolic |  | MCH3DLabAW | GEPACS | SYNGO_VIA | [7 empty] | BestDiast | 70%
|  | art Che c126 | [22 empty] | 2 | 2 | 101 | I30f medium smooth | Cardiac |  | Best Diastolic |  | GEPACS | [9 empty] | BestDiast | 70%
5130_Che_1_wo_cst_T06_00A (Adult)
5130_Che_1_wo_cst_T06_00A (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Che |  | 120 | 120 | 150 | 110 |  |  | On | 7 | On | CARE Dose4D | 10.12 | 30s | 25% | 50 |  | 0.5 |  | 4 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | Che | [22 empty] | 3 | 2 | 145 | I41f medium | Chest |  |  |  | GEPACS | [9 empty] | manual
|  | Che sag | [22 empty] | 3 | 2 | 191 | I41f medium | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Che mip | [22 empty] | 10 | 3 | 121 | I41f medium | Lung |  |  |  | GEPACS | [9 empty] | manual
|  | Che c130 3dt | [22 empty] | 1 | 1 | 289 | I50f medium sharp ASA | Lung |  |  |  | ROUTER | GEPACS | [8 empty] | manual
5133_Che_1_Abd_wo_cst (Adult)
5133_Che_1_Abd_wo_cst (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Che |  | 120 | 120 | 150 | 110 |  |  | On | 7 | On | CARE Dose4D | 10.12 | 30s | 25% | 50 |  | 0.5 |  | 4 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | Che | [22 empty] | 3 | 2 | 145 | I41f medium | Chest |  |  |  | GEPACS | [9 empty] | manual
|  | Che sag | [22 empty] | 3 | 2 | 191 | I41f medium | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Che mip | [22 empty] | 10 | 3 | 121 | I41f medium | Lung |  |  |  | GEPACS | [9 empty] | manual
|  | Che c130 3dt | [22 empty] | 1 | 1 | 289 | I50f medium sharp ASA | Lung |  |  |  | ROUTER | GEPACS | [8 empty] | manual
| Abd |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 8 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | Abd a190 | [22 empty] | 3 | 3 | 131 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | Abd cor | [22 empty] | 3 | 3 | 67 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | Abd a190 :-D | [22 empty] | 0.6 | 0.6 | 653 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
5137_Che_1_AbdPel_wo_cst (Adult)
5137_Che_1_AbdPel_wo_cst (Adult) | Topogram | Topogram |  | 100 |  | 393 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 393 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Che |  | 120 | 120 | 150 | 110 |  |  | On | 7 | On | CARE Dose4D | 10.12 | 30s | 25% | 50 |  | 0.5 |  | 4 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | Che | [22 empty] | 3 | 2 | 145 | I41f medium | Chest |  |  |  | GEPACS | [9 empty] | manual
|  | Che sag | [22 empty] | 3 | 2 | 191 | I41f medium | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Che mip | [22 empty] | 10 | 3 | 121 | I41f medium | Lung |  |  |  | GEPACS | [9 empty] | manual
|  | Che c130 3dt | [22 empty] | 1 | 1 | 289 | I50f medium sharp ASA | Lung |  |  |  | ROUTER | GEPACS | [8 empty] | manual
| AbdPel |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 8 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | AbdPel a200 | [22 empty] | 3 | 3 | 131 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | AbdPel cor | [22 empty] | 3 | 3 | 67 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | AbdPel sag | [22 empty] | 3 | 3 | 121 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | AbdPel a200 :-D | [22 empty] | 0.6 | 0.6 | 653 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
5140_Che_2 (Adult)
5140_Che_2 (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| art Che |  | 120 | 120 | 150 | 110 |  |  | On | 7 | On | CARE Dose4D | 10.12 | 30s | 25% | 50 |  | 0.5 |  | 25 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | art Che | [22 empty] | 3 | 2 | 145 | I41f medium | Chest |  |  |  | GEPACS | [9 empty] | manual
|  | art Che sag | [22 empty] | 3 | 2 | 191 | I41f medium | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | art Che mip | [22 empty] | 10 | 3 | 121 | I41f medium | Lung |  |  |  | GEPACS | [9 empty] | manual
|  | art Che c140 3dt | [22 empty] | 1 | 1 | 289 | I50f medium sharp ASA | Lung |  |  |  | ROUTER | GEPACS | [8 empty] | manual
5150_Che_2_Abd (Adult)
5150_Che_2_Abd (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| art Che |  | 120 | 120 | 150 | 110 |  |  | On | 7 | On | CARE Dose4D | 10.12 | 30s | 25% | 50 |  | 0.5 |  | 25 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | art Che | [22 empty] | 3 | 2 | 145 | I41f medium | Chest |  |  |  | GEPACS | [9 empty] | manual
|  | art Che sag | [22 empty] | 3 | 2 | 191 | I41f medium | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | art Che mip | [22 empty] | 10 | 3 | 121 | I41f medium | Lung |  |  |  | GEPACS | [9 empty] | manual
|  | art Che c140 3dt | [22 empty] | 1 | 1 | 289 | I50f medium sharp ASA | Lung |  |  |  | ROUTER | GEPACS | [8 empty] | manual
| ven Abd |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 60 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven Abd a010 | [22 empty] | 3 | 3 | 131 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | ven Abd cor | [22 empty] | 3 | 3 | 67 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven Abd a010 :-D | [22 empty] | 0.6 | 0.6 | 653 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
5170_Che_2_AbdPel (Adult)
5170_Che_2_AbdPel (Adult) | Topogram | Topogram |  | 100 |  | 393 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 393 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| art Che |  | 120 | 120 | 150 | 110 |  |  | On | 7 | On | CARE Dose4D | 10.12 | 30s | 25% | 50 |  | 0.5 |  | 25 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | art Che | [22 empty] | 3 | 2 | 145 | I41f medium | Chest |  |  |  | GEPACS | [9 empty] | manual
|  | art Che sag | [22 empty] | 3 | 2 | 191 | I41f medium | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | art Che mip | [22 empty] | 10 | 3 | 121 | I41f medium | Lung |  |  |  | GEPACS | [9 empty] | manual
|  | art Che c140 3dt | [22 empty] | 1 | 1 | 289 | I50f medium sharp ASA | Lung |  |  |  | ROUTER | GEPACS | [8 empty] | manual
| ven AbdPel |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 60 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven AbdPel a030 | [22 empty] | 3 | 3 | 131 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel cor | [22 empty] | 3 | 3 | 67 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel sag | [22 empty] | 3 | 3 | 121 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel a030 :-D | [22 empty] | 0.6 | 0.6 | 653 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
5190_Che_2_AbdPel_ent (Adult)
5190_Che_2_AbdPel_ent (Adult) | Topogram | Topogram |  | 100 |  | 393 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 393 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| art Che |  | 120 | 120 | 150 | 110 |  |  | On | 7 | On | CARE Dose4D | 10.12 | 30s | 25% | 50 |  | 0.5 |  | 25 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | art Che | [22 empty] | 3 | 2 | 145 | I41f medium | Chest |  |  |  | GEPACS | [9 empty] | manual
|  | art Che sag | [22 empty] | 3 | 2 | 191 | I41f medium | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | art Che mip | [22 empty] | 10 | 3 | 121 | I41f medium | Lung |  |  |  | GEPACS | [9 empty] | manual
|  | art Che c140 3dt | [22 empty] | 1 | 1 | 289 | I50f medium sharp ASA | Lung |  |  |  | ROUTER | GEPACS | [8 empty] | manual
| ven AbdPel |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.14 | 30s | 25% | 50 |  | 0.5 |  | 40 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven AbdPel a230 | [22 empty] | 3 | 2 | 176 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel cor | [22 empty] | 3 | 2 | 139 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel sag | [22 empty] | 2 | 2 | 181 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel cor mip | [22 empty] | 16 | 5 | 73 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel a230 :-D | [22 empty] | 0.6 | 0.6 | 584 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
5200_Che_2_AbdPel_ho_tcc_w_fur (Adult)
5200_Che_2_AbdPel_ho_tcc_w_fur (Adult) | Topogram | Topogram |  | 100 |  | 393 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 393 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| art Che |  | 120 | 120 | 150 | 110 |  |  | On | 7 | On | CARE Dose4D | 10.12 | 30s | 25% | 50 |  | 0.5 |  | 25 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | art Che | [22 empty] | 3 | 2 | 145 | I41f medium | Chest |  |  |  | GEPACS | [9 empty] | manual
|  | art Che sag | [22 empty] | 3 | 2 | 191 | I41f medium | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | art Che mip | [22 empty] | 10 | 3 | 121 | I41f medium | Lung |  |  |  | GEPACS | [9 empty] | manual
|  | art Che c140 3dt | [22 empty] | 1 | 1 | 289 | I50f medium sharp ASA | Lung |  |  |  | ROUTER | GEPACS | [8 empty] | manual
| ven AbdPel |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 70 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven AbdPel a250 | [22 empty] | 3 | 3 | 104 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel cor | [22 empty] | 3 | 3 | 51 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel sag | [22 empty] | 3 | 3 | 121 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel a250 :-D | [22 empty] | 0.6 | 0.6 | 516 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
| 10m AbdPel |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 2 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | 10m AbdPel a250 | [22 empty] | 3 | 3 | 80 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | 10m AbdPel cor | [22 empty] | 3 | 3 | 67 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | 10m AbdPel cor 10x10 | [22 empty] | 10 | 10 | 39 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | 10m AbdPel cor mip | [22 empty] | 24 | 5 | 77 | I30f medium smooth | Abdomen MIP | [13 empty] | manual
|  | 10m AbdPel a250 :-D | [22 empty] | 0.6 | 0.6 | 400 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | [8 empty] | manual
5210_Che_2_AbdPel_uro_w_fur (Adult)
5210_Che_2_AbdPel_uro_w_fur (Adult) | Topogram | Topogram |  | 100 |  | 393 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 393 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Abd |  | 120 | 120 | 160 | 210 |  |  | On | 6 | On | CARE Dose4D | 10.8 | 30s | 25% | 50 |  | 0.5 |  | 2 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | Abd a300 | [22 empty] | 3 | 3 | 148 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | Abd a300 :-D | [22 empty] | 0.6 | 0.6 | 738 | I26f medium smooth ASA | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
| art Che |  | 120 | 120 | 150 | 110 |  |  | On | 7 | On | CARE Dose4D | 10.12 | 30s | 25% | 50 |  | 0.5 |  | 25 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | art Che | [22 empty] | 3 | 2 | 145 | I41f medium | Chest |  |  |  | GEPACS | [9 empty] | manual
|  | art Che sag | [22 empty] | 3 | 2 | 191 | I41f medium | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | art Che mip | [22 empty] | 10 | 3 | 121 | I41f medium | Lung |  |  |  | GEPACS | [9 empty] | manual
|  | art Che c140 3dt | [22 empty] | 1 | 1 | 289 | I50f medium sharp ASA | Lung |  |  |  | ROUTER | GEPACS | [8 empty] | manual
| ven AbdPel |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 70 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven AbdPel a300 | [22 empty] | 3 | 3 | 104 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel cor | [22 empty] | 3 | 3 | 51 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel sag | [22 empty] | 3 | 3 | 121 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel a300 :-D | [22 empty] | 0.6 | 0.6 | 516 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
| 10m AbdPel |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 2 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | 10m AbdPel a300 | [22 empty] | 3 | 3 | 80 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | 10m AbdPel cor | [22 empty] | 3 | 3 | 67 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | 10m AbdPel cor 10x10 | [22 empty] | 10 | 10 | 39 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | 10m AbdPel cor mip | [22 empty] | 24 | 5 | 77 | I30f medium smooth | Abdomen MIP |  |  |  | GEPACS | [9 empty] | manual
|  | 10m AbdPel a300 :-D | [22 empty] | 0.6 | 0.6 | 400 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | [8 empty] | manual
5211_Che_3 (Adult)
5211_Che_3 (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| ven Che |  | 120 | 120 | 150 | 110 |  |  | On | 7 | On | CARE Dose4D | 10.12 | 30s | 25% | 50 |  | 0.5 |  | 75 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven Che | [22 empty] | 3 | 2 | 145 | I41f medium | Chest |  |  |  | GEPACS | [9 empty] | manual
|  | ven Che sag | [22 empty] | 3 | 2 | 191 | I41f medium | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | ven Che mip | [22 empty] | 10 | 3 | 121 | I41f medium | Lung |  |  |  | GEPACS | [9 empty] | manual
|  | ven Che c211 3dt | [22 empty] | 1 | 1 | 289 | I50f medium sharp ASA | Lung |  |  |  | ROUTER | GEPACS | [8 empty] | manual
5213_Che_3_Abd (Adult)
5213_Che_3_Abd (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| ven Che |  | 120 | 120 | 150 | 110 |  |  | On | 7 | On | CARE Dose4D | 10.12 | 30s | 25% | 50 |  | 0.5 |  | 75 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven Che | [22 empty] | 3 | 2 | 145 | I41f medium | Chest |  |  |  | GEPACS | [9 empty] | manual
|  | ven Che sag | [22 empty] | 3 | 2 | 191 | I41f medium | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | ven Che mip | [22 empty] | 10 | 3 | 121 | I41f medium | Lung |  |  |  | GEPACS | [9 empty] | manual
|  | ven Che c211 3dt | [22 empty] | 1 | 1 | 289 | I50f medium sharp ASA | Lung |  |  |  | ROUTER | GEPACS | [8 empty] | manual
| ven Abd |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 10 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven Abd a010 | [22 empty] | 3 | 3 | 131 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | ven Abd cor | [22 empty] | 3 | 3 | 67 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven Abd a010 :-D | [22 empty] | 0.6 | 0.6 | 653 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
5215_Che_3_AbdPel (Adult)
5215_Che_3_AbdPel (Adult) | Topogram | Topogram |  | 100 |  | 393 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 393 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| ven Che |  | 120 | 120 | 150 | 110 |  |  | On | 7 | On | CARE Dose4D | 10.12 | 30s | 25% | 50 |  | 0.5 |  | 75 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven Che | [22 empty] | 3 | 2 | 145 | I41f medium | Chest |  |  |  | GEPACS | [9 empty] | manual
|  | ven Che sag | [22 empty] | 3 | 2 | 191 | I41f medium | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | ven Che mip | [22 empty] | 10 | 3 | 121 | I41f medium | Lung |  |  |  | GEPACS | [9 empty] | manual
|  | ven Che c211 3dt | [22 empty] | 1 | 1 | 289 | I50f medium sharp ASA | Lung |  |  |  | ROUTER | GEPACS | [8 empty] | manual
| ven AbdPel |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 10 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven AbdPel a030 | [22 empty] | 3 | 3 | 131 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel cor | [22 empty] | 3 | 3 | 67 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel sag | [22 empty] | 3 | 3 | 121 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel a030 :-D | [22 empty] | 0.6 | 0.6 | 653 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
5217_Che_3_AbdPel_ent (Adult)
5217_Che_3_AbdPel_ent (Adult) | Topogram | Topogram |  | 100 |  | 393 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 393 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| ven Che |  | 120 | 120 | 150 | 110 |  |  | On | 7 | On | CARE Dose4D | 10.12 | 30s | 25% | 50 |  | 0.5 |  | 55 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven Che | [22 empty] | 3 | 2 | 145 | I41f medium | Chest |  |  |  | GEPACS | [9 empty] | manual
|  | ven Che sag | [22 empty] | 3 | 2 | 191 | I41f medium | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | ven Che mip | [22 empty] | 10 | 3 | 121 | I41f medium | Lung |  |  |  | GEPACS | [9 empty] | manual
|  | ven Che c211 3dt | [22 empty] | 1 | 1 | 289 | I50f medium sharp ASA | Lung |  |  |  | ROUTER | GEPACS | [8 empty] | manual
| ven AbdPel |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.14 | 30s | 25% | 50 |  | 0.5 |  | 10 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven AbdPel a230 | [22 empty] | 3 | 2 | 176 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel cor | [22 empty] | 3 | 2 | 139 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel sag | [22 empty] | 2 | 2 | 181 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel cor mip | [22 empty] | 16 | 5 | 73 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel a230 :-D | [22 empty] | 0.6 | 0.6 | 584 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
5221_Che_3_AbdPel_ho_tcc_w_fur (Adult)
5221_Che_3_AbdPel_ho_tcc_w_fur (Adult) | Topogram | Topogram |  | 100 |  | 393 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 393 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| ven Che |  | 120 | 120 | 150 | 110 |  |  | On | 7 | On | CARE Dose4D | 10.12 | 30s | 25% | 50 |  | 0.5 |  | 75 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven Che | [22 empty] | 3 | 2 | 145 | I41f medium | Chest |  |  |  | GEPACS | [9 empty] | manual
|  | ven Che sag | [22 empty] | 3 | 2 | 191 | I41f medium | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | ven Che mip | [22 empty] | 10 | 3 | 121 | I41f medium | Lung |  |  |  | GEPACS | [9 empty] | manual
|  | ven Che c211 3dt | [22 empty] | 1 | 1 | 289 | I50f medium sharp ASA | Lung |  |  |  | ROUTER | GEPACS | [8 empty] | manual
| ven AbdPel |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 20 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven AbdPel a250 | [22 empty] | 3 | 3 | 104 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel cor | [22 empty] | 3 | 3 | 51 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel sag | [22 empty] | 3 | 3 | 121 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel a250 :-D | [22 empty] | 0.6 | 0.6 | 516 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
| 10m AbdPel |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 2 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | 10m AbdPel a250 | [22 empty] | 3 | 3 | 80 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | 10m AbdPel cor | [22 empty] | 3 | 3 | 67 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | 10m AbdPel cor 10x10 | [22 empty] | 10 | 10 | 39 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | 10m AbdPel cor mip | [22 empty] | 24 | 5 | 77 | I30f medium smooth | Abdomen MIP |  |  |  | GEPACS | [9 empty] | manual
|  | 10m AbdPel a250 :-D | [22 empty] | 0.6 | 0.6 | 400 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | [8 empty] | manual
5224_Che_3_AbdPel_uro_w_fur (Adult)
5224_Che_3_AbdPel_uro_w_fur (Adult) | Topogram | Topogram |  | 100 |  | 393 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 393 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Abd |  | 120 | 120 | 160 | 210 |  |  | On | 6 | On | CARE Dose4D | 10.8 | 30s | 25% | 50 |  | 0.5 |  | 2 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | Abd a300 | [22 empty] | 3 | 3 | 148 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | Abd a300 :-D | [22 empty] | 0.6 | 0.6 | 738 | I26f medium smooth ASA | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
| ven Che |  | 120 | 120 | 150 | 110 |  |  | On | 7 | On | CARE Dose4D | 10.12 | 30s | 25% | 50 |  | 0.5 |  | 75 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven Che | [22 empty] | 3 | 2 | 145 | I41f medium | Chest |  |  |  | GEPACS | [9 empty] | manual
|  | ven Che sag | [22 empty] | 3 | 2 | 191 | I41f medium | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | ven Che mip | [22 empty] | 10 | 3 | 121 | I41f medium | Lung |  |  |  | GEPACS | [9 empty] | manual
|  | ven Che c211 3dt | [22 empty] | 1 | 1 | 289 | I50f medium sharp ASA | Lung |  |  |  | ROUTER | GEPACS | [8 empty] | manual
| ven AbdPel |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 20 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven AbdPel a300 | [22 empty] | 3 | 3 | 104 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel cor | [22 empty] | 3 | 3 | 51 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel sag | [22 empty] | 3 | 3 | 121 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel a300 :-D | [22 empty] | 0.6 | 0.6 | 516 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
| 10m AbdPel |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 2 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | 10m AbdPel a300 | [22 empty] | 3 | 3 | 80 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | 10m AbdPel cor | [22 empty] | 3 | 3 | 67 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | 10m AbdPel cor 10x10 | [22 empty] | 10 | 10 | 39 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | 10m AbdPel cor mip | [22 empty] | 24 | 5 | 77 | I30f medium smooth | Abdomen MIP |  |  |  | GEPACS | [9 empty] | manual
|  | 10m AbdPel a300 :-D | [22 empty] | 0.6 | 0.6 | 400 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | [8 empty] | manual
5230_Che_3_Abd_liv_bi (Adult)
5230_Che_3_Abd_liv_bi (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| PreMonitoring |  | 120 | 120 | 200 | 20 |  |  | Off | 7 | Off | CARE Dose4D | 1.23 |  | 25% |  |  | 0.5 |  | 2 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | PreMonitoring | [22 empty] | 10 |  | 1 | B30s medium smooth | Abdomen | [13 empty] | manual
| Monitoring |  | 120 | 120 | 200 | 20 |  |  | Off | 7 | Off | CARE Dose4D | 36.98 |  | 25% |  |  | 0.5 |  | 12 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | Monitoring | [22 empty] | 10 |  | 30 | B30s medium smooth | Abdomen | [13 empty] | manual
| art Abd |  | 120 | 120 | 180 | 110 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 16 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | art Abd a070 | [22 empty] | 3 | 2 | 145 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd a070 :-D | [22 empty] | 0.6 | 0.6 | 481 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
| ven Che |  | 120 | 120 | 150 | 110 |  |  | On | 7 | On | CARE Dose4D | 10.12 | 30s | 25% | 50 |  | 0.5 |  | 2 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven Che | [22 empty] | 3 | 2 | 145 | I41f medium | Chest |  |  |  | GEPACS | [9 empty] | manual
|  | ven Che sag | [22 empty] | 3 | 2 | 191 | I41f medium | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | ven Che mip | [22 empty] | 10 | 3 | 121 | I41f medium | Lung |  |  |  | GEPACS | [9 empty] | manual
|  | ven Che c220 3dt | [22 empty] | 1 | 1 | 289 | I50f medium sharp ASA | Lung |  |  |  | ROUTER | GEPACS | [8 empty] | manual
| ven Abd |  | 120 | 120 | 240 | 110 |  |  | On | 6 | On | CARE Dose4D | 16.2 | 30s | 25% | 50 |  | 0.5 |  | 40 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven Abd a070 | [22 empty] | 3 | 3 | 97 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | ven Abd cor | [22 empty] | 3 | 3 | 127 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | ven Abd a070 :-D | [22 empty] | 0.6 | 0.6 | 481 | I30f medium smooth | Abdomen |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
5240_Che_3_Abd_liv_bi_Pel (Adult)
5240_Che_3_Abd_liv_bi_Pel (Adult) | Topogram | Topogram |  | 100 |  | 393 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 393 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| PreMonitoring |  | 120 | 120 | 200 | 20 |  |  | Off | 7 | Off | CARE Dose4D | 1.23 |  | 25% |  |  | 0.5 |  | 2 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | PreMonitoring | [22 empty] | 10 |  | 1 | B30s medium smooth | Abdomen | [13 empty] | manual
| Monitoring |  | 120 | 120 | 200 | 20 |  |  | Off | 7 | Off | CARE Dose4D | 36.98 |  | 25% |  |  | 0.5 |  | 12 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | Monitoring | [22 empty] | 10 |  | 30 | B30s medium smooth | Abdomen | [13 empty] | manual
| art Abd |  | 120 | 120 | 180 | 110 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 16 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | art Abd a070 | [22 empty] | 3 | 2 | 145 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd a070 :-D | [22 empty] | 0.6 | 0.6 | 481 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
| ven Che |  | 120 | 120 | 150 | 110 |  |  | On | 7 | On | CARE Dose4D | 10.12 | 30s | 25% | 50 |  | 0.5 |  | 2 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven Che | [22 empty] | 3 | 2 | 145 | I41f medium | Chest |  |  |  | GEPACS | [9 empty] | manual
|  | ven Che sag | [22 empty] | 3 | 2 | 191 | I41f medium | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | ven Che mip | [22 empty] | 10 | 3 | 121 | I41f medium | Lung |  |  |  | GEPACS | [9 empty] | manual
|  | ven Che c220 3dt | [22 empty] | 1 | 1 | 289 | I50f medium sharp ASA | Lung |  |  |  | ROUTER | GEPACS | [8 empty] | manual
| ven AbdPel |  | 120 | 120 | 240 | 110 |  |  | On | 6 | On | CARE Dose4D | 16.2 | 30s | 25% | 50 |  | 0.5 |  | 40 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven AbdPel a080 | [22 empty] | 3 | 3 | 97 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel cor | [22 empty] | 3 | 3 | 127 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel sag | [22 empty] | 3 | 3 | 127 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel a080 :-D | [22 empty] | 0.6 | 0.6 | 481 | I30f medium smooth | Abdomen |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
5255_Che_3_Abd_liv_tri (Adult)
5255_Che_3_Abd_liv_tri (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| PreMonitoring |  | 120 | 120 | 200 | 20 |  |  | Off | 7 | Off | CARE Dose4D | 1.23 |  | 25% |  |  | 0.5 |  | 2 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | PreMonitoring | [22 empty] | 10 |  | 1 | B30s medium smooth | Abdomen | [13 empty] | manual
| Monitoring |  | 120 | 120 | 200 | 20 |  |  | Off | 7 | Off | CARE Dose4D | 36.98 |  | 25% |  |  | 0.5 |  | 12 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | Monitoring | [22 empty] | 10 |  | 30 | B30s medium smooth | Abdomen | [13 empty] | manual
| art Abd |  | 120 | 120 | 180 | 110 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 16 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | art Abd a090 | [22 empty] | 3 | 2 | 145 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd cor | [22 empty] | 3 | 3 | 127 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd sag | [22 empty] | 3 | 3 | 127 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd a090 :-D | [22 empty] | 0.6 | 0.6 | 481 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
| ven Che |  | 120 | 120 | 150 | 110 |  |  | On | 7 | On | CARE Dose4D | 10.12 | 30s | 25% | 50 |  | 0.5 |  | 2 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven Che | [22 empty] | 3 | 2 | 145 | I41f medium | Chest |  |  |  | GEPACS | [9 empty] | manual
|  | ven Che sag | [22 empty] | 3 | 2 | 191 | I41f medium | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | ven Che mip | [22 empty] | 10 | 3 | 121 | I41f medium | Lung |  |  |  | GEPACS | [9 empty] | manual
|  | ven Che c220 3dt | [22 empty] | 1 | 1 | 289 | I50f medium sharp ASA | Lung |  |  |  | ROUTER | GEPACS | [8 empty] | manual
| ven Abd |  | 120 | 120 | 180 | 110 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 40 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven Abd a090 | [22 empty] | 3 | 3 | 97 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | ven Abd cor | [22 empty] | 3 | 3 | 127 | I30f medium smooth | Abdomen | [13 empty] | manual
|  | ven Abd a090 :-D | [22 empty] | 0.6 | 0.6 | 481 | I30f medium smooth | Abdomen |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
| 5m Abd |  | 120 | 120 | 180 | 110 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 2 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | 5m Abd a090 | [22 empty] | 3 | 3 | 97 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | 5m Abd a090 :-D | [22 empty] | 0.6 | 0.6 | 481 | I30f medium smooth | Abdomen |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
5270_Che_3_Abd_liv_tri_Pel (Adult)
5270_Che_3_Abd_liv_tri_Pel (Adult) | Topogram | Topogram |  | 100 |  | 393 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 393 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| PreMonitoring |  | 120 | 120 | 200 | 20 |  |  | Off | 7 | Off | CARE Dose4D | 1.23 |  | 25% |  |  | 0.5 |  | 2 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | PreMonitoring | [22 empty] | 10 |  | 1 | B30s medium smooth | Abdomen | [13 empty] | manual
| Monitoring |  | 120 | 120 | 200 | 20 |  |  | Off | 7 | Off | CARE Dose4D | 36.98 |  | 25% |  |  | 0.5 |  | 12 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | Monitoring | [22 empty] | 10 |  | 30 | B30s medium smooth | Abdomen | [13 empty] | manual
| art Abd |  | 120 | 120 | 180 | 110 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 16 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | art Abd a110 :-D | [22 empty] | 3 | 2 | 145 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd cor | [22 empty] | 3 | 3 | 127 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd sag | [22 empty] | 3 | 3 | 127 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd a110 | [22 empty] | 0.6 | 0.6 | 481 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
| ven Che |  | 120 | 120 | 150 | 110 |  |  | On | 7 | On | CARE Dose4D | 10.12 | 30s | 25% | 50 |  | 0.5 |  | 2 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven Che | [22 empty] | 3 | 2 | 145 | I41f medium | Chest |  |  |  | GEPACS | [9 empty] | manual
|  | ven Che sag | [22 empty] | 3 | 2 | 191 | I41f medium | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | ven Che mip | [22 empty] | 10 | 3 | 121 | I41f medium | Lung |  |  |  | GEPACS | [9 empty] | manual
|  | ven Che c220 3dt | [22 empty] | 1 | 1 | 289 | I50f medium sharp ASA | Lung |  |  |  | ROUTER | GEPACS | [8 empty] | manual
| ven AbdPel |  | 120 | 120 | 180 | 110 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 40 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven AbdPel a110 | [22 empty] | 3 | 3 | 97 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | ven Abd cor | [22 empty] | 3 | 3 | 127 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | ven Abd sag | [22 empty] | 3 | 3 | 127 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel a110 :-D | [22 empty] | 0.6 | 0.6 | 481 | I30f medium smooth | Abdomen |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
| 5m Abd |  | 120 | 120 | 180 | 110 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 2 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | 5m Abd a110 | [22 empty] | 3 | 3 | 97 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | 5m Abd a110 :-D | [22 empty] | 0.6 | 0.6 | 481 | I30f medium smooth | Abdomen |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
5285_Che_3_Abd_liv_qua_pos_TACE (Adult)
5285_Che_3_Abd_liv_qua_pos_TACE (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Abd |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 2 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | Abd a110 | [22 empty] | 3 | 3 | 131 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | ven Abd cor | [22 empty] | 3 | 3 | 67 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | Abd a110 :-D | [22 empty] | 0.6 | 0.6 | 653 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
| PreMonitoring |  | 120 | 120 | 200 | 20 |  |  | Off | 7 | Off | CARE Dose4D | 1.23 |  | 25% |  |  | 0.5 |  | 4 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | PreMonitoring | [22 empty] | 10 |  | 1 | B30s medium smooth | Abdomen | [13 empty] | manual
| Monitoring |  | 120 | 120 | 200 | 20 |  |  | Off | 7 | Off | CARE Dose4D | 36.98 |  | 25% |  |  | 0.5 |  | 12 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | Monitoring | [22 empty] | 10 |  | 30 | B30s medium smooth | Abdomen | [13 empty] | manual
| art Abd |  | 120 | 120 | 180 | 110 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 16 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | art Abd a110 | [22 empty] | 3 | 2 | 145 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd cor | [22 empty] | 3 | 3 | 127 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd sag | [22 empty] | 3 | 3 | 127 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd a110 :-D | [22 empty] | 0.6 | 0.6 | 481 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
| ven Che |  | 120 | 120 | 150 | 110 |  |  | On | 7 | On | CARE Dose4D | 10.12 | 30s | 25% | 50 |  | 0.5 |  | 2 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven Che | [22 empty] | 3 | 2 | 145 | I41f medium | Chest |  |  |  | GEPACS | [9 empty] | manual
|  | ven Che sag | [22 empty] | 3 | 2 | 191 | I41f medium | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | ven Che mip | [22 empty] | 10 | 3 | 121 | I41f medium | Lung |  |  |  | GEPACS | [9 empty] | manual
|  | ven Che c220 3dt | [22 empty] | 1 | 1 | 289 | I50f medium sharp ASA | Lung |  |  |  | ROUTER | GEPACS | [8 empty] | manual
| ven Abd |  | 120 | 120 | 180 | 110 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 40 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven Abd a110 | [22 empty] | 3 | 3 | 97 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | ven Abd cor | [22 empty] | 3 | 3 | 127 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | ven Abd a110 :-D | [22 empty] | 0.6 | 0.6 | 481 | I30f medium smooth | Abdomen |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
| 5m Abd |  | 120 | 120 | 180 | 110 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 2 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | 5m Abd a110 | [22 empty] | 3 | 3 | 97 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | 5m Abd a110 :-D | [22 empty] | 0.6 | 0.6 | 481 | I30f medium smooth | Abdomen |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
5286_Che_3_Abd_liv_qua_pos_TACE_Pel (Adult)
5286_Che_3_Abd_liv_qua_pos_TACE_Pel (Adult) | Topogram | Topogram |  | 100 |  | 393 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 393 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Abd |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 2 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | Abd a110 | [22 empty] | 3 | 3 | 131 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | ven Abd cor | [22 empty] | 3 | 3 | 67 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | Abd a110 :-D | [22 empty] | 0.6 | 0.6 | 653 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
| PreMonitoring |  | 120 | 120 | 200 | 20 |  |  | Off | 7 | Off | CARE Dose4D | 1.23 |  | 25% |  |  | 0.5 |  | 4 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | PreMonitoring | [22 empty] | 10 |  | 1 | B30s medium smooth | Abdomen | [13 empty] | manual
| Monitoring |  | 120 | 120 | 200 | 20 |  |  | Off | 7 | Off | CARE Dose4D | 36.98 |  | 25% |  |  | 0.5 |  | 12 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | Monitoring | [22 empty] | 10 |  | 30 | B30s medium smooth | Abdomen | [13 empty] | manual
| art Abd |  | 120 | 120 | 180 | 110 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 16 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | art Abd a110 | [22 empty] | 3 | 2 | 145 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd cor | [22 empty] | 3 | 3 | 127 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd sag | [22 empty] | 3 | 3 | 127 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd a110 :-D | [22 empty] | 0.6 | 0.6 | 481 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
| ven Che |  | 120 | 120 | 150 | 110 |  |  | On | 7 | On | CARE Dose4D | 10.12 | 30s | 25% | 50 |  | 0.5 |  | 2 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven Che | [22 empty] | 3 | 2 | 145 | I41f medium | Chest |  |  |  | GEPACS | [9 empty] | manual
|  | ven Che sag | [22 empty] | 3 | 2 | 191 | I41f medium | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | ven Che mip | [22 empty] | 10 | 3 | 121 | I41f medium | Lung |  |  |  | GEPACS | [9 empty] | manual
|  | ven Che c220 3dt | [22 empty] | 1 | 1 | 289 | I50f medium sharp ASA | Lung |  |  |  | ROUTER | GEPACS | [8 empty] | manual
| ven AbdPel |  | 120 | 120 | 180 | 110 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 40 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven AbdPel a110 | [22 empty] | 3 | 3 | 97 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel cor | [22 empty] | 3 | 3 | 127 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel sag | [22 empty] | 3 | 3 | 127 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel a110 :-D | [22 empty] | 0.6 | 0.6 | 481 | I30f medium smooth | Abdomen |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
| 5m Abd |  | 120 | 120 | 180 | 110 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 2 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | 5m Abd a110 | [22 empty] | 3 | 3 | 97 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | 5m Abd a110 :-D | [22 empty] | 0.6 | 0.6 | 481 | I30f medium smooth | Abdomen |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
5290_Che_3_Abd_pan_bi (Adult)
5290_Che_3_Abd_pan_bi (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| PreMonitoring |  | 120 | 120 | 200 | 20 |  |  | Off | 7 | Off | CARE Dose4D | 1.23 |  | 25% |  |  | 0.5 |  | 2 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | PreMonitoring | [22 empty] | 10 |  | 1 | B30s medium smooth | Abdomen | [13 empty] | manual
| Monitoring |  | 120 | 120 | 200 | 20 |  |  | Off | 7 | Off | CARE Dose4D | 36.98 |  | 25% |  |  | 0.5 |  | 10 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | Monitoring | [22 empty] | 10 |  | 30 | B30s medium smooth | Abdomen | [13 empty] | manual
| art Abd |  | 120 | 120 | 180 | 110 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 19 | 0.8 |  | 2 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | art Abd a130 | [22 empty] | 2 | 2 | 145 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd cor | [22 empty] | 2 | 2 | 191 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd a130 :-D | [22 empty] | 0.6 | 0.6 | 481 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
| ven Che |  | 120 | 120 | 150 | 110 |  |  | On | 7 | On | CARE Dose4D | 10.12 | 30s | 25% | 50 |  | 0.5 |  | 6 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven Che | [22 empty] | 3 | 2 | 145 | I41f medium | Chest |  |  |  | GEPACS | [9 empty] | manual
|  | ven Che sag | [22 empty] | 3 | 2 | 191 | I41f medium | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | ven Che mip | [22 empty] | 10 | 3 | 121 | I41f medium | Lung |  |  |  | GEPACS | [9 empty] | manual
|  | ven Che c220 3dt | [22 empty] | 1 | 1 | 289 | I50f medium sharp ASA | Lung |  |  |  | ROUTER | GEPACS | [8 empty] | manual
| ven Abd |  | 120 | 120 | 180 | 110 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 50 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven Abd a130 | [22 empty] | 3 | 3 | 97 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | ven Abd cor | [22 empty] | 3 | 3 | 127 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | ven Abd a130 :-D | [22 empty] | 0.6 | 0.6 | 481 | I30f medium smooth | Abdomen |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
5295_Che_3_Abd_pan_bi_mod (Adult)
5295_Che_3_Abd_pan_bi_mod (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| ven Che |  | 120 | 120 | 150 | 110 |  |  | On | 7 | On | CARE Dose4D | 10.12 | 30s | 25% | 50 |  | 0.5 |  | 25 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven Che | [22 empty] | 3 | 2 | 145 | I41f medium | Chest |  |  |  | GEPACS | [9 empty] | manual
|  | ven Che sag | [22 empty] | 3 | 2 | 191 | I41f medium | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | ven Che mip | [22 empty] | 10 | 3 | 121 | I41f medium | Lung |  |  |  | GEPACS | [9 empty] | manual
|  | ven Che c220 3dt | [22 empty] | 1 | 1 | 289 | I50f medium sharp ASA | Lung |  |  |  | ROUTER | GEPACS | [8 empty] | manual
| artven Abd |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 45 | 0.8 |  | 2 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | artven Abd a135 | [22 empty] | 2 | 2 | 135 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | artven Abd cor | [22 empty] | 2 | 2 | 181 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | artven Abd a135 :-D | [22 empty] | 0.6 | 0.6 | 447 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
5300_Che_3_Abd_pan_bi_Pel (Adult)
5300_Che_3_Abd_pan_bi_Pel (Adult) | Topogram | Topogram |  | 100 |  | 393 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 393 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| PreMonitoring |  | 120 | 120 | 200 | 20 |  |  | Off | 7 | Off | CARE Dose4D | 1.23 |  | 25% |  |  | 0.5 |  | 2 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | PreMonitoring | [22 empty] | 10 |  | 1 | B30s medium smooth | Abdomen | [13 empty] | manual
| Monitoring |  | 120 | 120 | 200 | 20 |  |  | Off | 7 | Off | CARE Dose4D | 36.98 |  | 25% |  |  | 0.5 |  | 10 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | Monitoring | [22 empty] | 10 |  | 30 | B30s medium smooth | Abdomen | [13 empty] | manual
| art Abd |  | 120 | 120 | 180 | 110 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 19 | 0.8 |  | 2 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | art Abd a140 | [22 empty] | 2 | 2 | 145 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd cor | [22 empty] | 2 | 2 | 191 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd a140 :-D | [22 empty] | 0.6 | 0.6 | 481 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
| ven Che |  | 120 | 120 | 150 | 110 |  |  | On | 7 | On | CARE Dose4D | 10.12 | 30s | 25% | 50 |  | 0.5 |  | 6 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven Che | [22 empty] | 3 | 2 | 145 | I41f medium | Chest |  |  |  | GEPACS | [9 empty] | manual
|  | ven Che sag | [22 empty] | 3 | 2 | 191 | I41f medium | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | ven Che mip | [22 empty] | 10 | 3 | 121 | I41f medium | Lung |  |  |  | GEPACS | [9 empty] | manual
|  | ven Che c220 3dt | [22 empty] | 1 | 1 | 289 | I50f medium sharp ASA | Lung |  |  |  | ROUTER | GEPACS | [8 empty] | manual
| ven AbdPel |  | 120 | 120 | 180 | 110 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 50 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven AbdPel a140 | [22 empty] | 3 | 3 | 97 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel cor | [22 empty] | 3 | 3 | 127 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel sag | [22 empty] | 3 | 3 | 127 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel a140 :-D | [22 empty] | 0.6 | 0.6 | 481 | I30f medium smooth | Abdomen |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
5305_Che_3_Abd_pan_bi_mod_Pel (Adult)
5305_Che_3_Abd_pan_bi_mod_Pel (Adult) | Topogram | Topogram |  | 100 |  | 393 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 393 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| ven Che |  | 120 | 120 | 150 | 110 |  |  | On | 7 | On | CARE Dose4D | 10.12 | 30s | 25% | 50 |  | 0.5 |  | 25 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven Che | [22 empty] | 3 | 2 | 145 | I41f medium | Chest |  |  |  | GEPACS | [9 empty] | manual
|  | ven Che sag | [22 empty] | 3 | 2 | 191 | I41f medium | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | ven Che mip | [22 empty] | 10 | 3 | 121 | I41f medium | Lung |  |  |  | GEPACS | [9 empty] | manual
|  | ven Che c220 3dt | [22 empty] | 1 | 1 | 289 | I50f medium sharp ASA | Lung |  |  |  | ROUTER | GEPACS | [8 empty] | manual
| artven AbdPel |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 45 | 0.8 |  | 2 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | artven AbdPel a145 | [22 empty] | 2 | 2 | 135 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | artven AbdPel cor | [22 empty] | 2 | 2 | 181 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | artven AbdPel sag | [22 empty] | 2 | 2 | 181 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | artven AbdPel a145 :-D | [22 empty] | 0.6 | 0.6 | 447 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | [8 empty] | manual
5310_Che_3_ang_Abd_pan_bi (Adult)
5310_Che_3_ang_Abd_pan_bi (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| PreMonitoring |  | 120 | 120 | 350 | 50 |  |  | Off | 3 | Off | CARE Dose4D | 3.08 |  | 25% |  |  | 0.5 |  | 2 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | PreMonitoring | [22 empty] | 10 |  | 1 | B40s medium | 400/40 Abd | [13 empty] | manual
| Monitoring |  | 120 | 120 | 350 | 50 |  |  | Off | 3 | Off | CARE Dose4D | 123.28 |  | 25% |  |  | 0.5 |  | 10 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | Monitoring | [22 empty] | 10 |  | 40 | B40s medium | 400/40 Abd | [13 empty] | manual
| art Abd |  | 120 | 80 | 700 | 210 |  |  | Off | 8 | On | CARE Dose4D | 12.7 | 30s | 25% | 50 |  | 0.5 |  | 18 | 0.5 |  | 2 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | art Abd a350 | [22 empty] | 2 | 1 | 240 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd cor | [22 empty] | 2 | 2 | 191 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd cor mip | [22 empty] | 24 | 5 | 77 | I30f medium smooth | Abdomen MIP |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd a350 :-D | [22 empty] | 0.6 | 0.6 | 400 | I30f medium smooth | Abdomen |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
| ven Abd |  | 120 | 80 | 700 | 210 |  |  | Off | 6 | On | CARE Dose4D | 12.7 | 30s | 25% | 50 |  | 0.5 |  | 15 | 0.5 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven Abd a350 | [22 empty] | 3 | 3 | 80 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven Abd cor | [22 empty] | 3 | 3 | 127 | I30f medium smooth | 400/40 Abd | [13 empty] | manual
|  | ven Abd cor mip | [22 empty] | 24 | 5 | 77 | I30f medium smooth | Abdomen MIP |  |  |  | GEPACS | [9 empty] | manual
|  | ven Abd a350 :-D | [22 empty] | 0.6 | 0.6 | 400 | I30f medium smooth | Abdomen |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
| ven Che |  | 120 | 120 | 150 | 110 |  |  | On | 7 | On | CARE Dose4D | 10.12 | 30s | 25% | 50 |  | 0.5 |  | 6 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven Che | [22 empty] | 3 | 2 | 145 | I41f medium | Chest |  |  |  | GEPACS | [9 empty] | manual
|  | ven Che sag | [22 empty] | 3 | 2 | 191 | I41f medium | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | ven Che mip | [22 empty] | 10 | 3 | 121 | I41f medium | Lung |  |  |  | GEPACS | [9 empty] | manual
|  | ven Che c220 3dt | [22 empty] | 1 | 1 | 289 | I50f medium sharp ASA | Lung |  |  |  | ROUTER | GEPACS | [8 empty] | manual
5320_Che_3_ang_Abd_pan_bi_Pel (Adult)
5320_Che_3_ang_Abd_pan_bi_Pel (Adult) | Topogram | Topogram |  | 100 |  | 393 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 393 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| PreMonitoring |  | 120 | 120 | 350 | 50 |  |  | Off | 3 | Off | CARE Dose4D | 3.08 |  | 25% |  |  | 0.5 |  | 2 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | PreMonitoring | [22 empty] | 10 |  | 1 | B40s medium | 400/40 Abd | [13 empty] | manual
| Monitoring |  | 120 | 120 | 350 | 50 |  |  | Off | 3 | Off | CARE Dose4D | 123.28 |  | 25% |  |  | 0.5 |  | 10 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | Monitoring | [22 empty] | 10 |  | 40 | B40s medium | 400/40 Abd | [13 empty] | manual
| art Abd |  | 120 | 80 | 700 | 210 |  |  | Off | 8 | On | CARE Dose4D | 12.7 | 30s | 25% | 50 |  | 0.5 |  | 18 | 0.5 |  | 2 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | art Abd a360 | [22 empty] | 2 | 1 | 240 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd cor | [22 empty] | 2 | 2 | 191 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd cor mip | [22 empty] | 24 | 5 | 77 | I30f medium smooth | Abdomen MIP |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd a360 :-D | [22 empty] | 0.6 | 0.6 | 400 | I30f medium smooth | Abdomen |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
| ven AbdPel |  | 120 | 80 | 700 | 210 |  |  | Off | 6 | On | CARE Dose4D | 12.7 | 30s | 25% | 50 |  | 0.5 |  | 15 | 0.5 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven AbdPel a360 | [22 empty] | 3 | 3 | 80 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel cor | [22 empty] | 3 | 3 | 127 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel cor mip | [22 empty] | 24 | 5 | 77 | I30f medium smooth | Abdomen MIP |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel sag | [22 empty] | 3 | 3 | 127 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel a360 :-D | [22 empty] | 0.6 | 0.6 | 400 | I30f medium smooth | Abdomen |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
| ven Che |  | 120 | 120 | 150 | 110 |  |  | On | 7 | On | CARE Dose4D | 10.12 | 30s | 25% | 50 |  | 0.5 |  | 6 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven Che | [22 empty] | 3 | 2 | 145 | I41f medium | Chest |  |  |  | GEPACS | [9 empty] | manual
|  | ven Che sag | [22 empty] | 3 | 2 | 191 | I41f medium | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | ven Che mip | [22 empty] | 10 | 3 | 121 | I41f medium | Lung |  |  |  | GEPACS | [9 empty] | manual
|  | ven Che c220 3dt | [22 empty] | 1 | 1 | 289 | I50f medium sharp ASA | Lung |  |  |  | ROUTER | GEPACS | [8 empty] | manual
5322_Che_eso_can_wo_cst (Adult)
5322_Che_eso_can_wo_cst (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Che |  | 120 | 120 | 150 | 110 |  |  | On | 7 | On | CARE Dose4D | 10.12 | 30s | 25% | 50 |  | 0.5 |  | 25 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | Che | [22 empty] | 3 | 2 | 145 | I41f medium | Chest |  |  |  | GEPACS | [9 empty] | manual
|  | Che sag | [22 empty] | 3 | 2 | 191 | I41f medium | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Che mip | [22 empty] | 10 | 3 | 121 | I41f medium | Lung |  |  |  | GEPACS | [9 empty] | manual
|  | Che c130 3dt | [22 empty] | 1 | 1 | 289 | I50f medium sharp ASA | Lung |  |  |  | ROUTER | GEPACS | [8 empty] | manual
5324_Che_3_eso_can (Adult)
5324_Che_3_eso_can (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| ven Che |  | 120 | 120 | 150 | 110 |  |  | On | 7 | On | CARE Dose4D | 10.12 | 30s | 25% | 50 |  | 0.5 |  | 75 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven Che | [22 empty] | 3 | 2 | 145 | I41f medium | Chest |  |  |  | GEPACS | [9 empty] | manual
|  | ven Che sag | [22 empty] | 3 | 2 | 191 | I41f medium | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | ven Che sag | [22 empty] | 2 | 2 | 191 | I41f medium | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | ven Che c324 :-D | [22 empty] | 0.6 | 0.6 | 481 | I41f medium | 1500/600 Lung |  |  |  | ROUTER | GEPACS | [8 empty] | manual
5326_Che_3_eso_can_Abd (Adult)
5326_Che_3_eso_can_Abd (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| ven Che |  | 120 | 120 | 150 | 110 |  |  | On | 7 | On | CARE Dose4D | 10.12 | 30s | 25% | 50 |  | 0.5 |  | 75 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven Che | [22 empty] | 3 | 2 | 145 | I41f medium | Chest |  |  |  | GEPACS | [9 empty] | manual
|  | ven Che sag | [22 empty] | 3 | 2 | 191 | I41f medium | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | ven Che mip | [22 empty] | 10 | 3 | 121 | I41f medium | Lung |  |  |  | GEPACS | [9 empty] | manual
|  | ven Che c220 3dt | [22 empty] | 1 | 1 | 289 | I50f medium sharp ASA | Lung |  |  |  | ROUTER | GEPACS | [8 empty] | manual
| ven Abd |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 10 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven Abd a010 | [22 empty] | 3 | 3 | 131 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | ven Abd cor | [22 empty] | 3 | 3 | 67 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven Abd a010 :-D | [22 empty] | 0.6 | 0.6 | 653 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
5328_Che_3_eso_can_AbdPel (Adult)
5328_Che_3_eso_can_AbdPel (Adult) | Topogram | Topogram |  | 100 |  | 393 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 393 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| ven Che |  | 120 | 120 | 150 | 110 |  |  | On | 7 | On | CARE Dose4D | 10.12 | 30s | 25% | 50 |  | 0.5 |  | 75 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven Che | [22 empty] | 3 | 2 | 145 | I41f medium | Chest |  |  |  | GEPACS | [9 empty] | manual
|  | ven Che sag | [22 empty] | 3 | 2 | 191 | I41f medium | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | ven Che mip | [22 empty] | 10 | 3 | 121 | I41f medium | Lung |  |  |  | GEPACS | [9 empty] | manual
|  | ven Che c220 3dt | [22 empty] | 1 | 1 | 289 | I50f medium sharp ASA | Lung |  |  |  | ROUTER | GEPACS | [8 empty] | manual
| ven AbdPel |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 10 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven AbdPel a030 | [22 empty] | 3 | 3 | 131 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel cor | [22 empty] | 3 | 3 | 67 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel sag | [22 empty] | 3 | 3 | 121 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel a030 :-D | [22 empty] | 0.6 | 0.6 | 653 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | [8 empty] | manual
5330_Che_ild_wo_cst (Adult)
5330_Che_ild_wo_cst (Adult) | Supine | Supine |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Supine | Supine |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Che |  | 120 | 120 | 150 | 110 |  |  | On | 7 | On | CARE Dose4D | 10.12 | 30s | 25% | 50 |  | 0.5 |  | 4 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | Che | [22 empty] | 3 | 2 | 145 | I41f medium | Chest |  |  |  | GEPACS | [9 empty] | manual
|  | Che sag | [22 empty] | 3 | 2 | 191 | I41f medium | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Che mip | [22 empty] | 10 | 3 | 97 | I41f medium | Lung |  |  |  | GEPACS | [9 empty] | manual
|  | Che c130 3dt | [22 empty] | 1 | 1 | 289 | I50f medium sharp ASA | Lung |  |  |  | GEPACS | [9 empty] | manual
| exp Che |  | 120 | 120 | 150 | 110 |  |  | On | 7 | On | CARE Dose4D | 10.12 | 30s | 25% | 50 |  | 0.5 |  | 2 | 1 |  | 1 | 128x0.6mm | [5 empty] | 962494eb-db00-4e54-b01b-9e42ecd67603 | [9 empty] | Off
|  | exp Che c330 | [22 empty] | 1 | 1 | 289 | I41f medium | Lung |  |  |  | ROUTER | GEPACS | [8 empty] | manual
| Prone | Prone |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Prone | Prone |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| pro Che |  | 120 | 120 | 150 | 110 |  |  | On | 7 | On | CARE Dose4D | 1.03 |  | 25% | 50 |  | 0.5 |  | 2 | 0 | 20 | 1 | 2x1.0mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | pro Che c330 | [22 empty] | 1 |  | 14 | B41s medium + | Lung |  |  |  | GEPACS | [9 empty] | manual
5334_Che_lun_can_scr_wo_cst_T07_08A (Adult)
5334_Che_lun_can_scr_wo_cst_T07_08A (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Che |  | 120 | 120 | 30 | 110 |  |  | Off | 2 | On | CARE Dose4D | 2.04 | 30s | 25% | 50 |  | 0.285 |  | 5 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | Che | [22 empty] | 3 | 2 | 90 | I41f medium | Chest |  |  |  | GEPACS | [9 empty] | manual
|  | Che sag | [22 empty] | 3 | 2 | 191 | I41f medium | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Che cor | [22 empty] | 3 | 2 | 181 | I41f medium | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Che mip | [22 empty] | 10 | 3 | 121 | I41f medium | Lung |  |  |  | GEPACS | [9 empty] | manual
|  | Che c334 3dt | [22 empty] | 1 | 1 | 180 | I50f medium sharp ASA | Lung |  |  |  | ROUTER | GEPACS | [8 empty] | manual
5335_Che_pec_wo_cst (Adult)
5335_Che_pec_wo_cst (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Che |  | 120 | 120 | 150 | 110 |  |  | On | 7 | On | CARE Dose4D | 10.12 | 30s | 25% | 50 |  | 0.5 |  | 4 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | Che | [22 empty] | 3 | 2 | 145 | I41f medium | Chest |  |  |  | GEPACS | [9 empty] | manual
|  | Che sag | [22 empty] | 3 | 2 | 191 | I41f medium | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Che mip | [22 empty] | 10 | 3 | 121 | I41f medium | Lung |  |  |  | GEPACS | [9 empty] | manual
|  | Che c335 3dt | [22 empty] | 1 | 1 | 289 | I50f medium sharp ASA | Lung |  |  |  | ROUTER | GEPACS | [8 empty] | manual
| exp Che |  | 120 | 120 | 150 | 110 |  |  | On | 7 | On | CARE Dose4D | 10.12 | 30s | 25% | 50 |  | 0.5 |  | 2 | 1 |  | 3 | 128x0.6mm | [5 empty] | 962494eb-db00-4e54-b01b-9e42ecd67603 | [9 empty] | Off
|  | exp Che | [22 empty] | 3 | 2 | 145 | I41f medium | Chest |  |  |  | GEPACS | [9 empty] | manual
|  | exp Che c335 3dt | [22 empty] | 1 | 1 | 289 | I50f medium sharp ASA | Lung |  |  |  | ROUTER | GEPACS | [8 empty] | manual
5337_Che_pec_ld_wo_cst (Adult)
5337_Che_pec_ld_wo_cst (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Che |  | 120 | 120 | 150 | 110 |  |  | On | 7 | On | CARE Dose4D | 10.12 | 30s | 25% | 50 |  | 0.5 |  | 4 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | Che | [22 empty] | 3 | 2 | 145 | I41f medium | Chest |  |  |  | GEPACS | [9 empty] | manual
|  | Che sag | [22 empty] | 3 | 2 | 191 | I41f medium | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Che mip | [22 empty] | 10 | 3 | 121 | I41f medium | Lung |  |  |  | GEPACS | [9 empty] | manual
|  | Che c335 3dt | [22 empty] | 1 | 1 | 289 | I50f medium sharp ASA | Lung |  |  |  | ROUTER | GEPACS | [8 empty] | manual
| exp Che |  | 120 | 120 | 150 | 110 |  |  | On | 7 | On | CARE Dose4D | 10.12 | 30s | 25% | 50 |  | 0.5 |  | 2 | 1 |  | 3 | 128x0.6mm | [5 empty] | 962494eb-db00-4e54-b01b-9e42ecd67603 | [9 empty] | Off
|  | exp Che | [22 empty] | 3 | 2 | 145 | I41f medium | Chest |  |  |  | GEPACS | [9 empty] | manual
|  | exp Che c335 3dt | [22 empty] | 1 | 1 | 289 | I50f medium sharp ASA | Lung |  |  |  | ROUTER | GEPACS | [8 empty] | manual
5338_Che_pro_gat_wo_cst (Adult)
5338_Che_pro_gat_wo_cst (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Che |  | 100 | 100 | 290 |  | 290 |  | Semi | 11 | On | CARE Dose4D | 6.32 |  | 25% | 50 |  | 0.285 |  | 10 | 0 | 34 | 0.75 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [7 empty] | 70% | 70% | Off | 70% | 70%
|  | Che c338 :-D | [22 empty] | 0.75 |  | 413 | I30f medium smooth | Cardiac |  | Best Diastolic |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual | 70%
|  | Che c338 | [22 empty] | 2 |  | 103 | I30f medium smooth | Cardiac |  |  |  | GEPACS | [9 empty] | manual | 70%
5339_Che_ret_gat_wo_cst (Adult)
5339_Che_ret_gat_wo_cst (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | ff800867-8571-4a15-997f-1d4ffbea68b7 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | ff800867-8571-4a15-997f-1d4ffbea68b7 |  |  | GEPACS |  | TRACKER
| Cor_Spiral |  | 100 | 100 | 290 |  | 220 |  | Semi | 11 | On | CARE Dose4D | 25.02 |  | 25% | 150 |  | 0.285 |  | 15 | 0.17 |  | 0.75 | 128x0.6mm | [5 empty] | ff800867-8571-4a15-997f-1d4ffbea68b7 | [9 empty] | Advance | 70% | 70%
|  | Che c339 | [22 empty] | 0.75 | 0.5 | 401 | I41f medium | Cardiac |  | Best Diastolic |  | MCH3DLabAW | GEPACS | SYNGO_VIA | [7 empty] | BestDiast | 70%
|  | Che c339 | [22 empty] | 3 | 2 | 101 | I41f medium | Cardiac |  | Best Diastolic |  | MCH3DLabAW | GEPACS | SYNGO_VIA | [7 empty] | BestDiast | 70%
5340_Che_spn (Adult)
5340_Che_spn (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Che c340 |  | 120 | 120 | 100 | 110 |  |  | On | 7 | On | CARE Dose4D | 6.75 | 30s | 25% | 50 |  | 0.5 |  | 4 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | Che | [22 empty] | 3 | 2 | 145 | I30f medium smooth | Chest |  |  |  | GEPACS | TERARECON | [8 empty] | manual
|  | Che sag | [22 empty] | 3 | 2 | 191 | I30f medium smooth | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Che mip | [22 empty] | 10 | 3 | 121 | I30f medium smooth | Lung |  |  |  | GEPACS | [9 empty] | manual
|  | Che c130 3dt | [22 empty] | 1 | 1 | 289 | I30f medium smooth | Lung |  |  |  | ROUTER | GEPACS | [8 empty] | manual
| Che nfov |  | 120 | 120 | 150 | 110 |  |  | On | 7 | On | CARE Dose4D | 10.12 | 30s | 25% | 50 |  | 0.5 |  | 4 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | Che nfov c340 | [22 empty] | 3 | 2 | 145 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | Che nfov c340 :-D | [22 empty] | 0.6 | 0.6 | 481 | I30f medium smooth | Abdomen |  |  |  | ROUTER | GEPACS | TERARECON | [7 empty] | manual
| 1m Che |  |  | 80 | 300 |  | 205 | 87 |  | 3 | On | CARE Dose4D | 11.8 | 11s | 25% | 50 |  | 0.5 |  | 60 | 1 |  | 3 | 32x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | 1m Che :-D | [22 empty] | 3 | 2 | 151 | I30f medium smooth | Abdomen |  |  |  | GEPACS | SYNGO_VIA | [8 empty] | manual
|  | 1m Che F0_4 | [22 empty] | 3 | 2 | 151 | I30f medium smooth | Abdomen |  |  |  | GEPACS | TERARECON | [8 empty] | manual
| 2m Che |  |  | 80 | 300 |  | 205 | 87 |  | 3 | On | CARE Dose4D | 11.8 | 11s | 25% | 50 |  | 0.5 |  | 60 | 1 |  | 3 | 32x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | 2m Che :-D | [22 empty] | 3 | 2 | 151 | I30f medium smooth | Abdomen |  |  |  | GEPACS | SYNGO_VIA | [8 empty] | manual
|  | 2m Che F0_4 | [22 empty] | 3 | 2 | 151 | I30f medium smooth | Abdomen |  |  |  | GEPACS | TERARECON | [8 empty] | manual
| 3m Che |  | 120 | 120 | 150 | 110 |  |  | On | 7 | On | CARE Dose4D | 10.12 | 30s | 25% | 50 |  | 0.5 |  | 60 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | 3m Che c340 | [22 empty] | 3 | 2 | 145 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | 3m Che c340 | [22 empty] | 0.6 | 0.6 | 481 | I30f medium smooth | Abdomen |  |  |  | TERARECON | [9 empty] | manual
| 4m Che |  | 120 | 120 | 150 | 110 |  |  | On | 7 | On | CARE Dose4D | 10.12 | 30s | 25% | 50 |  | 0.5 |  | 60 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | 4m Che c340 | [22 empty] | 3 | 2 | 145 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | 4m Che c340 | [22 empty] | 0.6 | 0.6 | 481 | I30f medium smooth | Abdomen |  |  |  | TERARECON | [9 empty] | manual
5360_Che_vir_bro_wo_cst (Adult)
5360_Che_vir_bro_wo_cst (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Che |  | 120 | 120 | 150 | 110 |  |  | On | 7 | On | CARE Dose4D | 10.12 | 30s | 25% | 50 |  | 0.5 |  | 4 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | Che | [22 empty] | 3 | 2 | 145 | I41f medium | Chest |  |  |  | GEPACS | [9 empty] | manual
|  | Che sag | [22 empty] | 2 | 2 | 191 | I41f medium | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Che mip | [22 empty] | 10 | 3 | 121 | I41f medium | Lung |  |  |  | GEPACS | [9 empty] | manual
|  | Che c360 3dt | [22 empty] | 1 | 0.8 | 361 | I50f medium sharp ASA | Lung |  |  |  | ROUTER | GEPACS | [8 empty] | manual
5370_Hrt_CACS (Adult)
5370_Hrt_CACS (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | ff800867-8571-4a15-997f-1d4ffbea68b7 |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | ff800867-8571-4a15-997f-1d4ffbea68b7 |  |  | GEPACS | ROEMIR001a | TRACKER
| CaScSeq |  | 120 | 120 | 100 |  | 80 |  | Semi | 3 | On | CARE Dose4D | 3.75 |  | 25% | 50 |  | 0.285 |  | 2 | 0 | 34 | 3 | 32x1.2mm | [5 empty] | ff800867-8571-4a15-997f-1d4ffbea68b7 | [7 empty] | 70% | 70% | Off | 70% | 70%
|  | CaScSeq 3.0 I30f 1 70% | [22 empty] | 3 |  | 103 | I30f medium smooth | Mediastinum |  |  |  | GEPACS | MCH3DLabAW | [8 empty] | manual | 70%
|  | CaScSeq 3.0 nfov I30f 1 70% | [22 empty] | 3 |  | 103 | I30f medium smooth | Mediastinum |  |  |  | GEPACS | [9 empty] | manual | 70%
DSXXL_Thorax (Adult)
DSXXL_Thorax (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body |  |  |  | GEPACS | ROEMIR001a | TRACKER
Flash_Thorax (Adult)
Flash_Thorax (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body |  |  |  | GEPACS | ROEMIR001a | TRACKER
| Fl_Thorax |  | 120 | 120 | 130 |  | 130 |  | On | 7 | On | CARE Dose4D | 7.2 | 2s | 25% |  |  | 0.285 |  | 2 | 3 |  | 5 | 128x0.6mm
|  | Fl_Thorax 5.0 B31f | [22 empty] | 5 | 5 | 61 | B31f medium smooth + | Mediastinum
|  | Fl_Thorax 5.0 B80f | [22 empty] | 5 | 5 | 61 | B80f ultra sharp | Lung
|  | Fl_Thorax 1.0 B31f | [22 empty] | 1 | 0.7 | 429 | B31f medium smooth + | Mediastinum
|  | Fl_Thorax 1.0 B70f | [22 empty] | 1 | 0.7 | 429 | B70f very sharp | Lung
Flash_Thorax_IR (Adult)
Flash_Thorax_IR (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Fl_Thorax |  | 120 | 120 | 68 |  | 68 |  | On | 7 | On | CARE Dose4D | 3.76 | 5s | 25% |  |  | 0.285 |  | 2 | 3 |  | 5 | 128x0.6mm
|  | Fl_Thorax 5.0 I31f 3 | [22 empty] | 5 | 5 | 61 | I31f medium smooth | Mediastinum
|  | Fl_Thorax 5.0 I70f 3 | [22 empty] | 5 | 5 | 61 | I70f very sharp ASA | Lung
|  | Fl_Thorax 1.0 I31f 3 | [22 empty] | 1 | 0.7 | 429 | I31f medium smooth | Mediastinum
|  | Fl_Thorax 1.0 I70f 3 | [22 empty] | 1 | 0.7 | 429 | I70f very sharp ASA | Lung
LungCARE (Adult)
LungCARE (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body |  |  |  | GEPACS | ROEMIR001a | TRACKER
| LungCARE |  | 120 | 120 | 20 | 20 |  |  | On | 1 | On | CARE Dose4D | 1.35 | 6s | 25% |  |  | 0.5 |  | 2 | 1.2 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | LungCARE 5.0 B31f | [22 empty] | 5 | 5 | 61 | B31f medium smooth + | Mediastinum | [13 empty] | manual
|  | LungCARE 5.0 B80f | [22 empty] | 5 | 5 | 61 | B80f ultra sharp | Lung | [13 empty] | manual
|  | LungCARE 1.0 B31f | [22 empty] | 1 | 0.7 | 429 | B31f medium smooth + | Mediastinum | [13 empty] | manual
|  | LungCARE 1.0 B60f | [22 empty] | 1 | 0.7 | 429 | B60f sharp | Lung | [13 empty] | manual
LungLowDose (Adult)
LungLowDose (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body |  |  |  | GEPACS | ROEMIR001a | TRACKER
| LungLowDose |  | 120 | 120 | 20 | 20 |  |  | On | 1 | On | CARE Dose4D | 1.35 | 6s | 25% |  |  | 0.5 |  | 2 | 1.2 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | LungLowDose 5.0 B31f | [22 empty] | 5 | 5 | 61 | B31f medium smooth + | Mediastinum | [13 empty] | manual
|  | LungLowDose 5.0 B80f | [22 empty] | 5 | 5 | 61 | B80f ultra sharp | Lung | [13 empty] | manual
|  | LungLowDose 1.0 B31f | [22 empty] | 1 | 0.7 | 429 | B31f medium smooth + | Mediastinum | [13 empty] | manual
|  | LungLowDose 1.0 B70f | [22 empty] | 1 | 0.7 | 429 | B70f very sharp | Lung | [13 empty] | manual
TEST_5060_ang_Che_pe (Adult)
TEST_5060_ang_Che_pe (Adult) | Topogram | Phantom 30cm |  | 100 |  | 73 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 3 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Topogram | Phantom 30cm |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body |  |  |  | PHYSICSMIDIA
| Flash 1.0 |  | 100 | 100 | 146 |  | 100 |  | On | 8 | On | CARE Dose4D | 4.84 | 2s | 25% | 50 |  | 0.285 |  | 6 | 2.2 |  | 3 | 128x0.6mm
|  | art Che FLASH1.0 | [22 empty] | 3 | 2 | 46 | I41f medium | Mediastinum |  |  |  | PHYSICSMIDIA
|  | art Che FLASH1.0 :-D | [22 empty] | 1 | 1 | 251 | I26f medium smooth ASA | Mediastinum |  |  |  | PHYSICSMIDIA
| Flash 2.0 |  | 120 | 100 | 250 |  | 100 |  | On | 8 | On | CARE Dose4D | 7.67 | 2s | 25% | 50 |  | 0.285 |  | 2 | 1.6 |  | 3 | 128x0.6mm
|  | art Che FLASH2.0 | [22 empty] | 3 | 2 | 46 | I41f medium | Mediastinum |  |  |  | PHYSICSMIDIA
|  | art Che FLASH2.0 :-D | [22 empty] | 1 | 1 | 251 | I26f medium smooth ASA | Mediastinum |  |  |  | PHYSICSMIDIA
| Helical |  | 120 | 120 | 150 | 110 |  |  | On | 8 | On | CARE Dose4D | 10.12 | 30s | 25% | 50 |  | 0.5 |  | 6 | 1 |  | 3 | 128x0.6mm | [15 empty] | Off
|  | art Che HELICAL | [22 empty] | 3 | 2 | 46 | I41f medium | Mediastinum |  |  |  | PHYSICSMIDIA | [9 empty] | manual
|  | art Che HELICAL :-D | [22 empty] | 1 | 1 | 289 | I26f medium smooth ASA | Abdomen |  |  |  | PHYSICSMIDIA | [9 empty] | manual
ThorAbd (Adult)
ThorAbd (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 275 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body |  |  |  | GEPACS | ROEMIR001a | TRACKER
| ThorAbd |  | 120 | 120 | 210 | 210 |  |  | On | 7 | On | CARE Dose4D | 14.19 | 30s | 25% |  |  | 0.5 |  | 2 | 0.6 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | Tho ThorAbd 5.0 B31f | [22 empty] | 5 | 5 | 41 | B31f medium smooth + | Mediastinum | [13 empty] | manual
|  | Lun ThorAbd 5.0 B80f | [22 empty] | 5 | 5 | 41 | B80f ultra sharp | Lung | [13 empty] | manual
|  | Abd ThorAbd 5.0 B30f | [22 empty] | 5 | 5 | 81 | B30f medium smooth | Abdomen | [13 empty] | manual
|  | MPR ThorAbd 1.0 B30f | [22 empty] | 1 | 0.7 | 858 | B30f medium smooth | Mediastinum | [13 empty] | manual
ThorAbd_XCARE (Adult)
ThorAbd_XCARE (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 275 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body |  |  |  | GEPACS | ROEMIR001a | TRACKER
| ThorAbd |  | 120 | 120 | 210 | 210 |  |  | On | 7 | On | CARE Dose4D | 14.19 | 30s | 25% |  |  | 0.5 |  | 2 | 0.6 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | Tho ThorAbd 5.0 B31f | [22 empty] | 5 | 5 | 41 | B31f medium smooth + | Mediastinum |  |  | X-CARE | [10 empty] | manual
|  | Lun ThorAbd 5.0 B80f | [22 empty] | 5 | 5 | 41 | B80f ultra sharp | Lung |  |  | X-CARE | [10 empty] | manual
|  | Abd ThorAbd 5.0 B30f | [22 empty] | 5 | 5 | 81 | B30f medium smooth | Abdomen |  |  | X-CARE | [10 empty] | manual
|  | MPR ThorAbd 1.0 B30f | [22 empty] | 1 | 0.7 | 858 | B30f medium smooth | Mediastinum |  |  | X-CARE | [10 empty] | manual
ThoraxECGSeqHR (Adult)
ThoraxECGSeqHR (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body |  |  |  | GEPACS | ROEMIR001a | TRACKER
| ThorECGHR |  | 120 | 120 | 110 | 110 |  |  | On | 3 | On | CARE Dose4D | 0.9 |  | 25% |  |  | 0.285 |  | 2 | 0 | 10 | 1 | 2x1.0mm | [13 empty] | 70% | 70% | Off
|  | ThorECGHR 1.0 B80f 70% | [22 empty] | 1 |  | 60 | B80f ultra sharp | Lung | [13 empty] | manual | 70%
ThoraxHR (Adult)
ThoraxHR (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body |  |  |  | GEPACS | ROEMIR001a | TRACKER
| ThorHR |  | 120 | 120 | 110 | 110 |  |  | On | 3 | On | CARE Dose4D | 7.43 | 6s | 25% |  |  | 0.5 |  | 2 | 1.2 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | ThorHR 5.0 B80f | [22 empty] | 5 | 5 | 61 | B80f ultra sharp | Lung | [13 empty] | manual
|  | ThorHR 5.0 B31f | [22 empty] | 5 | 5 | 61 | B31f medium smooth + | Mediastinum | [13 empty] | manual
|  | ThorHR 0.6 B70f | [22 empty] | 0.6 | 0.4 | 751 | B70f very sharp | Lung | [13 empty] | manual
|  | ThorHR 0.6 B31f | [22 empty] | 0.6 | 0.4 | 751 | B31f medium smooth + | Mediastinum | [13 empty] | manual
ThoraxHR_IR (Adult)
ThoraxHR_IR (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| ThorHR |  | 120 | 120 | 66 | 66 |  |  | On | 3 | On | CARE Dose4D | 4.48 | 10s | 25% |  |  | 0.5 |  | 2 | 1.2 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | ThorHR 5.0 I70f 3 | [22 empty] | 5 | 5 | 61 | I70f very sharp ASA | Lung | [13 empty] | manual
|  | ThorHR 5.0 I31f 3 | [22 empty] | 5 | 5 | 61 | I31f medium smooth | Mediastinum | [13 empty] | manual
|  | ThorHR 0.6 I70f 3 | [22 empty] | 0.6 | 0.4 | 751 | I70f very sharp ASA | Lung | [13 empty] | manual
|  | ThorHR 0.6 I31f 3 | [22 empty] | 0.6 | 0.4 | 751 | I31f medium smooth | Mediastinum | [13 empty] | manual
ThoraxIntervention (Adult)
ThoraxIntervention (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body |  |  |  | GEPACS | ROEMIR001a | TRACKER
| i-Spiral |  | 120 | 120 | 110 | 110 |  |  | On | 3 | On | CARE Dose4D | 7.6 | 6s | 25% |  |  | 0.5 |  | 2 | 1.2 |  | 5 | 32x1.2mm | [15 empty] | Off
|  | i-Spiral 5.0 B31f | [22 empty] | 5 | 3.5 | 86 | B31f medium smooth + | Mediastinum | [13 empty] | manual
ThoraxRoutine (Adult)
ThoraxRoutine (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body |  |  |  | GEPACS | ROEMIR001a | TRACKER
| Thorax |  | 120 | 120 | 110 | 110 |  |  | On | 7 | On | CARE Dose4D | 7.43 | 6s | 25% |  |  | 0.5 |  | 2 | 1.2 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | Thorax 5.0 B31f | [22 empty] | 5 | 5 | 61 | B31f medium smooth + | Mediastinum | [13 empty] | manual
|  | Thorax 5.0 B80f | [22 empty] | 5 | 5 | 61 | B80f ultra sharp | Lung | [13 empty] | manual
|  | Thorax 1.0 B31f | [22 empty] | 1 | 0.7 | 429 | B31f medium smooth + | Mediastinum | [13 empty] | manual
|  | Thorax 1.0 B70f | [22 empty] | 1 | 0.7 | 429 | B70f very sharp | Lung | [13 empty] | manual
ThoraxSeqHR (Adult)
ThoraxSeqHR (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body |  |  |  | GEPACS | ROEMIR001a | TRACKER
| ThorHRSeq |  | 120 | 120 | 110 | 110 |  |  | On | 3 | On | CARE Dose4D | 1.51 |  | 25% |  |  | 0.5 |  | 2 | 0 | 10 | 1 | 2x1.0mm | [15 empty] | Off
|  | ThorHRSeq 1.0 B80s | [22 empty] | 1 |  | 60 | B80s ultra sharp | Lung | [13 empty] | manual
Thorax_IR (Adult)
Thorax_IR (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Thorax |  | 120 | 120 | 66 | 66 |  |  | On | 7 | On | CARE Dose4D | 4.48 | 10s | 25% |  |  | 0.5 |  | 2 | 1.2 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | Thorax 5.0 I31f 3 | [22 empty] | 5 | 5 | 61 | I31f medium smooth | Mediastinum | [13 empty] | manual
|  | Thorax 5.0 I70f 3 | [22 empty] | 5 | 5 | 61 | I70f very sharp ASA | Lung | [13 empty] | manual
|  | Thorax 1.0 I31f 3 | [22 empty] | 1 | 0.7 | 429 | I31f medium smooth | Mediastinum | [13 empty] | manual
|  | Thorax 1.0 I70f 3 | [22 empty] | 1 | 0.7 | 429 | I70f very sharp ASA | Lung | [13 empty] | manual
Thorax_XCARE (Adult)
Thorax_XCARE (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body |  |  |  | GEPACS | ROEMIR001a | TRACKER
| Thorax |  | 120 | 120 | 110 | 110 |  |  | On | 7 | On | CARE Dose4D | 7.45 | 6s | 25% |  |  | 0.285 |  | 2 | 0.6 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | Thorax 5.0 B31f | [22 empty] | 5 | 5 | 61 | B31f medium smooth + | Mediastinum |  |  | X-CARE | [10 empty] | manual
|  | Thorax 5.0 B80f | [22 empty] | 5 | 5 | 61 | B80f ultra sharp | Lung |  |  | X-CARE | [10 empty] | manual
|  | Thorax 1.0 B31f | [22 empty] | 1 | 0.7 | 429 | B31f medium smooth + | Mediastinum |  |  | X-CARE | [10 empty] | manual
|  | Thorax 1.0 B70f | [22 empty] | 1 | 0.7 | 429 | B70f very sharp | Lung |  |  | X-CARE | [10 empty] | manual
Thorax_XCARE_IR (Adult)
Thorax_XCARE_IR (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Thorax |  | 120 | 120 | 66 | 66 |  |  | On | 7 | On | CARE Dose4D | 4.46 | 10s | 25% |  |  | 0.285 |  | 2 | 0.6 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | Thorax 5.0 I31f 3 | [22 empty] | 5 | 5 | 61 | I31f medium smooth | Mediastinum |  |  | X-CARE | [10 empty] | manual
|  | Thorax 5.0 I70f 3 | [22 empty] | 5 | 5 | 61 | I70f very sharp ASA | Lung |  |  | X-CARE | [10 empty] | manual
|  | Thorax 1.0 I31f 3 | [22 empty] | 1 | 0.7 | 429 | I31f medium smooth | Mediastinum |  |  | X-CARE | [10 empty] | manual
|  | Thorax 1.0 I70f 3 | [22 empty] | 1 | 0.7 | 429 | I70f very sharp ASA | Lung |  |  | X-CARE | [10 empty] | manual
Abdomen
6010_Abd (Adult)
6010_Abd (Adult) | Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| ven Abd |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 90 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven Abd a010 | [22 empty] | 3 | 3 | 131 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | ven Abd cor | [22 empty] | 3 | 3 | 67 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven Abd a010 :-D | [22 empty] | 0.6 | 0.6 | 653 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
6030_AbdPel_A02_00A (Adult)
6030_AbdPel_A02_00A (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| ven AbdPel |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 90 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven AbdPel a030 | [22 empty] | 3 | 3 | 131 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel cor | [22 empty] | 3 | 3 | 67 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel sag | [22 empty] | 3 | 3 | 121 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel a030 :-D | [22 empty] | 0.6 | 0.6 | 653 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
6031_AbdPel_DE (Adult)
6031_AbdPel_DE (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T81f sharp | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T81f sharp | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| DE ven AbdPel |  |  | 100 | 240 |  | 202 | 156 |  | 3 | On | CARE Dose4D | 18.61 | 30s | 25% | 50 |  | 0.5 |  | 90 | 0.6 |  | 3 | 32x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven AbdPel M0.4 | [22 empty] | 3 | 3 | 88 | Q40f medium | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel M0.4 cor | [22 empty] | 2 | 2 | 181 | Q40f medium | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel M0.4 sag | [22 empty] | 2 | 2 | 181 | Q40f medium | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel a031 :-D #PP | [22 empty] | 0.6 | 0.6 | 439 | Q40f medium | 400/40 Abd |  |  |  | ROUTER | GEPACS | SYNGO_VIA | [7 empty] | manual
6050_Abd_adr (Adult)
6050_Abd_adr (Adult) | Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Abd |  | 120 | 120 | 180 | 210 |  |  | Semi | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 4 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | Abd a050 | [22 empty] | 3 | 3 | 21 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | Abd a050 :-D | [22 empty] | 0.6 | 0.6 | 105 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
| ven Abd |  | 120 | 120 | 180 | 210 |  |  | Semi | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 70 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven Abd a050 | [22 empty] | 3 | 3 | 21 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven Abd cor | [22 empty] | 3 | 3 | 121 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven Abd a050 :-D | [22 empty] | 0.6 | 0.6 | 105 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
| 15m Abd |  | 120 | 120 | 180 | 210 |  |  | Semi | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 2 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | 15m Abd a050 | [22 empty] | 3 | 3 | 21 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | 15m Abd a050 :-D | [22 empty] | 0.6 | 0.6 | 105 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
6060_Abd_adr_Pel (Adult)
6060_Abd_adr_Pel (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Abd |  | 120 | 120 | 180 | 210 |  |  | Semi | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 4 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | Abd a060 | [22 empty] | 3 | 3 | 21 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | Abd a060 :-D | [22 empty] | 0.6 | 0.6 | 105 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
| ven AbdPel |  | 120 | 120 | 180 | 210 |  |  | Semi | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 70 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven AbdPel a060 | [22 empty] | 3 | 3 | 21 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel cor | [22 empty] | 3 | 3 | 121 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel sag | [22 empty] | 3 | 3 | 83 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel a060 :-D | [22 empty] | 0.6 | 0.6 | 105 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
| 15m Abd |  | 120 | 120 | 180 | 210 |  |  | Semi | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 2 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | 15m Abd a060 | [22 empty] | 3 | 3 | 21 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | 15m Abd a060 :-D | [22 empty] | 0.6 | 0.6 | 105 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
6070_Abd_liv_bi (Adult)
6070_Abd_liv_bi (Adult) | Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| PreMonitoring |  | 120 | 120 | 240 | 20 |  |  | Off | 8 | Off | CARE Dose4D | 1.23 |  | 25% |  |  | 0.5 |  | 2 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | PreMonitoring 10.0 B30s | [22 empty] | 10 |  | 1 | B30s medium smooth | 400/40 Abd | [13 empty] | manual
| Monitoring |  | 120 | 120 | 240 | 20 |  |  | Off | 8 | Off | CARE Dose4D | 36.98 |  | 25% |  |  | 0.5 |  | 10 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | Monitoring 10.0 B30s | [22 empty] | 10 |  | 30 | B30s medium smooth | 400/40 Abd | [13 empty] | manual
| art Abd |  | 120 | 120 | 180 | 210 |  |  | On | 8 | On | CARE Dose4D | 12.17 | 30s | 25% | 50 |  | 0.5 |  | 16 | 0.6 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | art Abd a070 | [22 empty] | 3 | 3 | 61 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd a070 :-D | [22 empty] | 0.6 | 0.6 | 301 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
| ven Abd |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.17 | 30s | 25% | 50 |  | 0.5 |  | 60 | 0.6 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven Abd a070 | [22 empty] | 3 | 3 | 45 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven Abd cor | [22 empty] | 3 | 3 | 84 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven Abd a070 :-D | [22 empty] | 0.6 | 0.6 | 223 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
6080_Abd_liv_bi_Pel (Adult)
6080_Abd_liv_bi_Pel (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| PreMonitoring |  | 120 | 120 | 240 | 20 |  |  | Off | 8 | Off | CARE Dose4D | 1.23 |  | 25% |  |  | 0.5 |  | 2 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | PreMonitoring 10.0 B30s | [22 empty] | 10 |  | 1 | B30s medium smooth | 400/40 Abd | [13 empty] | manual
| Monitoring |  | 120 | 120 | 240 | 20 |  |  | Off | 8 | Off | CARE Dose4D | 36.98 |  | 25% |  |  | 0.5 |  | 10 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | Monitoring 10.0 B30s | [22 empty] | 10 |  | 30 | B30s medium smooth | 400/40 Abd | [13 empty] | manual
| art Abd |  | 120 | 120 | 180 | 210 |  |  | On | 8 | On | CARE Dose4D | 12.17 | 30s | 25% | 50 |  | 0.5 |  | 16 | 0.6 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | art Abd a080 | [22 empty] | 3 | 3 | 61 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd a080 :-D | [22 empty] | 0.6 | 0.6 | 301 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
| ven AbdPel |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.17 | 30s | 25% | 50 |  | 0.5 |  | 60 | 0.6 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven AbdPel a080 | [22 empty] | 3 | 3 | 45 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel cor | [22 empty] | 3 | 3 | 84 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel sag | [22 empty] | 3 | 3 | 121 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel a080 :-D | [22 empty] | 0.6 | 0.6 | 223 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
6090_Abd_liv_tri (Adult)
6090_Abd_liv_tri (Adult) | Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| PreMonitoring |  | 120 | 120 | 240 | 20 |  |  | Off | 8 | Off | CARE Dose4D | 1.23 |  | 25% |  |  | 0.5 |  | 2 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | PreMonitoring 10.0 B30s | [22 empty] | 10 |  | 1 | B30s medium smooth | 400/40 Abd | [13 empty] | manual
| Monitoring |  | 120 | 120 | 240 | 20 |  |  | Off | 8 | Off | CARE Dose4D | 36.98 |  | 25% |  |  | 0.5 |  | 10 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | Monitoring 10.0 B30s | [22 empty] | 10 |  | 30 | B30s medium smooth | 400/40 Abd | [13 empty] | manual
| art Abd |  | 120 | 120 | 180 | 210 |  |  | On | 8 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 16 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | art Abd a090 | [22 empty] | 3 | 3 | 117 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd cor | [22 empty] | 3 | 3 | 128 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd sag | [22 empty] | 3 | 3 | 121 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd a090 :-D | [22 empty] | 0.6 | 0.6 | 584 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
| ven Abd |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 60 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven Abd a090 | [22 empty] | 3 | 3 | 117 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven Abd cor | [22 empty] | 3 | 3 | 135 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven Abd a090 :-D | [22 empty] | 0.6 | 0.6 | 584 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
| 5m Abd |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 4 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | 5m Abd a090 | [22 empty] | 3 | 3 | 117 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | 5m Abd a090 :-D | [22 empty] | 0.6 | 0.6 | 584 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
6110_Abd_liv_tri_Pel (Adult)
6110_Abd_liv_tri_Pel (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| PreMonitoring |  | 120 | 120 | 240 | 20 |  |  | Off | 8 | Off | CARE Dose4D | 1.23 |  | 25% |  |  | 0.5 |  | 2 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | PreMonitoring 10.0 B30s | [22 empty] | 10 |  | 1 | B30s medium smooth | 400/40 Abd | [13 empty] | manual
| Monitoring |  | 120 | 120 | 240 | 20 |  |  | Off | 8 | Off | CARE Dose4D | 36.98 |  | 25% |  |  | 0.5 |  | 10 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | Monitoring 10.0 B30s | [22 empty] | 10 |  | 30 | B30s medium smooth | 400/40 Abd | [13 empty] | manual
| art Abd |  | 120 | 120 | 180 | 210 |  |  | On | 8 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 16 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | art Abd a110 | [22 empty] | 3 | 3 | 117 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd cor | [22 empty] | 3 | 3 | 128 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd sag | [22 empty] | 3 | 3 | 121 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd a110 :-D | [22 empty] | 0.6 | 0.6 | 584 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
| ven AbdPel |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 60 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven AbdPel a110 | [22 empty] | 3 | 3 | 117 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel cor | [22 empty] | 3 | 3 | 135 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel sag | [22 empty] | 3 | 3 | 121 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel a110 :-D | [22 empty] | 0.6 | 0.6 | 584 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
| 5m Abd |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 4 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | 5m Abd a110 | [22 empty] | 3 | 3 | 117 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | 5m Abd a110 :-D | [22 empty] | 0.6 | 0.6 | 584 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
6125_Abd_liv_qua_pos_TACE (Adult)
6125_Abd_liv_qua_pos_TACE (Adult) | Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Abd |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 2 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | Abd a110 | [22 empty] | 3 | 3 | 131 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | ven Abd cor | [22 empty] | 3 | 3 | 67 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | Abd a110 :-D | [22 empty] | 0.6 | 0.6 | 653 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
| PreMonitoring |  | 120 | 120 | 200 | 20 |  |  | Off | 7 | Off | CARE Dose4D | 1.23 |  | 25% |  |  | 0.5 |  | 4 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | PreMonitoring | [22 empty] | 10 |  | 1 | B30s medium smooth | Abdomen | [13 empty] | manual
| Monitoring |  | 120 | 120 | 200 | 20 |  |  | Off | 7 | Off | CARE Dose4D | 36.98 |  | 25% |  |  | 0.5 |  | 12 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | Monitoring | [22 empty] | 10 |  | 30 | B30s medium smooth | Abdomen | [13 empty] | manual
| art Abd |  | 120 | 120 | 180 | 110 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 16 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | art Abd a110 | [22 empty] | 3 | 2 | 145 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd cor | [22 empty] | 3 | 3 | 127 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd sag | [22 empty] | 3 | 3 | 127 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd a110 :-D | [22 empty] | 0.6 | 0.6 | 481 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
| ven Abd |  | 120 | 120 | 180 | 110 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 60 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven Abd a110 | [22 empty] | 3 | 3 | 97 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | ven Abd cor | [22 empty] | 3 | 3 | 127 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | ven Abd a110 :-D | [22 empty] | 0.6 | 0.6 | 481 | I30f medium smooth | Abdomen |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
| 5m Abd |  | 120 | 120 | 180 | 110 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 2 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | 5m Abd a110 | [22 empty] | 3 | 3 | 97 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | 5m Abd a110 :-D | [22 empty] | 0.6 | 0.6 | 481 | I30f medium smooth | Abdomen |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
6126_Abd_liv_qua_pos_TACE_Pel (Adult)
6126_Abd_liv_qua_pos_TACE_Pel (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Abd |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 2 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | Abd a110 | [22 empty] | 3 | 3 | 131 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | ven Abd cor | [22 empty] | 3 | 3 | 67 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | Abd a110 :-D | [22 empty] | 0.6 | 0.6 | 653 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
| PreMonitoring |  | 120 | 120 | 200 | 20 |  |  | Off | 7 | Off | CARE Dose4D | 1.23 |  | 25% |  |  | 0.5 |  | 4 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | PreMonitoring | [22 empty] | 10 |  | 1 | B30s medium smooth | Abdomen | [13 empty] | manual
| Monitoring |  | 120 | 120 | 200 | 20 |  |  | Off | 7 | Off | CARE Dose4D | 36.98 |  | 25% |  |  | 0.5 |  | 12 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | Monitoring | [22 empty] | 10 |  | 30 | B30s medium smooth | Abdomen | [13 empty] | manual
| art Abd |  | 120 | 120 | 180 | 110 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 16 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | art Abd a110 | [22 empty] | 3 | 2 | 145 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd cor | [22 empty] | 3 | 3 | 127 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd sag | [22 empty] | 3 | 3 | 127 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd a110 :-D | [22 empty] | 0.6 | 0.6 | 481 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
| ven AbdPel |  | 120 | 120 | 180 | 110 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 60 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven AbdPel a110 | [22 empty] | 3 | 3 | 97 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel cor | [22 empty] | 3 | 3 | 127 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel sag | [22 empty] | 3 | 3 | 127 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel a110 :-D | [22 empty] | 0.6 | 0.6 | 481 | I30f medium smooth | Abdomen |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
| 5m Abd |  | 120 | 120 | 180 | 110 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 2 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | 5m Abd a110 | [22 empty] | 3 | 3 | 97 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | 5m Abd a110 :-D | [22 empty] | 0.6 | 0.6 | 481 | I30f medium smooth | Abdomen |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
6130_Abd_pan_bi (Adult)
6130_Abd_pan_bi (Adult) | Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| PreMonitoring |  | 120 | 120 | 240 | 20 |  |  | Off | 8 | Off | CARE Dose4D | 1.23 |  | 25% |  |  | 0.5 |  | 2 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | PreMonitoring 10.0 B30s | [22 empty] | 10 |  | 1 | B30s medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
| Monitoring |  | 120 | 120 | 240 | 20 |  |  | Off | 8 | Off | CARE Dose4D | 36.98 |  | 25% |  |  | 0.5 |  | 10 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | Monitoring 10.0 B30s | [22 empty] | 10 |  | 30 | B30s medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
| art Abd |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 19 | 0.8 |  | 2 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | art Abd a130 | [22 empty] | 2 | 2 | 135 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd cor | [22 empty] | 2 | 2 | 161 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd a130 :-D | [22 empty] | 0.6 | 0.6 | 450 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
| ven Abd |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 50 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven Abd a130 | [22 empty] | 3 | 3 | 90 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven Abd cor | [22 empty] | 3 | 3 | 121 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven Abd a130 :-D | [22 empty] | 0.6 | 0.6 | 447 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
6135_Abd_pan_bi_mod (Adult)
6135_Abd_pan_bi_mod (Adult) | Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| artven Abd |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 75 | 0.8 |  | 2 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | artven Abd a135 | [22 empty] | 2 | 2 | 135 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | artven Abd cor | [22 empty] | 2 | 2 | 181 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | artven Abd a135 :-D | [22 empty] | 0.6 | 0.6 | 447 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
6140_Abd_pan_bi_Pel (Adult)
6140_Abd_pan_bi_Pel (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| PreMonitoring |  | 120 | 120 | 240 | 20 |  |  | Off | 8 | Off | CARE Dose4D | 1.23 |  | 25% |  |  | 0.5 |  | 2 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | PreMonitoring 10.0 B30s | [22 empty] | 10 |  | 1 | B30s medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
| Monitoring |  | 120 | 120 | 240 | 20 |  |  | Off | 8 | Off | CARE Dose4D | 36.98 |  | 25% |  |  | 0.5 |  | 10 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | Monitoring 10.0 B30s | [22 empty] | 10 |  | 30 | B30s medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
| art Abd |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 19 | 0.8 |  | 2 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | art Abd a140 | [22 empty] | 2 | 2 | 135 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd cor | [22 empty] | 2 | 2 | 161 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd a130 :-D | [22 empty] | 0.6 | 0.6 | 450 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
| ven AbdPel |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 50 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven AbdPel a130 | [22 empty] | 3 | 3 | 90 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel cor | [22 empty] | 3 | 3 | 121 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel sag | [22 empty] | 3 | 3 | 121 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel a140 :-D | [22 empty] | 0.6 | 0.6 | 447 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
6145_Abd_pan_bi_mod_Pel (Adult)
6145_Abd_pan_bi_mod_Pel (Adult) | Topogram | Topogram 0.6 T20f |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| artven AbdPel |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 75 | 0.8 |  | 2 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | artven AbdPel a145 | [22 empty] | 2 | 2 | 135 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | artven AbdPel cor | [22 empty] | 2 | 2 | 181 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | artven AbdPel sag | [22 empty] | 2 | 2 | 181 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | artven AbdPel a145 :-D | [22 empty] | 0.6 | 0.6 | 447 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
6170_Abd_ren_mas (Adult)
6170_Abd_ren_mas (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T21f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T21f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Abd |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 2 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | Abd a170 | [22 empty] | 3 | 3 | 95 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | Abd a170 :-D | [22 empty] | 0.6 | 0.6 | 473 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
| ven Abd |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 90 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven Abd a170 | [22 empty] | 3 | 3 | 165 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven Abd cor | [22 empty] | 3 | 3 | 104 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven Abd a170 :-D | [22 empty] | 0.6 | 0.6 | 823 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
| 4m Abd |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 4 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | 4m Abd a170 | [22 empty] | 3 | 3 | 95 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | 4m Abd a170 :-D | [22 empty] | 0.6 | 0.6 | 474 | I40f medium | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
6180_Abd_ren_mas_Pel (Adult)
6180_Abd_ren_mas_Pel (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T21f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T21f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Abd |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 2 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | Abd a180 | [22 empty] | 3 | 3 | 95 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | Abd a180 :-D | [22 empty] | 0.6 | 0.6 | 473 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
| ven AbdPel |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 90 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven AbdPel a180 | [22 empty] | 3 | 3 | 165 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel cor | [22 empty] | 3 | 3 | 104 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel sag | [22 empty] | 3 | 3 | 121 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel a180 :-D | [22 empty] | 0.6 | 0.6 | 823 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
| 4m Abd |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 4 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | 4m Abd a180 | [22 empty] | 3 | 3 | 95 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | 4m Abd a180 :-D | [22 empty] | 0.6 | 0.6 | 474 | I40f medium | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
6183_Abd_ren_spa (Adult)
6183_Abd_ren_spa (Adult) | Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Abd |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 2 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | Abd a183 | [22 empty] | 3 | 3 | 131 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | Abd a183 :-D | [22 empty] | 0.6 | 0.6 | 653 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
| PreMonitoring |  | 120 | 120 | 240 | 20 |  |  | Off | 8 | Off | CARE Dose4D | 1.23 |  | 25% |  |  | 0.5 |  | 2 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | PreMonitoring 10.0 B30s | [22 empty] | 10 |  | 1 | B30s medium smooth | 400/40 Abd | [13 empty] | manual
| Monitoring |  | 120 | 120 | 240 | 20 |  |  | Off | 8 | Off | CARE Dose4D | 36.98 |  | 25% |  |  | 0.5 |  | 10 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | Monitoring 10.0 B30s | [22 empty] | 10 |  | 30 | B30s medium smooth | 400/40 Abd | [13 empty] | manual
| art Abd |  | 120 | 120 | 180 | 210 |  |  | On | 8 | On | CARE Dose4D | 12.17 | 30s | 25% | 50 |  | 0.5 |  | 16 | 0.6 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | art Abd a183 | [22 empty] | 3 | 3 | 61 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd a183 :-D | [22 empty] | 0.6 | 0.6 | 301 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
| ven Abd |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.17 | 30s | 25% | 50 |  | 0.5 |  | 60 | 0.6 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven Abd a183 | [22 empty] | 3 | 3 | 45 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven Abd cor | [22 empty] | 3 | 3 | 84 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven Abd sag | [22 empty] | 3 | 3 | 121 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven Abd a183 :-D | [22 empty] | 0.6 | 0.6 | 223 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
| 10m Abd |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 90 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | 10m Abd a183 | [22 empty] | 3 | 3 | 131 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | 10m Abd cor | [22 empty] | 3 | 3 | 67 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | 10m Abd cor mip | [22 empty] | 24 | 5 | 41 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | 10m Abd a183 :-D | [22 empty] | 0.6 | 0.6 | 653 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
6185_Abd_ren_spa_Pel (Adult)
6185_Abd_ren_spa_Pel (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Abd |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 2 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | Abd a185 | [22 empty] | 3 | 3 | 131 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | Abd a185 :-D | [22 empty] | 0.6 | 0.6 | 653 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
| PreMonitoring |  | 120 | 120 | 240 | 20 |  |  | Off | 8 | Off | CARE Dose4D | 1.23 |  | 25% |  |  | 0.5 |  | 2 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | PreMonitoring 10.0 B30s | [22 empty] | 10 |  | 1 | B30s medium smooth | 400/40 Abd | [13 empty] | manual
| Monitoring |  | 120 | 120 | 240 | 20 |  |  | Off | 8 | Off | CARE Dose4D | 36.98 |  | 25% |  |  | 0.5 |  | 10 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | Monitoring 10.0 B30s | [22 empty] | 10 |  | 30 | B30s medium smooth | 400/40 Abd | [13 empty] | manual
| art Abd |  | 120 | 120 | 180 | 210 |  |  | On | 8 | On | CARE Dose4D | 12.17 | 30s | 25% | 50 |  | 0.5 |  | 16 | 0.6 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | art Abd a185 | [22 empty] | 3 | 3 | 61 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd a185 :-D | [22 empty] | 0.6 | 0.6 | 301 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
| ven AbdPel |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.17 | 30s | 25% | 50 |  | 0.5 |  | 60 | 0.6 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven AbdPel a185 | [22 empty] | 3 | 3 | 45 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel cor | [22 empty] | 3 | 3 | 84 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel sag | [22 empty] | 3 | 3 | 121 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel a185 :-D | [22 empty] | 0.6 | 0.6 | 223 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
| 10m Abd |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 90 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | 10m Abd a185 | [22 empty] | 3 | 3 | 131 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | 10m Abd cor | [22 empty] | 3 | 3 | 67 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | 10m Abd cor mip | [22 empty] | 24 | 5 | 41 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | 10m Abd a185 :-D | [22 empty] | 0.6 | 0.6 | 653 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
6190_Abd_wo_cst (Adult)
6190_Abd_wo_cst (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Abd |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 3 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | Abd a190 | [22 empty] | 3 | 3 | 131 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | Abd cor | [22 empty] | 3 | 3 | 67 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | Abd a190 :-D | [22 empty] | 0.6 | 0.6 | 653 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
6200_AbdPel_wo_cst (Adult)
6200_AbdPel_wo_cst (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| AbdPel |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 3 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | AbdPel a200 | [22 empty] | 3 | 3 | 131 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | AbdPel cor | [22 empty] | 3 | 3 | 67 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | AbdPel sag | [22 empty] | 3 | 3 | 121 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | AbdPel a200 :-D | [22 empty] | 0.6 | 0.6 | 653 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
6205_AbdPel_cap_loc_wo_cst (Adult)
6205_AbdPel_cap_loc_wo_cst (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| AbdPel |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.14 | 30s | 25% | 50 |  | 0.5 |  | 4 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | AbdPel a205 | [22 empty] | 3 | 3 | 131 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | AbdPel cor | [22 empty] | 3 | 3 | 67 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | AbdPel sag | [22 empty] | 3 | 3 | 121 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | AbdPel a205 :-D | [22 empty] | 0.6 | 0.6 | 653 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
6211_AbdPel_col_nonVEO (Adult)
6211_AbdPel_col_nonVEO (Adult) | Supine | Supine Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | Supine |  | GEPACS |  | TRACKER
| Supine | Supine Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | Supine |  | GEPACS |  | TRACKER
| AbdPel |  | 120 | 120 | 50 | 210 |  |  | On | 6 | On | CARE Dose4D | 3.38 | 30s | 25% | 50 |  | 0.5 |  | 2 | 1.2 |  | 1 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | AbdPel a211 :- | [22 empty] | 1 | 0.8 | 374 | I26f medium smooth ASA | 400/40 Abd |  | Supine |  | GEPACS | MCB_GI_AW | MCHNeuroBodyAW | [7 empty] | manual
|  | AbdPel a211 | [22 empty] | 0.6 | 0.6 | 499 | I26f medium smooth ASA | 400/40 Abd |  | Supine |  | ROUTER | MCHNeuroBodyAW | [8 empty] | manual
| Prone | Prone Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | Prone |  | GEPACS |  | TRACKER
| Prone | Prone Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | Prone |  | GEPACS |  | TRACKER
| ven pro AbdPel |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 70 | 0.8 |  | 1 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven pro AbdPel a211 :- | [22 empty] | 1 | 0.8 | 374 | I26f medium smooth ASA | 400/40 Abd |  | Prone |  | GEPACS | MCB_GI_AW | MCHNeuroBodyAW | [7 empty] | manual
|  | ven pro AbdPel a211 | [22 empty] | 3 | 3 | 100 | I30f medium smooth | 400/40 Abd |  | Prone | [11 empty] | manual
|  | ven pro AbdPel cor | [22 empty] | 3 | 3 | 121 | I30f medium smooth | 400/40 Abd |  | Prone |  | GEPACS | [9 empty] | manual
|  | ven pro AbdPel a211 | [22 empty] | 0.6 | 0.6 | 499 | I30f medium smooth | 400/40 Abd |  | Axial through liver only Prone |  | ROUTER | MCB_GI_AW | MCHNeuroBodyAW | [7 empty] | manual
| Ldec | Ldec Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | Prone |  | GEPACS |  | TRACKER
| Ldec | Ldec Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | Prone |  | GEPACS |  | TRACKER
| ven Ldec AP |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 70 | 0.8 |  | 1 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven Ldec AbdPel a211 :- | [22 empty] | 1 | 0.8 | 374 | I26f medium smooth ASA | 400/40 Abd |  | Prone |  | GEPACS | MCB_GI_AW | MCHNeuroBodyAW | [7 empty] | manual
|  | ven Ldec AbdPel a211 | [22 empty] | 3 | 3 | 100 | I30f medium smooth | 400/40 Abd |  | Prone | [11 empty] | manual
|  | ven Ldec AbdPel cor | [22 empty] | 3 | 3 | 121 | I30f medium smooth | 400/40 Abd |  | Prone |  | GEPACS | [9 empty] | manual
|  | ven Ldec AbdPel a211 | [22 empty] | 0.6 | 0.6 | 499 | I30f medium smooth | 400/40 Abd |  | Axial through liver only Prone |  | ROUTER | MCB_GI_AW | MCHNeuroBodyAW | [7 empty] | manual
| Rdec | Rdec Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | Prone |  | GEPACS |  | TRACKER
| Rdec | Rdec Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | Prone |  | GEPACS |  | TRACKER
| 10m Rdec AP |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 6 | 0.8 |  | 1 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | 10m Rdec AbdPel a211 :- | [22 empty] | 1 | 0.8 | 374 | I26f medium smooth ASA | 400/40 Abd |  | Prone |  | GEPACS | MCB_GI_AW | MCHNeuroBodyAW | [7 empty] | manual
|  | 10m Rdec AbdPel a211 | [22 empty] | 3 | 3 | 100 | I30f medium smooth | 400/40 Abd |  | Prone | [11 empty] | manual
|  | 10m Rdec AbdPel cor | [22 empty] | 3 | 3 | 121 | I30f medium smooth | 400/40 Abd |  | Prone |  | GEPACS | [9 empty] | manual
|  | 10m Rdec AbdPel a211 | [22 empty] | 0.6 | 0.6 | 499 | I30f medium smooth | 400/40 Abd |  | Axial through liver only Prone |  | ROUTER | MCB_GI_AW | MCHNeuroBodyAW | [7 empty] | manual
6221_AbdPel_col_ld_wo_cst_nonVEO (Adult)
6221_AbdPel_col_ld_wo_cst_nonVEO (Adult) | Supine | Supine Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | Supine |  | GEPACS |  | TRACKER
| Supine | Supine Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | Supine |  | GEPACS |  | TRACKER
| AbdPel |  | 120 | 120 | 50 | 210 |  |  | On | 2 | On | CARE Dose4D | 3.38 | 30s | 25% | 50 |  | 0.5 |  | 2 | 1.2 |  | 1 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | AbdPel a221 :- | [22 empty] | 1 | 0.8 | 374 | I26f medium smooth ASA | 400/40 Abd |  | Supine |  | GEPACS | MCB_GI_AW | MCHNeuroBodyAW | [7 empty] | manual
|  | AbdPel a221 | [22 empty] | 0.6 | 0.6 | 499 | I30f medium smooth | 400/40 Abd |  | Supine |  | ROUTER | MCB_GI_AW | MCHNeuroBodyAW | [7 empty] | manual
| Prone | Prone Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | Prone |  | GEPACS |  | TRACKER
| Prone | Prone Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | Prone |  | GEPACS |  | TRACKER
| pro AbdPel |  | 120 | 120 | 50 | 210 |  |  | On | 2 | On | CARE Dose4D | 3.38 | 30s | 25% | 50 |  | 0.5 |  | 2 | 0.8 |  | 1 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | pro AbdPel a221 :- | [22 empty] | 1 | 0.8 | 374 | I26f medium smooth ASA | 400/40 Abd |  | Prone |  | GEPACS | MCB_GI_AW | MCHNeuroBodyAW | [7 empty] | manual
|  | pro AbdPel a221 | [22 empty] | 0.6 | 0.6 | 499 | I30f medium smooth | 400/40 Abd |  | Prone |  | ROUTER | MCB_GI_AW | MCHNeuroBodyAW | [7 empty] | manual
| Ldec | Ldec Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | Prone |  | GEPACS |  | TRACKER
| Ldec | Ldec Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | Prone |  | GEPACS |  | TRACKER
| Ldec AbdPel |  | 120 | 120 | 50 | 210 |  |  | On | 2 | On | CARE Dose4D | 3.38 | 30s | 25% | 50 |  | 0.5 |  | 2 | 0.8 |  | 1 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | Ldec AbdPel a221 :- | [22 empty] | 1 | 0.8 | 374 | I26f medium smooth ASA | 400/40 Abd |  | Prone |  | GEPACS | MCB_GI_AW | MCHNeuroBodyAW | [7 empty] | manual
|  | Ldec AbdPel a221 | [22 empty] | 0.6 | 0.6 | 499 | I30f medium smooth | 400/40 Abd |  | Prone |  | ROUTER | MCB_GI_AW | MCHNeuroBodyAW | [7 empty] | manual
| Rdec | Rdec Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | Prone |  | GEPACS |  | TRACKER
| Rdec | Rdec Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | Prone |  | GEPACS |  | TRACKER
| Rdec AbdPel |  | 120 | 120 | 50 | 210 |  |  | On | 2 | On | CARE Dose4D | 3.38 | 30s | 25% | 50 |  | 0.5 |  | 3 | 0.8 |  | 1 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | Rdec AbdPel a221 :- | [22 empty] | 1 | 0.8 | 374 | I26f medium smooth ASA | 400/40 Abd |  | Prone |  | GEPACS | MCB_GI_AW | MCHNeuroBodyAW | [7 empty] | manual
|  | Rdec AbdPel a221 | [22 empty] | 0.6 | 0.6 | 499 | I30f medium smooth | 400/40 Abd |  | Prone |  | ROUTER | MCB_GI_AW | MCHNeuroBodyAW | [7 empty] | manual
6230_AbdPel_ent_A06_00A (Adult)
6230_AbdPel_ent_A06_00A (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| ven AbdPel |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.14 | 30s | 25% | 50 |  | 0.5 |  | 70 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven AbdPel a230 | [22 empty] | 3 | 2 | 176 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel cor | [22 empty] | 3 | 2 | 139 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel sag | [22 empty] | 2 | 2 | 181 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel cor mip | [22 empty] | 16 | 5 | 73 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel a230 :-D | [22 empty] | 0.6 | 0.6 | 584 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
6240_AbdPel_ho_tcc (Adult)
6240_AbdPel_ho_tcc (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T21f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T21f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| ven AbdPel |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 100 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven AbdPel a240 | [22 empty] | 3 | 3 | 104 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel cor | [22 empty] | 3 | 3 | 51 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel sag | [22 empty] | 3 | 3 | 121 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel a240 :-D | [22 empty] | 0.6 | 0.6 | 516 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
| 10m AbdPel |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 2 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | 10m AbdPel a240 | [22 empty] | 3 | 3 | 80 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | 10m AbdPel cor | [22 empty] | 3 | 3 | 67 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | 10m AbdPel cor 10x10 | [22 empty] | 10 | 10 | 39 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | 10m AbdPel cor mip | [22 empty] | 24 | 5 | 77 | I30f medium smooth | Abdomen MIP |  |  |  | GEPACS | [9 empty] | manual
|  | 10m AbdPel a240 :-D | [22 empty] | 0.6 | 0.6 | 400 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | [8 empty] | manual
6250_AbdPel_ho_tcc_w_fur (Adult)
6250_AbdPel_ho_tcc_w_fur (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T21f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T21f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| ven AbdPel |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 100 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven AbdPel a250 | [22 empty] | 3 | 3 | 104 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel cor | [22 empty] | 3 | 3 | 51 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel sag | [22 empty] | 3 | 3 | 121 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel a250 :-D | [22 empty] | 0.6 | 0.6 | 516 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
| 10m AbdPel |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 2 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | 10m AbdPel a250 | [22 empty] | 3 | 3 | 80 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | 10m AbdPel cor | [22 empty] | 3 | 3 | 67 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | 10m AbdPel cor 10x10 | [22 empty] | 10 | 10 | 39 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | 10m AbdPel cor mip | [22 empty] | 24 | 5 | 77 | I30f medium smooth | Abdomen MIP |  |  |  | GEPACS | [9 empty] | manual
|  | 10m AbdPel a250 :-D | [22 empty] | 0.6 | 0.6 | 400 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | [8 empty] | manual
6253_AbdPel_ili_vei_thr (Adult)
6253_AbdPel_ili_vei_thr (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T21f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T21f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| ven AbdPel |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 90 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven AbdPel a253 | [22 empty] | 3 | 3 | 104 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel cor | [22 empty] | 3 | 3 | 51 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel cor mip | [22 empty] | 16 | 5 | 31 | I30f medium smooth | Abdomen MIP |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel sag | [22 empty] | 3 | 3 | 121 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel a253 :-D | [22 empty] | 0.6 | 0.6 | 516 | I26f medium smooth ASA | 400/40 Abd |  |  |  | ROUTER | GEPACS | [8 empty] | manual
| 4m AbdPel |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 2 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | 4m AbdPel a253 | [22 empty] | 3 | 3 | 80 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | 4m AbdPel cor | [22 empty] | 3 | 3 | 67 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | 4m AbdPel cor mip | [22 empty] | 24 | 5 | 77 | I30f medium smooth | Abdomen MIP |  |  |  | GEPACS | [9 empty] | manual
|  | 4m AbdPel a253 :-D | [22 empty] | 0.6 | 0.6 | 400 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
6255_AbdPel_ili_vei_thr_ld (Adult)
6255_AbdPel_ili_vei_thr_ld (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T21f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T21f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| ven Abd |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 90 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven Abd a255 | [22 empty] | 3 | 3 | 104 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven Abd cor | [22 empty] | 3 | 3 | 51 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven Abd cor mip | [22 empty] | 16 | 5 | 31 | I30f medium smooth | Abdomen MIP |  |  |  | GEPACS | [9 empty] | manual
|  | ven Abd a255 :-D | [22 empty] | 0.6 | 0.6 | 516 | I26f medium smooth ASA | 400/40 Abd |  |  |  | ROUTER | GEPACS | [8 empty] | manual
| 4m AbdPel |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 2 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | 4m AbdPel a255 | [22 empty] | 3 | 3 | 80 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | 4m AbdPel cor | [22 empty] | 3 | 3 | 67 | I26f medium smooth ASA | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | 4m AbdPel sag | [22 empty] | 3 | 3 | 121 | I26f medium smooth ASA | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | 4m AbdPel cor mip | [22 empty] | 16 | 5 | 40 | I26f medium smooth ASA | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | 4m AbdPel a255 :-D | [22 empty] | 0.6 | 0.6 | 400 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | [8 empty] | manual
6257_AbdPel_pel_ven_con (Adult)
6257_AbdPel_pel_ven_con (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| 2m AbdPel |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 120 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | 2m AbdPel a257 | [22 empty] | 3 | 3 | 131 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | 2m AbdPel cor | [22 empty] | 3 | 3 | 67 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | 2m AbdPel cor mip | [22 empty] | 16 | 5 | 41 | I30f medium smooth | Abdomen MIP |  |  |  | GEPACS | [9 empty] | manual
|  | 2m AbdPel sag | [22 empty] | 3 | 3 | 121 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | 2m AbdPel a257 :-D | [22 empty] | 0.6 | 0.6 | 653 | I30f medium smooth | Abdomen MIP |  |  |  | ROUTER | GEPACS | [8 empty] | manual
6260_AbdPel_rec_gas (Adult)
6260_AbdPel_rec_gas (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| ven AbdPel |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 90 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven AbdPel a260 | [22 empty] | 3 | 3 | 131 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel cor | [22 empty] | 3 | 3 | 67 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel sag | [22 empty] | 3 | 3 | 121 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel a260 :-D | [22 empty] | 0.6 | 0.6 | 653 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | [8 empty] | manual
6265_AbdPel_rec_gas_wo_cst (Adult)
6265_AbdPel_rec_gas_wo_cst (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| AbdPel |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 6 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | AbdPel a265 | [22 empty] | 3 | 3 | 131 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | AbdPel cor | [22 empty] | 3 | 3 | 67 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | AbdPel sag | [22 empty] | 3 | 3 | 121 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | AbdPel a265 :-D | [22 empty] | 0.6 | 0.6 | 653 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | [8 empty] | manual
6270_AbdPel_ren_sto_wo_cst_A08_04A (Adult)
6270_AbdPel_ren_sto_wo_cst_A08_04A (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| AbdPel |  | 120 | 120 | 120 | 210 |  |  | On | 8 | On | CARE Dose4D | 8.1 | 30s | 25% | 50 |  | 0.5 |  | 4 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | AbdPel a270 | [22 empty] | 3 | 3 | 100 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | AbdPel cor | [22 empty] | 2 | 2 | 171 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | AbdPel cor mip | [22 empty] | 24 | 5 | 69 | I30f medium smooth | Abdomen MIP |  |  |  | GEPACS | [9 empty] | manual
|  | AbdPel sag | [22 empty] | 2 | 2 | 181 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | AbdPel a270 :-D | [22 empty] | 0.6 | 0.6 | 500 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | [8 empty] | manual
6272_AbdPel_ren_sto_qua_wo_cst (Adult)
6272_AbdPel_ren_sto_qua_wo_cst (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| AbdPel |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 4 | 0.8 |  | 2 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | AbdPel RSQ | [22 empty] | 2 | 2 | 150 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | AbdPel cor | [22 empty] | 3 | 3 | 114 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | AbdPel cor mip | [22 empty] | 24 | 5 | 69 | I30f medium smooth | Abdomen MIP |  |  |  | GEPACS | [9 empty] | manual
|  | AbdPel sag | [22 empty] | 3 | 3 | 101 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | AbdPel a272 :-D | [22 empty] | 0.6 | 0.6 | 500 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | [8 empty] | manual
6275_AbdPel_ren_sto_com_wo_cst (Adult)
6275_AbdPel_ren_sto_com_wo_cst (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T81f sharp | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T81f sharp | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| DE AbdPel |  |  | 100 | 240 |  | 202 | 156 |  | 3 | On | CARE Dose4D | 18.61 | 30s | 25% | 50 |  | 0.5 |  | 2 | 0.6 |  | 5 | 32x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | AbdPel a275 | [22 empty] | 5 | 2.5 | 106 | D40f medium | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | AbdPel a275 #PP | [22 empty] | 1 | 0.8 | 329 | D40f medium | 400/40 Abd |  |  |  | ROUTER | GEPACS | MMWP1 | [7 empty] | manual
|  | AbdPel a275 | [22 empty] | 2 | 2 | 132 | D40f medium | 400/40 Abd |  |  |  | GEPACS | MCH3DLabAW | [8 empty] | manual
|  | AbdPel cor | [22 empty] | 2 | 1.8 | 167 | D40f medium | 400/40 Abd |  |  |  | GEPACS | MCH3DLabAW | [8 empty] | manual
|  | AbdPel cor mip | [22 empty] | 24 | 24 | 13 | D40f medium | Abdomen MIP |  |  |  | GEPACS | MCH3DLabAW | [8 empty] | manual
|  | AbdPel sag | [22 empty] | 2 | 1.8 | 167 | D40f medium | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
6290_AbdPel_ret_ble_wo_cst (Adult)
6290_AbdPel_ret_ble_wo_cst (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| AbdPel |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 4 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | AbdPel a290 | [22 empty] | 3 | 3 | 131 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | AbdPel cor | [22 empty] | 3 | 3 | 67 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | AbdPel sag | [22 empty] | 3 | 3 | 121 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | AbdPel a290 :-D | [22 empty] | 0.6 | 0.6 | 653 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | [8 empty] | manual
6300_AbdPel_uro_w_fur (Adult)
6300_AbdPel_uro_w_fur (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T21f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T21f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Abd |  | 120 | 120 | 160 | 210 |  |  | On | 6 | On | CARE Dose4D | 10.8 | 30s | 25% | 50 |  | 0.5 |  | 5 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | Abd a300 | [22 empty] | 3 | 3 | 84 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | Abd a300 :-D | [22 empty] | 0.6 | 0.6 | 738 | I26f medium smooth ASA | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
| ven AbdPel |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 100 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven AbdPel a300 | [22 empty] | 3 | 3 | 104 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel cor | [22 empty] | 3 | 3 | 51 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel sag | [22 empty] | 3 | 3 | 121 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel a300 :-D | [22 empty] | 0.6 | 0.6 | 516 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | [8 empty] | manual
| 10m AbdPel |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 2 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | 10m AbdPel a300 | [22 empty] | 3 | 3 | 80 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | 10m AbdPel cor | [22 empty] | 3 | 3 | 67 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | 10m AbdPel cor 10x10 | [22 empty] | 10 | 10 | 39 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | 10m AbdPel cor mip | [22 empty] | 24 | 5 | 77 | I30f medium smooth | Abdomen MIP |  |  |  | GEPACS | [9 empty] | manual
|  | 10m AbdPel a300 :-D | [22 empty] | 0.6 | 0.6 | 400 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | [8 empty] | manual
6320_AbdPel_uro_mod (Adult)
6320_AbdPel_uro_mod (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T21f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T21f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| AbdPel |  | 120 | 120 | 160 | 210 |  |  | On | 6 | On | CARE Dose4D | 10.8 | 30s | 25% | 50 |  | 0.5 |  | 5 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | AbdPel a320 | [22 empty] | 3 | 3 | 148 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | AbdPel a320 :-D | [22 empty] | 0.6 | 0.6 | 738 | I26f medium smooth ASA | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
| ven10m AbdPel |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 100 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven10m AbdPel a320 | [22 empty] | 3 | 3 | 80 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven10m AbdPel cor | [22 empty] | 3 | 3 | 67 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven10m AbdPel cor 10x10 | [22 empty] | 10 | 10 | 39 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | ven10m AbdPel cor mip | [22 empty] | 24 | 5 | 77 | I30f medium smooth | Abdomen MIP |  |  |  | GEPACS | [9 empty] | manual
|  | ven10m AbdPel sag | [22 empty] | 3 | 3 | 121 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven10m AbdPel a320 :-D | [22 empty] | 0.6 | 0.6 | 400 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | [8 empty] | manual
6321_AbdPel_ven_her (Adult)
6321_AbdPel_ven_her (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| ven AbdPel |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 90 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven AbdPel a321 | [22 empty] | 3 | 3 | 131 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel cor | [22 empty] | 3 | 3 | 67 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel sag | [22 empty] | 3 | 3 | 121 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel a321 :-D | [22 empty] | 0.6 | 0.6 | 653 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
6324_AbdPel_ven_her_wo_cst (Adult)
6324_AbdPel_ven_her_wo_cst (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| AbdPel |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 4 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | AbdPel a324 | [22 empty] | 3 | 3 | 131 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | AbdPel cor | [22 empty] | 3 | 3 | 67 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | AbdPel sag | [22 empty] | 3 | 3 | 121 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | AbdPel a324 :-D | [22 empty] | 0.6 | 0.6 | 653 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
6325_Abd_cho (Adult)
6325_Abd_cho (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| 60m Abd |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 4 | 0.8 |  | 2 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | 60m Abd a325 | [22 empty] | 2 | 2 | 176 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | 60m Abd cor | [22 empty] | 2 | 2 | 191 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | 60m Abd sag | [22 empty] | 2 | 2 | 191 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | 60m Abd mip | [22 empty] | 24 | 5 | 71 | I30f medium smooth | Abdomen MIP |  |  |  | GEPACS | [9 empty] | manual
|  | 60m Abd cor mip | [22 empty] | 24 | 5 | 77 | I30f medium smooth | Abdomen MIP |  |  |  | GEPACS | [9 empty] | manual
|  | 60m Abd sag mip | [22 empty] | 24 | 5 | 77 | I30f medium smooth | Abdomen MIP |  |  |  | GEPACS | [9 empty] | manual
|  | 60m Abd a325 :-D | [22 empty] | 0.6 | 0.6 | 584 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
|  | vrt thins a325 | [22 empty] | 0.6 | 0.6 | 584 | I26f medium smooth ASA | 400/40 Abd |  |  |  | MCH3DLabAW | [9 empty] | manual
6330_ang_Abd_lld (Adult)
6330_ang_Abd_lld (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Abd |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 2 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | Abd a330 | [22 empty] | 3 | 3 | 131 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | Abd a330 :-D | [22 empty] | 0.6 | 0.6 | 653 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
| PreMonitoring |  | 120 | 120 | 350 | 50 |  |  | Off | 3 | Off | CARE Dose4D | 3.08 |  | 25% |  |  | 0.5 |  | 2 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | PreMonitoring | [22 empty] | 10 |  | 1 | B40s medium | 400/40 Abd | [13 empty] | manual
| Monitoring |  | 120 | 120 | 350 | 50 |  |  | Off | 3 | Off | CARE Dose4D | 123.28 |  | 25% |  |  | 0.5 |  | 10 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | Monitoring | [22 empty] | 10 |  | 40 | B40s medium | 400/40 Abd | [13 empty] | manual
| art Abd |  | 120 | 120 | 180 | 210 |  |  | On | 8 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 10 | 0.8 |  | 2 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | art Abd a330 | [22 empty] | 2 | 2 | 120 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd cor | [22 empty] | 2 | 2 | 191 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd sag | [22 empty] | 2 | 2 | 191 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd mip | [22 empty] | 24 | 4 | 96 | I30f medium smooth | Abdomen MIP |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd cor mip | [22 empty] | 24 | 4 | 96 | I30f medium smooth | Abdomen MIP |  |  |  | GEPACS | [9 empty] | manual
|  | vrt thins a330 | [22 empty] | 0.6 | 0.6 | 400 | I26f medium smooth ASA | Abdomen |  |  |  | MCH3DLabAW | [9 empty] | manual
|  | art Abd a330 :-D | [22 empty] | 0.6 | 0.6 | 400 | I30f medium smooth | Abdomen |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
| ven Abd |  | 120 | 120 | 350 | 210 |  |  | On | 6 | On | CARE Dose4D | 23.65 | 30s | 25% | 50 |  | 0.5 |  | 50 | 0.6 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven Abd a330 | [22 empty] | 3 | 3 | 79 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven Abd cor | [22 empty] | 3 | 3 | 127 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven Abd mip | [22 empty] | 24 | 4 | 60 | I30f medium smooth | Abdomen MIP |  |  |  | GEPACS | [9 empty] | manual
|  | ven Abd cor mip | [22 empty] | 24 | 4 | 96 | I30f medium smooth | Abdomen MIP |  |  |  | GEPACS | [9 empty] | manual
|  | ven Abd a330 :-D | [22 empty] | 0.6 | 0.6 | 395 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
6335_ang_Abd_lld_Pel (Adult)
6335_ang_Abd_lld_Pel (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Abd |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 2 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | Abd a335 | [22 empty] | 3 | 3 | 131 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | Abd a335 :-D | [22 empty] | 0.6 | 0.6 | 653 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
| PreMonitoring |  | 120 | 120 | 350 | 50 |  |  | Off | 3 | Off | CARE Dose4D | 3.08 |  | 25% |  |  | 0.5 |  | 2 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | PreMonitoring | [22 empty] | 10 |  | 1 | B40s medium | 400/40 Abd | [13 empty] | manual
| Monitoring |  | 120 | 120 | 350 | 50 |  |  | Off | 3 | Off | CARE Dose4D | 123.28 |  | 25% |  |  | 0.5 |  | 10 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | Monitoring | [22 empty] | 10 |  | 40 | B40s medium | 400/40 Abd | [13 empty] | manual
| art Abd |  | 120 | 120 | 180 | 210 |  |  | On | 8 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 10 | 0.8 |  | 2 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | art Abd a335 | [22 empty] | 2 | 2 | 120 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd cor | [22 empty] | 2 | 2 | 191 | I30f medium smooth | 400/40 Abd | [13 empty] | manual
|  | art Abd sag | [22 empty] | 2 | 2 | 191 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd mip | [22 empty] | 24 | 5 | 77 | I30f medium smooth | Abdomen MIP |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd cor mip | [22 empty] | 24 | 5 | 77 | I30f medium smooth | Abdomen MIP | [13 empty] | manual
|  | art Abd a335 :-D | [22 empty] | 0.6 | 0.6 | 400 | I30f medium smooth | Abdomen MIP |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
|  | vrt thins a335 | [22 empty] | 0.6 | 0.6 | 400 | I26f medium smooth ASA | Abdomen MIP |  |  |  | MCH3DLabAW | [9 empty] | manual
| ven AbdPel |  | 120 | 120 | 350 | 210 |  |  | On | 6 | On | CARE Dose4D | 23.65 | 30s | 25% | 50 |  | 0.5 |  | 50 | 0.6 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven AbdPel a335 | [22 empty] | 3 | 2 | 119 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel cor | [22 empty] | 3 | 3 | 127 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel mip | [22 empty] | 24 | 5 | 48 | I30f medium smooth | Abdomen MIP |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel cor mip | [22 empty] | 24 | 5 | 77 | I30f medium smooth | Abdomen MIP |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel sag | [22 empty] | 3 | 3 | 127 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel a335 :-D | [22 empty] | 0.6 | 0.6 | 395 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | [8 empty] | manual
6341_ang_Abd_lrd_Pel_nonVEO (Adult)
6341_ang_Abd_lrd_Pel_nonVEO (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T21f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T21f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| AbdPel |  | 120 | 120 | 180 | 210 |  |  | On | 8 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 5 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | AbdPel a285 | [22 empty] | 3 | 3 | 317 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | AbdPel a285 :-D | [22 empty] | 0.6 | 0.6 | 1585 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
| PreMonitoring |  | 120 | 120 | 240 | 50 |  |  | Off | 3 | Off | CARE Dose4D | 3.08 |  | 25% |  |  | 0.5 |  | 2 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | PreMonitoring | [22 empty] | 10 |  | 1 | B40s medium | 400/40 Abd | [13 empty] | manual
| Monitoring |  | 120 | 120 | 240 | 50 |  |  | Off | 3 | Off | CARE Dose4D | 123.28 |  | 25% |  |  | 0.5 |  | 10 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | Monitoring | [22 empty] | 10 |  | 40 | B40s medium | 400/40 Abd | [13 empty] | manual
| art Abd |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 5 | 0.8 |  | 2 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | art Abd a340 | [22 empty] | 2 | 2 | 438 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd cor | [22 empty] | 2 | 2 | 78 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd sag | [22 empty] | 2 | 2 | 185 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd cor mip | [22 empty] | 24 | 5 | 73 | I30f medium smooth | Abdomen MIP |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd a340 :-D | [22 empty] | 0.6 | 0.6 | 1460 | I26f medium smooth ASA | Abdomen |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
| ven Abd |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 60 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven Abd a340 | [22 empty] | 3 | 3 | 277 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven Abd a340 :-D | [22 empty] | 0.6 | 0.6 | 1383 | I26f medium smooth ASA | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
| 10m AbdPel |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 2 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | 10m AbdPel a340 | [22 empty] | 3 | 3 | 277 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | 10m AbdPel cor | [22 empty] | 3 | 3 | 121 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | 10m AbdPel cor mip | [22 empty] | 24 | 5 | 73 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | 10m AbdPel a340 :-D | [22 empty] | 0.6 | 0.6 | 1383 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
6351_ang_Abd_pan_bi_nonVEO (Adult)
6351_ang_Abd_pan_bi_nonVEO (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| PreMonitoring |  | 120 | 120 | 350 | 50 |  |  | Off | 3 | Off | CARE Dose4D | 3.08 |  | 25% |  |  | 0.5 |  | 2 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | PreMonitoring | [22 empty] | 10 |  | 1 | B40s medium | 400/40 Abd | [13 empty] | manual
| Monitoring |  | 120 | 120 | 350 | 50 |  |  | Off | 3 | Off | CARE Dose4D | 123.28 |  | 25% |  |  | 0.5 |  | 10 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | Monitoring | [22 empty] | 10 |  | 40 | B40s medium | 400/40 Abd | [13 empty] | manual
| art Abd |  | 120 | 80 | 700 | 210 |  |  | Off | 8 | On | CARE Dose4D | 12.7 | 30s | 50% | 50 |  | 0.5 |  | 18 | 0.5 |  | 2 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | art Abd a350 | [22 empty] | 2 | 1 | 240 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd cor | [22 empty] | 2 | 2 | 191 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd cor mip | [22 empty] | 24 | 5 | 77 | I30f medium smooth | Abdomen MIP |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd a350 :-D | [22 empty] | 0.6 | 0.6 | 400 | I30f medium smooth | Abdomen |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
| ven Abd |  | 120 | 80 | 700 | 210 |  |  | Off | 6 | On | CARE Dose4D | 12.7 | 30s | 50% | 50 |  | 0.5 |  | 15 | 0.5 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven Abd a350 | [22 empty] | 3 | 3 | 80 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven Abd cor | [22 empty] | 3 | 3 | 127 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven Abd cor mip | [22 empty] | 24 | 5 | 77 | I30f medium smooth | Abdomen MIP |  |  |  | GEPACS | [9 empty] | manual
|  | ven Abd a350 :-D | [22 empty] | 0.6 | 0.6 | 400 | I30f medium smooth | Abdomen |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
6361_ang_Abd_pan_bi_Pel_nonVEO (Adult)
6361_ang_Abd_pan_bi_Pel_nonVEO (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| PreMonitoring |  | 120 | 120 | 350 | 50 |  |  | Off | 3 | Off | CARE Dose4D | 3.08 |  | 25% |  |  | 0.5 |  | 2 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | PreMonitoring | [22 empty] | 10 |  | 1 | B40s medium | 400/40 Abd | [13 empty] | manual
| Monitoring |  | 120 | 120 | 350 | 50 |  |  | Off | 3 | Off | CARE Dose4D | 123.28 |  | 25% |  |  | 0.5 |  | 10 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | Monitoring | [22 empty] | 10 |  | 40 | B40s medium | 400/40 Abd | [13 empty] | manual
| art Abd |  | 120 | 80 | 700 | 210 |  |  | Off | 8 | On | CARE Dose4D | 12.7 | 30s | 50% | 50 |  | 0.5 |  | 18 | 0.5 |  | 2 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | art Abd a360 | [22 empty] | 2 | 1 | 240 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd cor | [22 empty] | 2 | 2 | 191 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd cor mip | [22 empty] | 24 | 5 | 77 | I30f medium smooth | Abdomen MIP |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd a360 | [22 empty] | 0.6 | 0.6 | 400 | I30f medium smooth | Abdomen |  |  |  | MCHNeuroBodyAW | [9 empty] | manual
| ven AbdPel |  | 120 | 80 | 700 | 7 |  |  | Off | 6 | On | CARE Dose4D | 12.7 | 30s | 50% | 50 |  | 0.5 |  | 15 | 0.5 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven AbdPel a360 | [22 empty] | 3 | 3 | 80 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel cor | [22 empty] | 3 | 3 | 127 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel cor mip | [22 empty] | 24 | 5 | 77 | I30f medium smooth | Abdomen MIP |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel sag | [22 empty] | 3 | 3 | 127 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel a360 :-D | [22 empty] | 0.6 | 0.6 | 400 | I30f medium smooth | Abdomen |  |  |  | ROUTER | GEPACS | [8 empty] | manual
6370_ang_Abd_ren (Adult)
6370_ang_Abd_ren (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| PreMonitoring |  | 120 | 120 | 350 | 50 |  |  | Off | 3 | Off | CARE Dose4D | 3.08 |  | 25% |  |  | 0.5 |  | 2 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | PreMonitoring | [22 empty] | 10 |  | 1 | B40s medium | 400/40 Abd | [13 empty] | manual
| Monitoring |  | 120 | 120 | 350 | 50 |  |  | Off | 3 | Off | CARE Dose4D | 123.28 |  | 25% |  |  | 0.5 |  | 10 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | Monitoring | [22 empty] | 10 |  | 40 | B40s medium | 400/40 Abd | [13 empty] | manual
| art Abd |  | 120 | 120 | 180 | 210 |  |  | On | 8 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 6 | 0.8 |  | 2 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | art Abd a370 | [22 empty] | 2 | 2 | 120 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd cor | [22 empty] | 2 | 2 | 191 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd sag | [22 empty] | 2 | 2 | 191 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd cor mip | [22 empty] | 24 | 5 | 77 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd a370 :-D | [22 empty] | 0.6 | 0.6 | 400 | I26f medium smooth ASA | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
| ven Abd |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 50 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven Abd a370 | [22 empty] | 3 | 3 | 80 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven Abd cor | [22 empty] | 3 | 3 | 127 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven Abd a370 :-D | [22 empty] | 0.6 | 0.6 | 400 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
6380_ang_Abd_ren_Pel (Adult)
6380_ang_Abd_ren_Pel (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| PreMonitoring |  | 120 | 120 | 350 | 50 |  |  | Off | 3 | Off | CARE Dose4D | 3.08 |  | 25% |  |  | 0.5 |  | 2 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | PreMonitoring | [22 empty] | 10 |  | 1 | B40s medium | 400/40 Abd | [13 empty] | manual
| Monitoring |  | 120 | 120 | 350 | 50 |  |  | Off | 3 | Off | CARE Dose4D | 123.28 |  | 25% |  |  | 0.5 |  | 10 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | Monitoring | [22 empty] | 10 |  | 40 | B40s medium | 400/40 Abd | [13 empty] | manual
| art Abd |  | 120 | 120 | 180 | 210 |  |  | On | 8 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 7 | 0.8 |  | 2 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | art Abd a380 | [22 empty] | 2 | 2 | 120 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd cor | [22 empty] | 2 | 2 | 191 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd sag | [22 empty] | 2 | 2 | 191 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd cor mip | [22 empty] | 24 | 5 | 77 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd a380 :-D | [22 empty] | 0.6 | 0.6 | 400 | I26f medium smooth ASA | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
| ven AbdPel |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 50 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven AbdPel a380 | [22 empty] | 3 | 3 | 80 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel cor | [22 empty] | 3 | 3 | 127 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel sag | [22 empty] | 3 | 3 | 127 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel a380 :-D | [22 empty] | 0.6 | 0.6 | 400 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
6390_ang_AbdPel_diep (Adult)
6390_ang_AbdPel_diep (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| PreMonitoring |  | 120 | 120 | 350 | 50 |  |  | Off | 3 | Off | CARE Dose4D | 3.08 |  | 25% |  |  | 0.5 |  | 2 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | PreMonitoring | [22 empty] | 10 |  | 1 | B40s medium | 400/40 Abd | [13 empty] | manual
| Monitoring |  | 120 | 120 | 350 | 50 |  |  | Off | 3 | Off | CARE Dose4D | 123.28 |  | 25% |  |  | 0.5 |  | 10 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | Monitoring | [22 empty] | 10 |  | 40 | B40s medium | 400/40 Abd | [13 empty] | manual
| art AbdPel |  | 120 | 120 | 180 | 210 |  |  | On | 8 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 20 | 0.8 |  | 2 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | art AbdPel a390 | [22 empty] | 2 | 2 | 120 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art AbdPel cor | [22 empty] | 2 | 2 | 191 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art AbdPel sag | [22 empty] | 2 | 2 | 191 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art AbdPel mip | [22 empty] | 20 | 10 | 24 | I30f medium smooth | Abdomen MIP |  |  |  | GEPACS | [9 empty] | manual
|  | art AbdPel cor mip | [22 empty] | 20 | 10 | 39 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art AbdPel sag mip | [22 empty] | 20 | 10 | 39 | I30f medium smooth | Abdomen MIP |  |  |  | GEPACS | [9 empty] | manual
|  | art AbdPel a390 :-D | [22 empty] | 0.6 | 0.6 | 400 | I26f medium smooth ASA | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
6400_ang_AbdPel_acu_gas_ble (Adult)
6400_ang_AbdPel_acu_gas_ble (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| AbdPel |  | 120 | 100 | 180 | 210 |  |  | Off | 6 | On | CARE Dose4D | 7.41 | 30s | 25% | 50 |  | 0.5 |  | 2 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | AbdPel a400 | [22 empty] | 3 | 2 | 235 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | AbdPel cor | [22 empty] | 3 | 2 | 104 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | AbdPel a400 :-D | [22 empty] | 0.6 | 0.6 | 781 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
|  | AbdPel cor mip | [22 empty] | 16 | 5 | 77 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
| PreMonitoring |  | 120 | 120 | 240 | 50 |  |  | Off | 8 | Off | CARE Dose4D | 3.08 |  | 25% |  |  | 0.5 |  | 2 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | PreMonitoring | [22 empty] | 10 |  | 1 | B40s medium | 400/40 Abd | [13 empty] | manual
| Monitoring |  | 120 | 120 | 240 | 50 |  |  | Off | 8 | Off | CARE Dose4D | 123.28 |  | 25% |  |  | 0.5 |  | 10 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | Monitoring | [22 empty] | 10 |  | 40 | B40s medium | 400/40 Abd | [13 empty] | manual
| art AbdPel |  | 120 | 100 | 180 | 210 |  |  | Off | 8 | On | CARE Dose4D | 7.41 | 30s | 25% | 50 |  | 0.5 |  | 5 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | art AbdPel a400 | [22 empty] | 3 | 2 | 235 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art AbdPel cor | [22 empty] | 3 | 2 | 104 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art AbdPel sag | [22 empty] | 2 | 2 | 191 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art AbdPel cor mip | [22 empty] | 16 | 5 | 77 | I30f medium smooth | Abdomen MIP |  |  |  | GEPACS | [9 empty] | manual
|  | art AbdPel a400 :-D | [22 empty] | 0.6 | 0.6 | 781 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
| ven AbdPel |  | 120 | 100 | 180 | 210 |  |  | Off | 6 | On | CARE Dose4D | 7.41 | 30s | 25% | 50 |  | 0.5 |  | 65 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven AbdPel a400 | [22 empty] | 3 | 2 | 235 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel cor | [22 empty] | 3 | 2 | 104 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel cor mip | [22 empty] | 16 | 5 | 42 | I30f medium smooth | Abdomen MIP |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel a400 :-D | [22 empty] | 0.6 | 0.6 | 781 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
6405_ang_AbdPel_chr_gas_ble (Adult)
6405_ang_AbdPel_chr_gas_ble (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| AbdPel |  | 120 | 100 | 180 | 210 |  |  | Off | 6 | On | CARE Dose4D | 7.41 | 30s | 25% | 50 |  | 0.5 |  | 2 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | AbdPel a405 | [22 empty] | 3 | 2 | 235 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | AbdPel cor | [22 empty] | 3 | 2 | 104 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | AbdPel cor mip | [22 empty] | 16 | 5 | 42 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | AbdPel a405 :-D | [22 empty] | 0.6 | 0.6 | 781 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | [8 empty] | manual
| PreMonitoring |  | 120 | 120 | 240 | 50 |  |  | Off | 8 | Off | CARE Dose4D | 3.08 |  | 25% |  |  | 0.5 |  | 2 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | PreMonitoring | [22 empty] | 10 |  | 1 | B40s medium | 400/40 Abd | [13 empty] | manual
| Monitoring |  | 120 | 120 | 240 | 50 |  |  | Off | 8 | Off | CARE Dose4D | 123.28 |  | 25% |  |  | 0.5 |  | 10 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | Monitoring | [22 empty] | 10 |  | 40 | B40s medium | 400/40 Abd | [13 empty] | manual
| art AbdPel |  | 120 | 100 | 180 | 210 |  |  | Off | 8 | On | CARE Dose4D | 7.41 | 30s | 25% | 50 |  | 0.5 |  | 5 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | art AbdPel a405 | [22 empty] | 3 | 2 | 235 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art AbdPel cor | [22 empty] | 3 | 2 | 104 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art AbdPel sag | [22 empty] | 2 | 2 | 191 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art AbdPel cor mip | [22 empty] | 16 | 5 | 77 | I30f medium smooth | Abdomen MIP |  |  |  | GEPACS | [9 empty] | manual
|  | art AbdPel a405 :-D | [22 empty] | 0.6 | 0.6 | 781 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
| ven AbdPel |  | 120 | 100 | 180 | 210 |  |  | Off | 6 | On | CARE Dose4D | 7.41 | 30s | 25% | 50 |  | 0.5 |  | 65 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven AbdPel a405 | [22 empty] | 3 | 2 | 235 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel cor | [22 empty] | 3 | 2 | 104 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel cor mip | [22 empty] | 16 | 5 | 42 | I30f medium smooth | Abdomen MIP |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel a405 :-D | [22 empty] | 0.6 | 0.6 | 781 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
6407_ang_AbdPel_mes_art (Adult)
6407_ang_AbdPel_mes_art (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| PreMonitoring |  | 120 | 120 | 240 | 50 |  |  | Off | 8 | Off | CARE Dose4D | 3.08 |  | 25% |  |  | 0.5 |  | 2 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | PreMonitoring | [22 empty] | 10 |  | 1 | B40s medium | 400/40 Abd | [13 empty] | manual
| Monitoring |  | 120 | 120 | 240 | 50 |  |  | Off | 8 | Off | CARE Dose4D | 123.28 |  | 25% |  |  | 0.5 |  | 10 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | Monitoring | [22 empty] | 10 |  | 40 | B40s medium | 400/40 Abd | [13 empty] | manual
| art AbdPel |  | 120 | 120 | 180 | 210 |  |  | On | 8 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 5 | 0.8 |  | 2 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | art AbdPel a407 | [22 empty] | 2 | 2 | 235 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art AbdPel cor | [22 empty] | 2 | 2 | 104 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art AbdPel sag | [22 empty] | 2 | 2 | 191 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art AbdPel cor mip | [22 empty] | 24 | 5 | 77 | I30f medium smooth | Abdomen MIP |  |  |  | GEPACS | [9 empty] | manual
|  | art AbdPel 1x1 | [22 empty] | 1 | 1 | 469 | I26f medium smooth ASA | Abdomen |  |  |  | ROUTER | MCH3DLabAW | [8 empty] | manual
|  | art AbdPel a407 :-D | [22 empty] | 0.6 | 0.6 | 781 | I26f medium smooth ASA | Abdomen MIP |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
| ven AbdPel |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 65 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven AbdPel a407 | [22 empty] | 3 | 3 | 157 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel cor | [22 empty] | 3 | 3 | 70 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel a407 :-D | [22 empty] | 0.6 | 0.6 | 781 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
6410_ang_AbdPel_mes_art_ld (Adult)
6410_ang_AbdPel_mes_art_ld (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| PreMonitoring |  | 120 | 120 | 240 | 50 |  |  | Off | 8 | Off | CARE Dose4D | 3.08 |  | 25% |  |  | 0.5 |  | 2 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | PreMonitoring | [22 empty] | 10 |  | 1 | B40s medium | 400/40 Abd | [13 empty] | manual
| Monitoring |  | 120 | 120 | 240 | 50 |  |  | Off | 8 | Off | CARE Dose4D | 123.28 |  | 25% |  |  | 0.5 |  | 10 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | Monitoring | [22 empty] | 10 |  | 40 | B40s medium | 400/40 Abd | [13 empty] | manual
| art AbdPel |  | 120 | 120 | 180 | 210 |  |  | On | 8 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 5 | 0.8 |  | 2 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | art AbdPel a410 | [22 empty] | 2 | 2 | 235 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art AbdPel cor | [22 empty] | 2 | 2 | 104 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art AbdPel sag | [22 empty] | 2 | 2 | 191 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art AbdPel cor mip | [22 empty] | 24 | 5 | 77 | I30f medium smooth | Abdomen MIP |  |  |  | GEPACS | [9 empty] | manual
|  | art AbdPel 1x1 | [22 empty] | 1 | 1 | 469 | I26f medium smooth ASA | Abdomen |  |  |  | ROUTER | MCH3DLabAW | [8 empty] | manual
|  | art AbdPel a410 :-D | [22 empty] | 0.6 | 0.6 | 781 | I26f medium smooth ASA | Abdomen MIP |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
| ven Abd |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 65 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven Abd a410 | [22 empty] | 3 | 3 | 157 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven Abd cor | [22 empty] | 3 | 3 | 70 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven Abd a410 :-D | [22 empty] | 0.6 | 0.6 | 781 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
6415_ang_Abd_spl_art (Adult)
6415_ang_Abd_spl_art (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| PreMonitoring |  | 120 | 120 | 240 | 50 |  |  | Off | 8 | Off | CARE Dose4D | 3.08 |  | 25% |  |  | 0.5 |  | 2 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | PreMonitoring | [22 empty] | 10 |  | 1 | B40s medium | 400/40 Abd | [13 empty] | manual
| Monitoring |  | 120 | 120 | 240 | 50 |  |  | Off | 8 | Off | CARE Dose4D | 123.28 |  | 25% |  |  | 0.5 |  | 10 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | Monitoring | [22 empty] | 10 |  | 40 | B40s medium | 400/40 Abd | [13 empty] | manual
| art Abd |  | 120 | 120 | 180 | 210 |  |  | On | 8 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 5 | 0.8 |  | 2 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | art Abd a415 | [22 empty] | 2 | 2 | 235 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd cor | [22 empty] | 2 | 2 | 104 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd sag | [22 empty] | 2 | 2 | 191 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd cor mip | [22 empty] | 24 | 5 | 77 | I30f medium smooth | Abdomen MIP |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd 1x1 | [22 empty] | 1 | 1 | 469 | I26f medium smooth ASA | Abdomen |  |  |  | ROUTER | MCH3DLabAW | [8 empty] | manual
|  | art Abd a415 :-D | [22 empty] | 0.6 | 0.6 | 781 | I26f medium smooth ASA | Abdomen MIP |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
| ven Abd |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 65 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven Abd a415 | [22 empty] | 3 | 3 | 157 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven Abd cor | [22 empty] | 3 | 3 | 70 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven Abd a415 :-D | [22 empty] | 0.6 | 0.6 | 781 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
6420_ang_AbdPel_pos_ste (Adult)
6420_ang_AbdPel_pos_ste (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| AbdPel |  | 120 | 120 | 100 | 210 |  |  | On | 6 | On | CARE Dose4D | 6.75 | 30s | 25% | 50 |  | 0.5 |  | 3 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | AbdPel a420 | [22 empty] | 3 | 3 | 131 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | AbdPel a420 :-D | [22 empty] | 0.6 | 0.6 | 653 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
| PreMonitoring |  | 120 | 120 | 350 | 50 |  |  | Off | 3 | Off | CARE Dose4D | 3.08 |  | 25% |  |  | 0.5 |  | 2 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | PreMonitoring | [22 empty] | 10 |  | 1 | B40s medium | 400/40 Abd | [13 empty] | manual
| Monitoring |  | 120 | 120 | 350 | 50 |  |  | Off | 3 | Off | CARE Dose4D | 123.28 |  | 25% |  |  | 0.5 |  | 10 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | Monitoring | [22 empty] | 10 |  | 40 | B40s medium | 400/40 Abd | [13 empty] | manual
| art AbdPel |  | 120 | 120 | 180 | 210 |  |  | On | 8 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 6 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | art AbdPel a420 | [22 empty] | 3 | 3 | 80 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art AbdPel cor | [22 empty] | 2 | 2 | 191 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd sag | [22 empty] | 2 | 2 | 191 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art AbdPel cor mip | [22 empty] | 24 | 5 | 77 | I30f medium smooth | Abdomen MIP |  |  |  | GEPACS | [9 empty] | manual
|  | art AbdPel a420 :-D | [22 empty] | 0.6 | 0.6 | 400 | I26f medium smooth ASA | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
| ven AbdPel |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 40 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven AbdPel a420 | [22 empty] | 3 | 3 | 80 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel cor | [22 empty] | 2 | 2 | 191 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel a420 :-D | [22 empty] | 0.6 | 0.6 | 400 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
6430_ang_AbdPel_pos_ste_ld (Adult)
6430_ang_AbdPel_pos_ste_ld (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| AbdPel |  | 120 | 120 | 100 | 210 |  |  | On | 6 | On | CARE Dose4D | 6.75 | 30s | 25% | 50 |  | 0.5 |  | 3 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | AbdPel a430 | [22 empty] | 3 | 3 | 131 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | AbdPel a430 :-D | [22 empty] | 0.6 | 0.6 | 653 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
| PreMonitoring |  | 120 | 120 | 350 | 50 |  |  | Off | 3 | Off | CARE Dose4D | 3.08 |  | 25% |  |  | 0.5 |  | 2 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | PreMonitoring | [22 empty] | 10 |  | 1 | B40s medium | 400/40 Abd | [13 empty] | manual
| Monitoring |  | 120 | 120 | 350 | 50 |  |  | Off | 3 | Off | CARE Dose4D | 123.28 |  | 25% |  |  | 0.5 |  | 10 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | Monitoring | [22 empty] | 10 |  | 40 | B40s medium | 400/40 Abd | [13 empty] | manual
| art AbdPel |  | 120 | 120 | 180 | 210 |  |  | On | 8 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 6 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | art AbdPel a430 | [22 empty] | 3 | 3 | 80 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art AbdPel cor | [22 empty] | 2 | 2 | 191 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd sag | [22 empty] | 2 | 2 | 191 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art AbdPel cor mip | [22 empty] | 24 | 5 | 77 | I30f medium smooth | Abdomen MIP |  |  |  | GEPACS | [9 empty] | manual
|  | art AbdPel a430 :-D | [22 empty] | 0.6 | 0.6 | 400 | I26f medium smooth ASA | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
| ven AbdPel |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 40 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven AbdPel a430 | [22 empty] | 3 | 3 | 80 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel cor | [22 empty] | 2 | 2 | 191 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel a430 :-D | [22 empty] | 0.6 | 0.6 | 400 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
6435_ang_AbdPel_pre_ren_tra (Adult)
6435_ang_AbdPel_pre_ren_tra (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| PreMonitoring |  | 120 | 120 | 240 | 50 |  |  | Off | 8 | Off | CARE Dose4D | 3.08 |  | 25% |  |  | 0.5 |  | 2 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | PreMonitoring | [22 empty] | 10 |  | 1 | B40s medium | 400/40 Abd | [13 empty] | manual
| Monitoring |  | 120 | 120 | 240 | 50 |  |  | Off | 8 | Off | CARE Dose4D | 123.28 |  | 25% |  |  | 0.5 |  | 10 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | Monitoring | [22 empty] | 10 |  | 40 | B40s medium | 400/40 Abd | [13 empty] | manual
| art AbdPel |  | 120 | 120 | 180 | 210 |  |  | On | 8 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 5 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | art AbdPel a435 | [22 empty] | 3 | 3 | 157 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art AbdPel cor | [22 empty] | 2 | 2 | 104 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art AbdPel sag | [22 empty] | 2 | 2 | 191 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art AbdPel cor mip | [22 empty] | 24 | 5 | 77 | I30f medium smooth | Abdomen MIP |  |  |  | GEPACS | [9 empty] | manual
|  | art AbdPel a435 :-D | [22 empty] | 0.6 | 0.6 | 781 | I26f medium smooth ASA | Abdomen MIP |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
| ven Abd |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 60 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven Abd a435 | [22 empty] | 3 | 3 | 157 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven Abd cor | [22 empty] | 3 | 3 | 70 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven Abd a435 :-D | [22 empty] | 0.6 | 0.6 | 781 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
| 4m Pel |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 4 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | 4m Pel a435 | [22 empty] | 3 | 3 | 95 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | 4m Pel cor | [22 empty] | 3 | 3 | 121 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | 4m Pel a435 :-D | [22 empty] | 0.6 | 0.6 | 474 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
6440_ang_AbdPel_pre_ste (Adult)
6440_ang_AbdPel_pre_ste (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| PreMonitoring |  | 120 | 120 | 350 | 50 |  |  | Off | 3 | Off | CARE Dose4D | 3.08 |  | 25% |  |  | 0.5 |  | 2 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | PreMonitoring | [22 empty] | 10 |  | 1 | B40s medium | 400/40 Abd | [13 empty] | manual
| Monitoring |  | 120 | 120 | 350 | 50 |  |  | Off | 3 | Off | CARE Dose4D | 123.28 |  | 25% |  |  | 0.5 |  | 10 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | Monitoring | [22 empty] | 10 |  | 40 | B40s medium | 400/40 Abd | [13 empty] | manual
| art AbdPel |  | 120 | 120 | 180 | 210 |  |  | On | 8 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 6 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | art AbdPel a440 | [22 empty] | 3 | 3 | 80 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art AbdPel cor | [22 empty] | 2 | 2 | 191 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art AbdPel sag | [22 empty] | 2 | 2 | 191 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art AbdPel cor mip | [22 empty] | 24 | 5 | 77 | I30f medium smooth | Abdomen MIP |  |  |  | GEPACS | [9 empty] | manual
|  | art AbdPel a440 :-D | [22 empty] | 0.6 | 0.6 | 400 | I26f medium smooth ASA | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
| ven AbdPel |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 40 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven AbdPel a440 | [22 empty] | 3 | 3 | 80 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel cor | [22 empty] | 3 | 3 | 127 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel a440 :-D | [22 empty] | 0.6 | 0.6 | 400 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
6450_ang_AbdPelThiLeg_run (Adult)
6450_ang_AbdPelThiLeg_run (Adult) | Topogram | Topogram |  | 100 |  | 994 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 994 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| PreMonitoring |  | 120 | 120 | 350 | 50 |  |  | Off | 3 | Off | CARE Dose4D | 3.08 |  | 25% |  |  | 0.5 |  | 2 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | PreMonitoring | [22 empty] | 10 |  | 1 | B40s medium | 400/40 Abd | [13 empty] | manual
| Monitoring |  | 120 | 120 | 350 | 50 |  |  | Off | 3 | Off | CARE Dose4D | 123.28 |  | 25% |  |  | 0.5 |  | 10 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | Monitoring | [22 empty] | 10 |  | 40 | B40s medium | 400/40 Abd | [13 empty] | manual
| artAbdPelThiLe |  | 120 | 120 | 180 | 210 |  |  | On | 8 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 6 | 0.8 |  | 2 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | art AbdPelThiLeg a450 | [22 empty] | 2 | 2 | 120 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art AbdPelThiLeg a450 :-D | [22 empty] | 0.6 | 0.6 | 400 | I26f medium smooth ASA | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
| art Leg |  | 120 | 120 | 180 | 210 |  |  | On | 8 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 2 | 0.8 |  | 2 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | art Leg a450 | [22 empty] | 2 | 2 | 120 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art Leg a450 :-D | [22 empty] | 0.6 | 0.6 | 400 | I26f medium smooth ASA | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
| ven Abd |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 40 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven Abd a450 | [22 empty] | 3 | 3 | 80 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven Abd cor | [22 empty] | 3 | 3 | 127 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven Abd a450 :-D | [22 empty] | 0.6 | 0.6 | 400 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
AbdMultiPhase (Adult)
AbdMultiPhase (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body |  |  |  | GEPACS | ROEMIR001a | TRACKER
| Non Contrast |  | 120 | 120 | 190 | 190 |  |  | On | 3 | On | CARE Dose4D | 13.14 | 20s | 25% |  |  | 0.5 |  | 2 | 0.6 |  | 5 | 32x1.2mm | [15 empty] | Off
|  | Non Contrast 5.0 B30f | [22 empty] | 5 | 5 | 41 | B30f medium smooth | Abdomen | [13 empty] | manual
| Arterial Phase |  | 120 | 120 | 210 | 210 |  |  | On | 7 | On | CARE Dose4D | 14.19 | 20s | 25% |  |  | 0.5 |  | 2 | 0.6 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | Arterial Phase 5.0 B30f | [22 empty] | 5 | 5 | 41 | B30f medium smooth | Abdomen | [13 empty] | manual
|  | Arterial Phase 1.0 B20f | [22 empty] | 1 | 0.7 | 286 | B20f smooth | Abdomen | [13 empty] | manual
| Venous Phase |  | 120 | 120 | 190 | 190 |  |  | On | 7 | On | CARE Dose4D | 13.14 | 20s | 25% |  |  | 0.5 |  | 25 | 0.6 |  | 5 | 32x1.2mm | [15 empty] | Off
|  | Venous Phase 5.0 B30f | [22 empty] | 5 | 5 | 41 | B30f medium smooth | Abdomen | [13 empty] | manual
AbdMultiPhase_Customized (Adult)
AbdMultiPhase_Customized (Adult) | Topogram | Topogram 0.6 T20s |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body |  |  |  | GEPACS | ROEMIR001a | TRACKER
| Non Contrast |  | 120 | 120 | 190 | 190 |  |  | Off | 3 | On | CARE Dose4D | 13.14 | 30s | 25% | 50 |  | 0.5 |  | 2 | 0.6 |  | 5 | 32x1.2mm | [15 empty] | Off
|  | Non Contrast 5.0 B30f | [22 empty] | 5 | 5 | 41 | B30f medium smooth | Abdomen | [13 empty] | manual
| Arterial Phase |  | 120 | 120 | 210 | 210 |  |  | Off | 3 | On | CARE Dose4D | 14.19 | 30s | 25% | 50 |  | 0.5 |  | 2 | 0.6 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | Arterial Phase 5.0 B30f | [22 empty] | 5 | 5 | 41 | B30f medium smooth | Abdomen | [13 empty] | manual
|  | Arterial Phase 1.0 B20f | [22 empty] | 1 | 0.7 | 286 | B20f smooth | Abdomen | [13 empty] | manual
| Venous Phase |  | 120 | 120 | 190 | 190 |  |  | Off | 3 | On | CARE Dose4D | 13.14 | 30s | 25% | 50 |  | 0.5 |  | 25 | 0.6 |  | 5 | 32x1.2mm | [15 empty] | Off
|  | Venous Phase 5.0 B30f | [22 empty] | 5 | 5 | 41 | B30f medium smooth | Abdomen | [13 empty] | manual
AbdomenIntervention (Adult)
AbdomenIntervention (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body |  |  |  | GEPACS | ROEMIR001a | TRACKER
| i-Spiral |  | 120 | 120 | 210 | 210 |  |  | On | 3 | On | CARE Dose4D | 14.5 | 20s | 25% |  |  | 0.5 |  | 2 | 0.8 |  | 5 | 32x1.2mm | [15 empty] | Off
|  | i-Spiral 5.0 B30f | [22 empty] | 5 | 3.5 | 86 | B30f medium smooth | Abdomen | [13 empty] | manual
AbdomenIntervention_Customized (Adult)
AbdomenIntervention_Customized (Adult) | Topogram | Topogram 0.6 T20s |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body |  |  |  | GEPACS | ROEMIR001a | TRACKER
| i-Spiral |  | 120 | 120 | 210 | 210 |  |  | Off | 3 | On | CARE Dose4D | 14.5 | 30s | 25% |  |  | 0.5 |  | 2 | 0.8 |  | 5 | 32x1.2mm | [15 empty] | Off
|  | i-Spiral 5.0 B30f | [22 empty] | 5 | 3.5 | 86 | B30f medium smooth | Abdomen | [13 empty] | manual
AbdomenRoutine (Adult)
AbdomenRoutine (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body |  |  |  | GEPACS | ROEMIR001a | TRACKER
| Abdomen |  | 120 | 120 | 210 | 210 |  |  | On | 7 | On | CARE Dose4D | 14.19 | 20s | 25% |  |  | 0.5 |  | 2 | 0.6 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | Abdomen 5.0 B30f | [22 empty] | 5 | 5 | 41 | B30f medium smooth | Abdomen | [13 empty] | manual
|  | Abdomen 1.0 B20f | [22 empty] | 1 | 0.7 | 286 | B20f smooth | Abdomen | [13 empty] | manual
AbdomenRoutine_Customized_Customized (Adult)
AbdomenRoutine_Customized_Customized (Adult) | Topogram | Topogram 0.6 T20s |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body |  |  |  | GEPACS | ROEMIR001a | TRACKER
| Abdomen |  | 120 | 120 | 210 | 210 |  |  | Off | 3 | On | CARE Dose4D | 14.19 | 30s | 25% | 50 |  | 0.5 |  | 2 | 0.6 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | Abdomen 5.0 B30f | [22 empty] | 5 | 5 | 41 | B30f medium smooth | Abdomen | [13 empty] | manual
|  | Abdomen 1.0 B20f | [22 empty] | 1 | 0.7 | 286 | B20f smooth | Abdomen | [13 empty] | manual
AbdomenSeq (Adult)
AbdomenSeq (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body |  |  |  | GEPACS | ROEMIR001a | TRACKER
| AbdSeq |  | 120 | 120 | 210 | 210 |  |  | On | 7 | On | CARE Dose4D | 16.13 |  | 25% |  |  | 0.5 |  | 2 | 0 | 34 | 5 | 32x1.2mm | [15 empty] | Off
|  | AbdSeq 5.0 B30f | [22 empty] | 5 |  | 41 | B30f medium smooth | Abdomen | [13 empty] | manual
AbdomenSeq_Customized (Adult)
AbdomenSeq_Customized (Adult) | Topogram | Topogram 0.6 T20s |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body |  |  |  | GEPACS | ROEMIR001a | TRACKER
| AbdSeq |  | 120 | 120 | 230 | 230 |  |  | Off | 3 | On | CARE Dose4D | 17.67 |  | 25% | 50 |  | 0.5 |  | 2 | 0 | 34 | 5 | 32x1.2mm | [15 empty] | Off
|  | AbdSeq 5.0 B30f | [22 empty] | 5 |  | 41 | B30f medium smooth | Abdomen | [13 empty] | manual
Abdomen_IR (Adult)
Abdomen_IR (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Abdomen |  | 120 | 120 | 147 | 147 |  |  | On | 7 | On | CARE Dose4D | 9.97 | 30s | 25% |  |  | 0.5 |  | 2 | 0.6 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | Abdomen 5.0 I30f 3 | [22 empty] | 5 | 5 | 41 | I30f medium smooth | Abdomen | [13 empty] | manual
|  | Abdomen 1.0 I26f 3 | [22 empty] | 1 | 0.7 | 286 | I26f medium smooth ASA | Abdomen | [13 empty] | manual
Colonography (Adult)
Colonography (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body |  |  |  | GEPACS | ROEMIR001a | TRACKER
| Colo_supine |  | 120 | 120 | 55 | 55 |  |  | On | 1 | On | CARE Dose4D | 3.72 | 20s | 25% |  |  | 0.5 |  | 2 | 0.9 |  | 1 | 128x0.6mm | [15 empty] | Off
|  | Colo_supine 1.0 B20f | [22 empty] | 1 | 0.7 | 572 | B20f smooth | Abdomen | [13 empty] | manual
| Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body |  |  |  | GEPACS | ROEMIR001a | TRACKER
| Colo_prone |  | 120 | 120 | 30 | 30 |  |  | On | 1 | On | CARE Dose4D | 2.03 | 20s | 25% |  |  | 0.5 |  | 2 | 1.4 |  | 1 | 128x0.6mm | [15 empty] | Off
|  | Colo_prone 1.0 B10f | [22 empty] | 1 | 0.7 | 572 | B10f very smooth | Abdomen | [13 empty] | manual
Colonography_Customized (Adult)
Colonography_Customized (Adult) | Topogram | Topogram 0.6 T20s |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body |  |  |  | GEPACS | ROEMIR001a | TRACKER
| Colo_supine |  | 120 | 120 | 55 | 55 |  |  | Off | 3 | Off | CARE Dose4D | 3.72 | 30s | 25% | 50 |  | 0.5 |  | 2 | 0.9 |  | 1 | 128x0.6mm | [15 empty] | Off
|  | Colo_supine 1.0 B20f | [22 empty] | 1 | 0.7 | 572 | B20f smooth | Abdomen | [13 empty] | manual
| Topogram | Topogram 0.6 T20s |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body |  |  |  | GEPACS | ROEMIR001a | TRACKER
| Colo_prone |  | 120 | 120 | 30 | 30 |  |  | Off | 3 | Off | CARE Dose4D | 2.03 | 30s | 25% | 50 |  | 0.5 |  | 2 | 1.4 |  | 1 | 128x0.6mm | [15 empty] | Off
|  | Colo_prone 1.0 B10f | [22 empty] | 1 | 0.7 | 572 | B10f very smooth | Abdomen | [13 empty] | manual
DSXXL_Abdomen (Adult)
DSXXL_Abdomen (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body |  |  |  | GEPACS | ROEMIR001a | TRACKER
DSXXL_Abdomen_Customized (Adult)
DSXXL_Abdomen_Customized (Adult) | Topogram | Topogram 0.6 T20s |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body |  |  |  | GEPACS | ROEMIR001a | TRACKER
Renal_Stone_Customized (Adult)
Renal_Stone_Customized (Adult) | Topogram | Topogram 0.6 T20s |  | 100 |  | 318 |  |  |  | 3 | Off | CARE Dose | 0.15 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 67092aa9-b6e0-47d9-8c91-ba48e48b197a |  |  |  | ROEMIR001a
| Abdomen |  | 120 | 100 | 170 | 210 |  |  | Off | 3 | On | CARE Dose4D | 7 | 30s | 25% |  |  | 0.5 |  | 2 | 0.6 |  | 5 | 128x0.6mm | [5 empty] | 67092aa9-b6e0-47d9-8c91-ba48e48b197a | [9 empty] | Off
|  | Abdomen 5.0 B40f | [22 empty] | 5 | 5 | 36 | B40f medium | 400/40 Abd |  |  |  | AWCIC | [9 empty] | manual
|  | Abdomen 3.0 MPR ax | [22 empty] | 3 | 2.5 | 81 | B40f medium | 400/40 Abd |  |  |  | AWCIC | [9 empty] | manual
Spine
3010_ang_CspTspLsp (Adult)
3010_ang_CspTspLsp (Adult) | Topogram | Topogram |  | 100 |  | 393 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 393 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| art CspTspLsp |  | 120 | 140 | 260 | 260 |  |  | Off | 3 | On | CARE Dose4D | 25.53 | 30s | 25% | 50 |  | 1 |  | 9 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | art CspTspLsp s010 | [22 empty] | 3 | 3 | 54 | B31s medium smooth + | Abdomen |  |  |  | GE_PACS | [9 empty] | manual
|  | art CspTspLsp sag | [22 empty] | 3 | 1.5 | 167 | B31s medium smooth + | 400/40 Spine Std |  |  |  | GE_PACS | [9 empty] | manual
|  | art CspTspLsp :-D | [22 empty] | 0.6 | 0.6 | 267 | B31s medium smooth + | 400/40 Spine Std |  |  |  | ROUTER | GEPACS | [8 empty] | manual
3011_ang_Tsp (Adult)
3011_ang_Tsp (Adult) | Topogram | Topogram |  | 100 |  | 393 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 393 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| PreMonitoring |  | 120 | 120 | 260 | 40 |  |  | Off | 3 | Off | CARE Dose4D | 2.47 |  | 25% |  |  | 1 |  | 2 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | PreMonitoring 10.0 B31s | [22 empty] | 10 |  | 1 | B31s medium smooth + | Abdomen | [13 empty] | manual
| Monitoring |  | 120 | 120 | 260 | 40 |  |  | Off | 3 | Off | CARE Dose4D | 73.97 |  | 25% |  |  | 1 |  | 10 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | Monitoring 10.0 B31s | [22 empty] | 10 |  | 30 | B31s medium smooth + | Abdomen | [13 empty] | manual
| art Tsp |  | 120 | 120 | 260 | 260 |  |  | Off | 3 | On | CARE Dose4D | 17.54 | 30s | 25% | 50 |  | 1 |  | 9 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | art Tsp s011 | [22 empty] | 3 | 3 | 54 | B31s medium smooth + | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | art Tsp sag | [22 empty] | 2 | 2 | 126 | B31s medium smooth + | 400/40 Spine Std | [13 empty] | manual
|  | art Tsp cor | [22 empty] | 2 | 2 | 126 | B31s medium smooth + | 400/40 Spine Std |  |  |  | GEPACS | [9 empty] | manual
|  | art Tsp cor mip | [22 empty] | 10 | 10 | 26 | B31s medium smooth + | 400/40 Spine Std |  |  |  | GEPACS | [9 empty] | manual
|  | art Tsp s011 :-D | [22 empty] | 0.75 | 0.7 | 229 | B31s medium smooth + | Abdomen |  |  |  | GEPACS | ROUTER | MCHNeuroBodyAW | [7 empty] | manual
| ven Tsp |  | 120 | 120 | 260 | 260 |  |  | Off | 3 | On | CARE Dose4D | 17.57 | 30s | 25% | 50 |  | 1 |  | 50 | 0.9 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven Tsp bon s080 | [22 empty] | 3 | 2 | 81 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | ven Tsp cor bon | [22 empty] | 2 | 2 | 76 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | ven Tsp sag bon | [22 empty] | 2 | 2 | 76 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | ven Tsp | [22 empty] | 3 | 3 | 54 | B31s medium smooth + | Abdomen | [13 empty] | manual
|  | ven Tsp sag | [22 empty] | 2 | 2 | 76 | B31s medium smooth + | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | ven Tsp bon s080 :-D | [22 empty] | 0.6 | 0.6 | 267 | B70s very sharp | Bone |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
| 2m Tsp |  | 120 | 120 | 260 | 260 |  |  | Off | 3 | On | CARE Dose4D | 17.57 | 30s | 25% | 50 |  | 1 |  | 6 | 0.9 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | 2m Tsp bon s080 | [22 empty] | 3 | 2 | 81 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | 2m Tsp cor bon | [22 empty] | 2 | 2 | 76 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | 2m Tsp sag bon | [22 empty] | 2 | 2 | 76 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | 2m Tsp | [22 empty] | 3 | 2 | 81 | B31s medium smooth + | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | 2m Tsp sag | [22 empty] | 2 | 2 | 76 | B31s medium smooth + | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | 2m Tsp bon s080 :-D | [22 empty] | 0.6 | 0.6 | 267 | B70s very sharp | Bone |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
3015_CspTspLsp_mye (Adult)
3015_CspTspLsp_mye (Adult) | Topogram | Topogram 0.6 T80f |  | 140 |  | 393 |  |  |  | 3 | Off | CARE Dose | 0.29 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram 0.6 T80f |  | 140 |  | 393 |  |  |  | 3 | Off | CARE Dose | 0.29 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Needle Place |  | 120 | 140 | 200 | 200 |  |  | Off | 5 | On | CARE Dose4D | 19.66 | 20s | 25% |  |  | 1 |  | 2 | 1.1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | Lsp s015 | [22 empty] | 3 | 3 | 72 | B31s medium smooth + | Spine |  |  |  | GEPACS | [9 empty] | manual
|  | Needle Place 0.6 B31s | [22 empty] | 0.6 | 0.6 | 358 | B31s medium smooth + | Spine | [13 empty] | manual
| Spine 1 ca-cr |  | 120 | 140 | 200 | 200 |  |  | Off | 5 | On | CARE Dose4D | 19.66 | 20s | 25% | 50 |  | 1 |  | 2 | 1.1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | 1st cst CspTspLsp s015 | [22 empty] | 3 | 3 | 212 | B31s medium smooth + | Spine | [13 empty] | manual
|  | 1st cst CspTspLsp s015 :-D | [22 empty] | 0.6 | 0.6 | 1056 | B31s medium smooth + | Spine |  |  |  | ROUTER | GEPACS | [8 empty] | manual
|  | 1st cst CspTspLsp cor | [22 empty] | 2 | 2 | 151 | B31s medium smooth + | Spine | [13 empty] | manual
|  | 1st cst CspTspLsp sag | [22 empty] | 2 | 2 | 151 | B31s medium smooth + | Spine | [13 empty] | manual
| Spine 2 cr-ca |  | 120 | 140 | 200 | 200 |  |  | Off | 5 | On | CARE Dose4D | 19.66 | 20s | 25% | 50 |  | 1 |  | 2 | 1.1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | 2nd cst CspTspLsp s015 | [22 empty] | 3 | 3 | 212 | B31s medium smooth + | Spine | [13 empty] | manual
|  | 2nd cst CspTspLsp s015 :-D | [22 empty] | 0.6 | 0.6 | 1056 | B31s medium smooth + | Spine |  |  |  | ROUTER | GEPACS | [8 empty] | manual
|  | 2nd cst CspTspLsp cor | [22 empty] | 2 | 2 | 151 | B31s medium smooth + | Spine | [13 empty] | manual
|  | 2nd cst CspTspLsp sag | [22 empty] | 2 | 2 | 151 | B31s medium smooth + | Spine | [13 empty] | manual
| Spine 3 ca-cr |  | 120 | 140 | 200 | 200 |  |  | Off | 5 | On | CARE Dose4D | 19.66 | 20s | 25% | 50 |  | 1 |  | 2 | 1.1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | 3rd cst CspTspLsp s015 | [22 empty] | 3 | 3 | 212 | B31s medium smooth + | Spine | [13 empty] | manual
|  | 3rd cst CspTspLsp s015 :-D | [22 empty] | 0.6 | 0.6 | 1056 | B31s medium smooth + | Spine |  |  |  | ROUTER | GEPACS | [8 empty] | manual
|  | 3rd cst CspTspLsp cor | [22 empty] | 2 | 2 | 151 | B31s medium smooth + | Spine | [13 empty] | manual
|  | 3rd cst CspTspLsp sag | [22 empty] | 2 | 2 | 151 | B31s medium smooth + | Spine | [13 empty] | manual
| Spine 4 cr-ca |  | 120 | 140 | 200 | 200 |  |  | Off | 5 | On | CARE Dose4D | 19.66 | 20s | 25% | 50 |  | 1 |  | 2 | 1.1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | 4th cst CspTspLsp s015 | [22 empty] | 3 | 3 | 212 | B31s medium smooth + | Spine | [13 empty] | manual
|  | 4th cst CspTspLsp s015 :-D | [22 empty] | 0.6 | 0.6 | 1056 | B31s medium smooth + | Spine |  |  |  | ROUTER | GEPACS | [8 empty] | manual
|  | 4th cst CspTspLsp cor | [22 empty] | 2 | 2 | 151 | B31s medium smooth + | Spine | [13 empty] | manual
|  | 4th cst CspTspLsp sag | [22 empty] | 2 | 2 | 151 | B31s medium smooth + | Spine | [13 empty] | manual
| Spine 5 ca-cr |  | 120 | 140 | 200 | 200 |  |  | Off | 5 | On | CARE Dose4D | 19.66 | 20s | 25% | 50 |  | 1 |  | 2 | 1.1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | 5th cst CspTspLsp s015 | [22 empty] | 3 | 3 | 212 | B31s medium smooth + | Spine | [13 empty] | manual
|  | 5th cst CspTspLsp s015 :-D | [22 empty] | 0.6 | 0.6 | 1056 | B31s medium smooth + | Spine |  |  |  | ROUTER | GEPACS | [8 empty] | manual
|  | 5th cst CspTspLsp cor | [22 empty] | 0.6 | 0.6 | 501 | B31s medium smooth + | Spine | [13 empty] | manual
|  | 5th cst CspTspLsp sag | [22 empty] | 0.6 | 0.6 | 501 | B31s medium smooth + | Spine | [13 empty] | manual
| Spine 6 cr-ca |  | 120 | 140 | 200 | 200 |  |  | Off | 5 | On | CARE Dose4D | 19.66 | 20s | 25% | 50 |  | 1 |  | 2 | 1.1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | 6th cst CspTspLsp s015 | [22 empty] | 3 | 3 | 212 | B31s medium smooth + | Spine | [13 empty] | manual
|  | 6th cst CspTspLsp s015 :-D | [22 empty] | 0.6 | 0.6 | 1056 | B31s medium smooth + | Spine |  |  |  | ROUTER | GEPACS | [8 empty] | manual
|  | 6th cst CspTspLsp cor | [22 empty] | 2 | 2 | 151 | B31s medium smooth + | Spine | [13 empty] | manual
|  | 6th cst CspTspLsp sag | [22 empty] | 2 | 2 | 151 | B31s medium smooth + | Spine | [13 empty] | manual
| Post Myelo |  | 120 | 140 | 200 | 200 |  |  | Off | 5 | On | CARE Dose4D | 19.66 | 20s | 25% | 50 |  | 1 |  | 2 | 1.1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | Lsp pos cst s015 | [22 empty] | 3 | 3 | 72 | B31s medium smooth + | Spine | [13 empty] | manual
|  | Post Myelo 0.6 B31s | [22 empty] | 0.6 | 0.6 | 358 | B31s medium smooth + | Spine | [13 empty] | manual
3020_Csp (Adult)
3020_Csp (Adult) | Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| ven Csp |  | 120 | 140 | 280 | 280 |  |  | Off | 3 | On | CARE Dose4D | 27.54 | 30s | 25% | 50 |  | 1 |  | 70 | 0.9 |  | 2 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven Csp bon s020 | [22 empty] | 2 | 2 | 81 | B70s very sharp | Bone |  |  |  | GE_PACS | [9 empty] | manual
|  | ven Csp cor bon | [22 empty] | 2 | 2 | 76 | B70s very sharp | Bone | [13 empty] | manual
|  | ven Csp sag bon | [22 empty] | 2 | 2 | 76 | B70s very sharp | Bone |  |  |  | GE_PACS | [9 empty] | manual
|  | ven Csp | [22 empty] | 2 | 2 | 81 | B30s medium smooth | Abdomen |  |  |  | GE_PACS | [9 empty] | manual
|  | ven Csp sag | [22 empty] | 2 | 2 | 76 | B30s medium smooth | Abdomen |  |  |  | GE_PACS | [9 empty] | manual
|  | ven Csp bon s020 :-D | [22 empty] | 0.6 | 0.6 | 267 | B70s very sharp | Bone |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
3030_Csp_ste_wo_cst (Adult)
3030_Csp_ste_wo_cst (Adult) | Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Csp |  | 120 | 140 | 280 | 280 |  |  | Off | 3 | On | CARE Dose4D | 27.54 | 30s | 25% | 50 |  | 1 |  | 3 | 0.9 |  | 2 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | Csp bon s030 | [22 empty] | 2 | 2 | 81 | B70s very sharp | Bone |  |  |  | GE_PACS | [9 empty] | manual
|  | Csp cor bon | [22 empty] | 2 | 2 | 76 | B70s very sharp | Bone | [13 empty] | manual
|  | Csp sag bon | [22 empty] | 2 | 2 | 76 | B70s very sharp | Bone |  |  |  | GE_PACS | [9 empty] | manual
|  | Csp | [22 empty] | 2 | 2 | 81 | B30s medium smooth | Abdomen |  |  |  | GE_PACS | [9 empty] | manual
|  | Csp sag | [22 empty] | 2 | 2 | 76 | B30s medium smooth | Abdomen |  |  |  | GE_PACS | [9 empty] | manual
|  | Csp bon s030 :-D | [22 empty] | 0.6 | 0.6 | 267 | B70s very sharp | Bone |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
3040_Csp_wo_cst (Adult)
3040_Csp_wo_cst (Adult) | Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Csp |  | 120 | 140 | 280 | 280 |  |  | Off | 3 | On | CARE Dose4D | 27.54 | 30s | 25% | 50 |  | 1 |  | 3 | 0.9 |  | 2 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | Csp bon s040 | [22 empty] | 2 | 2 | 81 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Csp cor bon | [22 empty] | 2 | 2 | 76 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Csp sag bon | [22 empty] | 2 | 2 | 76 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Csp | [22 empty] | 2 | 2 | 81 | B30s medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | Csp sag | [22 empty] | 2 | 2 | 76 | B30s medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | Csp bon s040 :-D | [22 empty] | 0.6 | 0.6 | 267 | B70s very sharp | Bone |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
3050_Lsp (Adult)
3050_Lsp (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| ven Lsp |  | 120 | 120 | 260 | 260 |  |  | Off | 3 | On | CARE Dose4D | 17.57 | 30s | 25% | 50 |  | 1 |  | 70 | 0.9 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven Lsp bon s050 | [22 empty] | 3 | 2 | 81 | B70s very sharp | Bone |  |  |  | GE_PACS | [9 empty] | manual
|  | ven Lsp cor bon | [22 empty] | 2 | 2 | 76 | B70s very sharp | Bone |  |  |  | GE_PACS | [9 empty] | manual
|  | ven Lsp sag bon | [22 empty] | 2 | 2 | 76 | B70s very sharp | Bone |  |  |  | GE_PACS | [9 empty] | manual
|  | ven Lsp | [22 empty] | 3 | 2 | 81 | B31s medium smooth + | Abdomen |  |  |  | GE_PACS | [9 empty] | manual
|  | ven Lsp sag | [22 empty] | 2 | 2 | 91 | B31s medium smooth + | Abdomen |  |  |  | GE_PACS | [9 empty] | manual
|  | ven Lsp bon s050 :-D | [22 empty] | 0.6 | 0.6 | 267 | B70s very sharp | Bone |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
3060_Lsp_ste_wo_cst (Adult)
3060_Lsp_ste_wo_cst (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Lsp |  | 120 | 120 | 260 | 260 |  |  | Off | 3 | On | CARE Dose4D | 17.57 | 30s | 25% | 50 |  | 1 |  | 2 | 0.9 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | Lsp bon s060 | [22 empty] | 3 | 2 | 81 | B70s very sharp | Bone |  |  |  | GE_PACS | [9 empty] | manual
|  | Lsp cor bon | [22 empty] | 2 | 2 | 76 | B70s very sharp | Bone |  |  |  | GE_PACS | [9 empty] | manual
|  | Lsp sag bon | [22 empty] | 2 | 2 | 76 | B70s very sharp | Bone |  |  |  | GE_PACS | [9 empty] | manual
|  | Lsp | [22 empty] | 3 | 2 | 81 | B31s medium smooth + | Abdomen |  |  |  | GE_PACS | [9 empty] | manual
|  | Lsp sag | [22 empty] | 2 | 2 | 91 | B31s medium smooth + | Abdomen |  |  |  | GE_PACS | [9 empty] | manual
|  | Lsp bon s060 :-D | [22 empty] | 0.6 | 0.6 | 267 | B70s very sharp | Bone |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
3070_Lsp_wo_cst (Adult)
3070_Lsp_wo_cst (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Lsp |  | 120 | 120 | 260 | 260 |  |  | Off | 3 | On | CARE Dose4D | 17.57 | 30s | 25% | 50 |  | 1 |  | 2 | 0.9 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | Lsp bon s070 | [22 empty] | 3 | 2 | 81 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Lsp cor bon | [22 empty] | 2 | 2 | 76 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Lsp sag bon | [22 empty] | 2 | 2 | 76 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Lsp | [22 empty] | 3 | 2 | 81 | B31s medium smooth + | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | Lsp sag | [22 empty] | 2 | 2 | 91 | B31s medium smooth + | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | Lsp bon s070 :-D | [22 empty] | 0.6 | 0.6 | 267 | B70s very sharp | Bone |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
3080_Tsp (Adult)
3080_Tsp (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| ven Tsp |  | 120 | 120 | 260 | 260 |  |  | Off | 3 | On | CARE Dose4D | 17.57 | 30s | 25% | 50 |  | 1 |  | 70 | 0.9 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven Tsp bon s080 | [22 empty] | 3 | 2 | 81 | B70s very sharp | Bone |  |  |  | GE_PACS | [9 empty] | manual
|  | ven Tsp cor bon | [22 empty] | 2 | 2 | 76 | B70s very sharp | Bone |  |  |  | GE_PACS | [9 empty] | manual
|  | ven Tsp sag bon | [22 empty] | 2 | 2 | 76 | B70s very sharp | Bone |  |  |  | GE_PACS | [9 empty] | manual
|  | ven Tsp | [22 empty] | 3 | 2 | 81 | B31s medium smooth + | Abdomen |  |  |  | GE_PACS | [9 empty] | manual
|  | ven Tsp sag | [22 empty] | 2 | 2 | 76 | B31s medium smooth + | Abdomen |  |  |  | GE_PACS | [9 empty] | manual
|  | ven Tsp bon s080 :-D | [22 empty] | 0.6 | 0.6 | 251 | B70s very sharp | Bone |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
3090_Tsp_ste_wo_cst (Adult)
3090_Tsp_ste_wo_cst (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Tsp |  | 120 | 120 | 260 | 260 |  |  | Off | 3 | On | CARE Dose4D | 17.57 | 30s | 25% | 50 |  | 1 |  | 3 | 0.9 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | Tsp bon s090 | [22 empty] | 3 | 2 | 81 | B70s very sharp | Bone |  |  |  | GE_PACS | [9 empty] | manual
|  | Tsp cor bon | [22 empty] | 2 | 2 | 76 | B70s very sharp | Bone |  |  |  | GE_PACS | [9 empty] | manual
|  | Tsp sag bon | [22 empty] | 2 | 2 | 76 | B70s very sharp | Bone |  |  |  | GE_PACS | [9 empty] | manual
|  | Tsp | [22 empty] | 3 | 2 | 81 | B31s medium smooth + | Abdomen |  |  |  | GE_PACS | [9 empty] | manual
|  | Tsp sag | [22 empty] | 2 | 2 | 76 | B31s medium smooth + | Abdomen |  |  |  | GE_PACS | [9 empty] | manual
|  | Tsp bon s090 :-D | [22 empty] | 0.6 | 0.6 | 251 | B70s very sharp | Bone |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
3100_Tsp_wo_cst (Adult)
3100_Tsp_wo_cst (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Tsp |  | 120 | 120 | 260 | 260 |  |  | Off | 3 | On | CARE Dose4D | 17.57 | 30s | 25% | 50 |  | 1 |  | 3 | 0.9 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | Tsp bon s100 | [22 empty] | 3 | 2 | 81 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Tsp cor bon | [22 empty] | 2 | 2 | 76 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Tsp sag bon | [22 empty] | 2 | 2 | 76 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Tsp | [22 empty] | 3 | 2 | 81 | B31s medium smooth + | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | Tsp sag | [22 empty] | 2 | 2 | 76 | B31s medium smooth + | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | Tsp bon s100 :-D | [22 empty] | 0.6 | 0.6 | 251 | B70s very sharp | Bone |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
3110_LD_ske_sur_wo_cst (Adult)
3110_LD_ske_sur_wo_cst (Adult) | Topogram | Topogram |  | 100 |  | 777 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | TRACKER
| Topogram | Topogram |  | 100 |  | 777 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | TRACKER
| Hea thru Kne |  | 120 | 120 | 70 | 210 |  |  | Off | 6 | On | CARE Dose4D | 4.72 | 30s | 25% | 50 |  | 0.5 |  | 7 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | Ske Sur bon s110 | [22 empty] | 3 | 3 | 622 | B70f very sharp | Bone |  |  |  | GEPACS | MCB3DLabAW | MCH3DLabAW | [7 empty] | manual
|  | Ske Sur s110 | [22 empty] | 2 | 1 | 1864 | B40f medium | Abdomen |  |  |  | GEPACS | MCB3DLabAW | MCH3DLabAW | [7 empty] | manual
|  | Ske Sur bon s110 :-D | [22 empty] | 1 | 1 | 1864 | B70f very sharp | Bone |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
C_Spine (Adult)
C_Spine (Adult) | Topogram | Topogram 0.6 T80f |  | 120 |  | 164 |  |  |  | 3 | Off | CARE Dose | 0.24 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Body |  |  |  | GEPACS | ROEMIR001a | TRACKER
| C_Spine |  | 120 | 120 | 275 | 275 |  |  | On | 3 | On | CARE Dose4D | 18.57 | 30s | 25% |  |  | 1 |  | 2 | 0.8 |  | 2 | 128x0.6mm | [15 empty] | Off
|  | C_Spine 2.0 B30s | [22 empty] | 2 | 2 | 81 | B30s medium smooth | Spine | [13 empty] | manual
|  | C_Spine 0.6 B20s | [22 empty] | 0.6 | 0.4 | 401 | B20s smooth | Spine | [13 empty] | manual
|  | C_Spine 0.6 B60s | [22 empty] | 0.6 | 0.4 | 401 | B60s sharp | Spine | [13 empty] | manual
C_Spine_IR (Adult)
C_Spine_IR (Adult) | Topogram | Topogram 0.6 T80f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Body
| C_Spine |  | 120 | 120 | 193 | 193 |  |  | On | 3 | On | CARE Dose4D | 13.08 | 30s | 25% |  |  | 1 |  | 2 | 0.8 |  | 2 | 128x0.6mm | [15 empty] | Off
|  | C_Spine 2.0 I30s 3 | [22 empty] | 2 | 2 | 81 | I30s medium smooth | Spine | [13 empty] | manual
|  | C_Spine 0.6 I30s 3 | [22 empty] | 0.6 | 0.4 | 401 | I30s medium smooth | Spine | [13 empty] | manual
|  | C_Spine 0.6 I50s 3 | [22 empty] | 0.6 | 0.4 | 401 | I50s medium sharp ASA | Spine | [13 empty] | manual
C_Spine_IR_Vol (Adult)
C_Spine_IR_Vol (Adult) | Topogram | Topogram 0.6 T80f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Body
| C_Spine |  | 120 | 120 | 193 | 193 |  |  | On | 3 | On | CARE Dose4D | 13.08 | 30s | 25% |  |  | 1 |  | 2 | 0.8 |  | 2 | 128x0.6mm | [15 empty] | Off
|  | C_Spine 2.0 I30s 3 | [22 empty] | 2 | 2 | 81 | I30s medium smooth | Spine | [13 empty] | manual
|  | C_Spine 0.6 MPR spine multi | [22 empty] | 0.6 | 0.4 | 0 | I26s medium smooth ASA | Spine | [13 empty] | manual
|  | C_Spine 2.0 MPR spine | [22 empty] | 2 | 2 | 17 | I30s medium smooth | Spine | [13 empty] | manual
|  | C_Spine 2.0 MPR spine | [22 empty] | 2 | 2 | 17 | I50s medium sharp ASA | Spine | [13 empty] | manual
C_Spine_Vol (Adult)
C_Spine_Vol (Adult) | Topogram | Topogram 0.6 T80f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Body
| C_Spine |  | 120 | 120 | 275 | 275 |  |  | On | 3 | On | CARE Dose4D | 18.57 | 30s | 25% |  |  | 1 |  | 2 | 0.8 |  | 2 | 128x0.6mm | [15 empty] | Off
|  | C_Spine 2.0 B30s | [22 empty] | 2 | 2 | 81 | B30s medium smooth | Spine | [13 empty] | manual
|  | C_Spine 0.6 MPR spine multi | [22 empty] | 0.6 | 0.4 | 0 | B20s smooth | Spine | [13 empty] | manual
|  | C_Spine 2.0 MPR spine | [22 empty] | 2 | 2 | 17 | B20s smooth | Spine | [13 empty] | manual
|  | C_Spine 2.0 MPR spine | [22 empty] | 2 | 2 | 17 | B60s sharp | Spine | [13 empty] | manual
DSXXL_Spine (Adult)
DSXXL_Spine (Adult) | Topogram | Topogram 0.6 T80f |  | 120 |  | 318 |  |  |  | 3 | Off | CARE Dose | 0.24 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Body |  |  |  | GEPACS | ROEMIR001a | TRACKER
DSXXL_Spine_Vol (Adult)
DSXXL_Spine_Vol (Adult) | Topogram | Topogram 0.6 T80f |  | 140 |  | 529 |  |  |  | 3 | Off | CARE Dose | 0.58 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Body
SpineIntervention (Adult)
SpineIntervention (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body |  |  |  | GEPACS | ROEMIR001a | TRACKER
| i-Spiral |  | 120 | 120 | 275 | 275 |  |  | On | 3 | On | CARE Dose4D | 18.99 | 30s | 25% |  |  | 0.5 |  | 2 | 0.8 |  | 5 | 32x1.2mm | [15 empty] | Off
|  | i-Spiral 5.0 B30f | [22 empty] | 5 | 3.5 | 46 | B30f medium smooth | Spine | [13 empty] | manual
SpineNeuro (Adult)
SpineNeuro (Adult) | Topogram | Topogram 0.6 T80f |  | 120 |  | 318 |  |  |  | 3 | Off | CARE Dose | 0.24 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Body |  |  |  | GEPACS | ROEMIR001a | TRACKER
| Spine |  | 120 | 120 | 330 | 330 |  |  | On | 3 | On | CARE Dose4D | 26.85 | 40s | 25% |  |  | 1 |  | 2 | 0.8 |  | 3 | 20x0.6mm | [15 empty] | Off
|  | Spine 3.0 B30s | [22 empty] | 3 | 3 | 54 | B30s medium smooth | Spine | [13 empty] | manual
|  | Spine 0.6 B20s | [22 empty] | 0.6 | 0.4 | 401 | B20s smooth | Spine | [13 empty] | manual
|  | Spine 0.6 B60s | [22 empty] | 0.6 | 0.4 | 401 | B60s sharp | Spine | [13 empty] | manual
SpineNeuro_Vol (Adult)
SpineNeuro_Vol (Adult) | Topogram | Topogram 0.6 T80f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Body
| Spine |  | 120 | 120 | 330 | 330 |  |  | On | 3 | On | CARE Dose4D | 26.85 | 40s | 25% |  |  | 1 |  | 2 | 0.8 |  | 3 | 20x0.6mm | [15 empty] | Off
|  | Spine 3.0 B30s | [22 empty] | 3 | 3 | 54 | B30s medium smooth | Spine | [13 empty] | manual
|  | Spine 0.6 MPR spine multi | [22 empty] | 0.6 | 0.4 | 0 | B20s smooth | Spine | [13 empty] | manual
|  | Spine 2.0 MPR spine | [22 empty] | 2 | 2 | 17 | B20s smooth | Spine | [13 empty] | manual
|  | Spine 2.0 MPR spine | [22 empty] | 2 | 2 | 17 | B60s sharp | Spine | [13 empty] | manual
SpineRoutine (Adult)
SpineRoutine (Adult) | Topogram | Topogram 0.6 T80f |  | 120 |  | 318 |  |  |  | 3 | Off | CARE Dose | 0.24 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Body |  |  |  | GEPACS | ROEMIR001a | TRACKER
| Spine |  | 120 | 120 | 330 | 330 |  |  | On | 3 | On | CARE Dose4D | 22.28 | 30s | 25% |  |  | 1 |  | 2 | 0.8 |  | 3 | 128x0.6mm | [15 empty] | Off
|  | Spine 3.0 B30s | [22 empty] | 3 | 3 | 54 | B30s medium smooth | Spine | [13 empty] | manual
|  | Spine 0.6 B20s | [22 empty] | 0.6 | 0.4 | 401 | B20s smooth | Spine | [13 empty] | manual
|  | Spine 0.6 B60s | [22 empty] | 0.6 | 0.4 | 401 | B60s sharp | Spine | [13 empty] | manual
SpineRoutine_Vol (Adult)
SpineRoutine_Vol (Adult) | Topogram | Topogram 0.6 T80f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Body
| Spine |  | 120 | 120 | 330 | 330 |  |  | On | 3 | On | CARE Dose4D | 22.28 | 40s | 25% |  |  | 1 |  | 2 | 0.8 |  | 3 | 128x0.6mm | [15 empty] | Off
|  | Spine 3.0 B30s | [22 empty] | 3 | 3 | 54 | B30s medium smooth | Spine | [13 empty] | manual
|  | Spine 0.6 MPR spine multi | [22 empty] | 0.6 | 0.4 | 0 | B20s smooth | Spine | [13 empty] | manual
|  | Spine 2.0 MPR spine | [22 empty] | 2 | 2 | 17 | B20s smooth | Spine | [13 empty] | manual
|  | Spine 2.0 MPR spine | [22 empty] | 2 | 2 | 17 | B60s sharp | Spine | [13 empty] | manual
SpineSeq (Adult)
SpineSeq (Adult) | Topogram | Topogram 0.6 T80f |  | 120 |  | 318 |  |  |  | 3 | Off | CARE Dose | 0.24 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Body |  |  |  | GEPACS | ROEMIR001a | TRACKER
| L3-4 |  | 120 | 120 | 220 | 220 |  |  | On | 3 | On | CARE Dose4D | 17.4 |  | 25% |  |  | 1 |  | 2 | 0 | 14 | 2.4 | 12x1.2mm | [15 empty] | Off
|  | L3-4 2.4 B30s | [22 empty] | 2.4 |  | 6 | B30s medium smooth | Spine | [13 empty] | manual
| L4-5 |  | 120 | 120 | 220 | 220 |  |  | On | 3 | On | CARE Dose4D | 17.4 |  | 25% |  |  | 1 |  | 2 | 0 | 14 | 2.4 | 12x1.2mm | [15 empty] | Off
|  | L4-5 2.4 B30s | [22 empty] | 2.4 |  | 6 | B30s medium smooth | Spine | [13 empty] | manual
| L5-S1 |  | 120 | 120 | 220 | 220 |  |  | On | 3 | On | CARE Dose4D | 17.4 |  | 25% |  |  | 1 |  | 2 | 0 | 14 | 2.4 | 12x1.2mm | [15 empty] | Off
|  | L5-S1 2.4 B30s | [22 empty] | 2.4 |  | 6 | B30s medium smooth | Spine | [13 empty] | manual
Spine_IR (Adult)
Spine_IR (Adult) | Topogram | Topogram 0.6 T80f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Body
| Spine |  | 120 | 120 | 231 | 231 |  |  | On | 3 | On | CARE Dose4D | 15.61 | 40s | 25% |  |  | 1 |  | 2 | 0.8 |  | 3 | 128x0.6mm | [15 empty] | Off
|  | Spine 3.0 I30s 3 | [22 empty] | 3 | 3 | 54 | I30s medium smooth | Spine | [13 empty] | manual
|  | Spine 0.6 I30s 3 | [22 empty] | 0.6 | 0.4 | 401 | I30s medium smooth | Spine | [13 empty] | manual
|  | Spine 0.6 I50s 3 | [22 empty] | 0.6 | 0.4 | 401 | I50s medium sharp ASA | Spine | [13 empty] | manual
Spine_IR_Vol (Adult)
Spine_IR_Vol (Adult) | Topogram | Topogram 0.6 T80f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Body
| Spine |  | 120 | 120 | 231 | 231 |  |  | On | 3 | On | CARE Dose4D | 15.61 | 40s | 25% |  |  | 1 |  | 2 | 0.8 |  | 3 | 128x0.6mm | [15 empty] | Off
|  | Spine 3.0 I30s 3 | [22 empty] | 3 | 3 | 54 | I30s medium smooth | Spine | [13 empty] | manual
|  | Spine 0.6 MPR spine multi | [22 empty] | 0.6 | 0.4 | 0 | I26s medium smooth ASA | Spine | [13 empty] | manual
|  | Spine 2.0 MPR spine | [22 empty] | 2 | 2 | 17 | I30s medium smooth | Spine | [13 empty] | manual
|  | Spine 2.0 MPR spine | [22 empty] | 2 | 2 | 17 | I50s medium sharp ASA | Spine | [13 empty] | manual
Pelvis
8000_ang_Pel (Adult)
8000_ang_Pel (Adult) | Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| PreMonitoring |  | 120 | 120 | 350 | 50 |  |  | Off | 3 | Off | CARE Dose4D | 3.08 |  | 25% |  |  | 0.5 |  | 2 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | PreMonitoring | [22 empty] | 10 |  | 1 | B40s medium | 400/40 Abd | [13 empty] | manual
| Monitoring |  | 120 | 120 | 350 | 50 |  |  | Off | 3 | Off | CARE Dose4D | 123.28 |  | 25% |  |  | 0.5 |  | 10 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | Monitoring | [22 empty] | 10 |  | 40 | B40s medium | 400/40 Abd | [13 empty] | manual
| art Pel |  | 120 | 120 | 240 | 210 |  |  | On | 8 | On | CARE Dose4D | 16.2 | 30s | 25% | 50 |  | 0.5 |  | 18 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | art Pel p000 | [22 empty] | 3 | 3 | 80 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art Pel cor | [22 empty] | 3 | 3 | 127 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art Pel cor mip | [22 empty] | 24 | 5 | 73 | I30f medium smooth | Abdomen MIP |  |  |  | GEPACS | [9 empty] | manual
|  | art Pel sag | [22 empty] | 3 | 3 | 121 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | art Pel p000 :-D | [22 empty] | 0.6 | 0.6 | 400 | I30f medium smooth | Abdomen |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
| ven Pel |  | 120 | 120 | 240 | 210 |  |  | On | 6 | On | CARE Dose4D | 16.2 | 30s | 25% | 50 |  | 0.5 |  | 60 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven Pel p000 | [22 empty] | 3 | 3 | 80 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven Pel cor | [22 empty] | 3 | 3 | 127 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven Pel sag | [22 empty] | 3 | 3 | 121 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven Pel p000 :-D | [22 empty] | 0.6 | 0.6 | 400 | I30f medium smooth | Abdomen |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
8010_Pel (Adult)
8010_Pel (Adult) | Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T21f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T21f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| ven Pel |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.17 | 30s | 25% | 50 |  | 0.5 |  | 90 | 1.2 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven Pel p010 | [22 empty] | 3 | 3 | 67 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven Pel cor | [22 empty] | 3 | 3 | 121 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven Pel sag | [22 empty] | 3 | 3 | 121 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven Pel p010 :-D | [22 empty] | 0.6 | 0.6 | 334 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | [8 empty] | manual
8020_Pel_cys_dir (Adult)
8020_Pel_cys_dir (Adult) | Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T21f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T21f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Pel |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.17 | 30s | 25% | 50 |  | 0.5 |  | 2 | 1.2 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | Pel p020 | [22 empty] | 3 | 3 | 67 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | Pel cor | [22 empty] | 3 | 3 | 121 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | Pel sag | [22 empty] | 3 | 3 | 121 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | Pel p020 :-D | [22 empty] | 0.6 | 0.6 | 334 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
| cst Pel |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.17 | 30s | 25% | 50 |  | 0.5 |  | 2 | 1.2 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | cst Pel p020 | [22 empty] | 3 | 3 | 67 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | cst Pel cor | [22 empty] | 3 | 3 | 134 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | cst Pel sag | [22 empty] | 3 | 3 | 134 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | cst Pel p020 :-D | [22 empty] | 0.6 | 0.6 | 334 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
8040_Pel_wo_cst (Adult)
8040_Pel_wo_cst (Adult) | Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T21f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T21f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Pel |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.17 | 30s | 25% | 50 |  | 0.5 |  | 2 | 1.2 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | Pel p040 | [22 empty] | 3 | 3 | 67 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | Pel cor | [22 empty] | 3 | 3 | 121 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | Pel sag | [22 empty] | 3 | 3 | 121 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | Pel p040 :-D | [22 empty] | 0.6 | 0.6 | 334 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
DSXXL_Pelvis (Adult)
DSXXL_Pelvis (Adult) | Topogram | Topogram 0.6 T80f |  | 120 |  | 318 |  |  |  | 3 | Off | CARE Dose | 0.24 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Body |  |  |  | GEPACS | ROEMIR001a | TRACKER
DSXXL_Pelvis_Customized (Adult)
DSXXL_Pelvis_Customized (Adult) | Topogram | Topogram 0.6 T80f |  | 120 |  | 318 |  |  |  | 3 | Off | CARE Dose | 0.24 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Body |  |  |  | GEPACS | ROEMIR001a | TRACKER
Hip (Adult)
Hip (Adult) | Topogram | Topogram 0.6 T80f |  | 120 |  | 164 |  |  |  | 3 | Off | CARE Dose | 0.24 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Body |  |  |  | GEPACS | ROEMIR001a | TRACKER
| Hip |  | 120 | 120 | 190 | 190 |  |  | On | 5 | On | CARE Dose4D | 12.83 | 20s | 25% |  |  | 1 |  | 2 | 0.8 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | Hip 5.0 B60s | [22 empty] | 5 | 5 | 21 | B60s sharp | Pelvis | [13 empty] | manual
|  | Hip 0.6 B60s | [22 empty] | 0.6 | 0.4 | 251 | B60s sharp | Pelvis | [13 empty] | manual
Hip_Customized (Adult)
Hip_Customized (Adult) | Topogram | Topogram 0.6 T80f |  | 120 |  | 164 |  |  |  | 3 | Off | CARE Dose | 0.24 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Body |  |  |  | GEPACS | ROEMIR001a | TRACKER
| Hip |  | 120 | 120 | 190 | 190 |  |  | On | 5 | On | CARE Dose4D | 12.83 | 20s | 25% | 50 |  | 1 |  | 2 | 0.8 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | Hip 5.0 B60s | [22 empty] | 5 | 5 | 21 | B60s sharp | Pelvis | [13 empty] | manual
|  | Hip 0.6 B60s | [22 empty] | 0.6 | 0.4 | 251 | B60s sharp | Pelvis | [13 empty] | manual
PelvisRoutine (Adult)
PelvisRoutine (Adult) | Topogram | Topogram 0.6 T80f |  | 120 |  | 318 |  |  |  | 3 | Off | CARE Dose | 0.24 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Body |  |  |  | GEPACS | ROEMIR001a | TRACKER
| Pelvis |  | 120 | 120 | 210 | 210 |  |  | On | 7 | On | CARE Dose4D | 14.5 | 20s | 25% |  |  | 0.5 |  | 2 | 0.8 |  | 5 | 32x1.2mm | [15 empty] | Off
|  | Pelvis 5.0 B31f | [22 empty] | 5 | 5 | 41 | B31f medium smooth + | Pelvis | [13 empty] | manual
PelvisRoutine_Customized (Adult)
PelvisRoutine_Customized (Adult) | Topogram | Topogram 0.6 T80f |  | 120 |  | 318 |  |  |  | 3 | Off | CARE Dose | 0.24 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Body |  |  |  | GEPACS | ROEMIR001a | TRACKER
| Pelvis |  | 120 | 120 | 210 | 210 |  |  | On | 7 | On | CARE Dose4D | 14.5 | 20s | 25% | 50 |  | 0.5 |  | 2 | 0.8 |  | 5 | 32x1.2mm | [15 empty] | Off
|  | Pelvis 5.0 B31f | [22 empty] | 5 | 5 | 41 | B31f medium smooth + | Pelvis | [13 empty] | manual
Pelvis_IR (Adult)
Pelvis_IR (Adult) | Topogram | Topogram 0.6 T80f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Body
| Pelvis |  | 120 | 120 | 147 | 147 |  |  | On | 7 | On | CARE Dose4D | 10.19 | 30s | 25% |  |  | 0.5 |  | 2 | 0.8 |  | 5 | 32x1.2mm | [15 empty] | Off
|  | Pelvis 5.0 I31f 3 | [22 empty] | 5 | 5 | 41 | I31f medium smooth | Pelvis | [13 empty] | manual
Specials
7275_Quantification_Phantom (Adult)
7275_Quantification_Phantom (Adult) | Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T81f sharp | 400/40 Abd |  |  |  | PHYSICSMIDIA
| Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T81f sharp | 400/40 Abd |  |  |  | PHYSICSMIDIA
| 120kVp |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 2 | 0.8 |  | 3 | 128x0.6mm | [15 empty] | Off
|  | Phantom 120kVp | [22 empty] | 3 | 3 | 131 | I30f medium smooth | Abdomen |  |  |  | PHYSICSMIDIA | TEST_PACS | SYNGO_VIA | [7 empty] | manual
|  | Phantom 120kVp :-D | [22 empty] | 0.6 | 0.6 | 653 | I30f medium smooth | 400/40 Abd |  |  |  | PHYSICSMIDIA | TEST_PACS | [8 empty] | manual
| DECT |  |  | 100 | 240 |  | 202 | 156 |  | 3 | On | CARE Dose4D | 18.61 | 30s | 25% | 50 |  | 0.5 |  | 2 | 0.6 |  | 3 | 32x0.6mm | [15 empty] | Off
|  | Phantom ABM | [22 empty] | 3 | 3 | 88 | Q40f medium | Abdomen |  |  |  | PHYSICSMIDIA | TEST_PACS | SYNGO_VIA | [7 empty] | manual
|  | Phantom 50k | [22 empty] | 3 | 3 | 88 | Q40f medium | Abdomen |  |  |  | PHYSICSMIDIA | TEST_PACS | SYNGO_VIA | [7 empty] | manual
|  | Phantom 70k | [22 empty] | 3 | 3 | 88 | Q40f medium | Abdomen |  |  |  | PHYSICSMIDIA | TEST_PACS | SYNGO_VIA | [7 empty] | manual
|  | Phantom IodCE | [22 empty] | 3 | 3 | 88 | Q40f medium | Abdomen |  |  |  | PHYSICSMIDIA | TEST_PACS | SYNGO_VIA | [7 empty] | manual
|  | Phantom WatVNC | [22 empty] | 3 | 3 | 88 | Q40f medium | Abdomen |  |  |  | PHYSICSMIDIA | TEST_PACS | SYNGO_VIA | [7 empty] | manual
|  | Phantom ABM 3dt | [22 empty] | 1 | 0.8 | 329 | Q40f medium | 400/40 Abd |  |  |  | PHYSICSMIDIA | TEST_PACS | SYNGO_VIA | [7 empty] | manual
75818_Che_3_Abd_cho_qua_Pel_IRB14000308 (Adult)
75818_Che_3_Abd_cho_qua_Pel_IRB14000308 (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 521 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Abd |  | 120 | 100 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 10.75 | 30s | 25% | 50 |  | 0.5 |  | 2 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | Abd a75818 | [22 empty] | 3 | 3 | 77 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | ven Abd cor | [22 empty] | 3 | 3 | 96 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | Abd a75818 :-D | [22 empty] | 0.6 | 0.6 | 383 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
| PreMonitoring |  | 120 | 120 | 200 | 20 |  |  | Off | 7 | Off | CARE Dose4D | 1.23 |  | 25% |  |  | 0.5 |  | 4 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | PreMonitoring | [22 empty] | 10 |  | 1 | B30s medium smooth | Abdomen | [13 empty] | manual
| Monitoring |  | 120 | 120 | 200 | 20 |  |  | Off | 7 | Off | CARE Dose4D | 36.98 |  | 25% |  |  | 0.5 |  | 12 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | Monitoring | [22 empty] | 10 |  | 30 | B30s medium smooth | Abdomen | [13 empty] | manual
| art Abd |  | 120 | 100 | 180 | 110 |  |  | On | 6 | On | CARE Dose4D | 10.49 | 30s | 25% | 50 |  | 0.5 |  | 16 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | art Abd a75181 | [22 empty] | 3 | 2 | 117 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd cor | [22 empty] | 3 | 3 | 94 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd sag | [22 empty] | 3 | 3 | 122 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art Abd a75818 :-D | [22 empty] | 0.6 | 0.6 | 389 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
| ven Che |  | 120 | 100 | 150 | 110 |  |  | On | 7 | On | CARE Dose4D | 8.31 | 30s | 25% | 50 |  | 0.5 |  | 2 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven Che | [22 empty] | 3 | 2 | 182 | I41f medium | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | ven Che sag | [22 empty] | 2 | 2 | 211 | I41f medium | 1500/600 Lung |  |  |  | GEPACS | [9 empty] | manual
|  | ven Che cor mip | [22 empty] | 10 | 10 | 29 | I41f medium | 1500/600 Lung |  |  |  | GEPACS | [9 empty] | manual
|  | ven Che c75181 :-D | [22 empty] | 0.6 | 0.6 | 605 | I41f medium | Abdomen |  |  |  | ROUTER | GEPACS | [8 empty] | manual
| ven AbdPel |  | 120 | 100 | 180 | 110 |  |  | On | 6 | On | CARE Dose4D | 10.49 | 30s | 25% | 50 |  | 0.5 |  | 60 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven AbdPel a75818 | [22 empty] | 3 | 3 | 148 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel cor | [22 empty] | 3 | 3 | 94 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel sag | [22 empty] | 3 | 3 | 130 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel a75818:-D | [22 empty] | 0.6 | 0.6 | 737 | I30f medium smooth | Abdomen |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
| 5m Abd |  | 120 | 100 | 180 | 110 |  |  | On | 6 | On | CARE Dose4D | 10.49 | 30s | 25% | 50 |  | 0.5 |  | 2 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | 5m Abd a75818 | [22 empty] | 3 | 3 | 80 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | 5m Abd a75818 :-D | [22 empty] | 0.6 | 0.6 | 397 | I30f medium smooth | Abdomen |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
BodyMultiPhase_VAngio (Adult)
BodyMultiPhase_VAngio (Adult) | DynMulti4D_L |  |  | 100 |  | 150 |  |  |  | 3 | Off | CARE Dose | 76.8 |  | 25% |  |  | 0.285 |  | 2 | 1 |  | 5 | 32x1.2mm
|  | DynMulti4D_L 5.0 B30f | [22 empty] | 5 | 3 | 710 | B30f medium smooth | Abdomen
|  | DynMulti4D_L 3.0 B20f | [22 empty] | 3 | 2 | 1070 | B20f smooth | Abdomen
|  | DynMulti4D_L 1.5 B20f | [22 empty] | 1.5 | 1 | 2150 | B20f smooth | Abdomen
BodyPCT (Adult)
BodyPCT (Adult) | DynMulti |  |  | 100 |  | 150 |  |  |  | 3 | Off | CARE Dose | 252.56 |  | 25% |  |  | 1 |  | 2 | 0 |  | 10 | 32x1.2mm | [15 empty] | Off
|  | DynMulti 10.0 B30s | [22 empty] | 10 | 7 | 195 | B30s medium smooth | Abdomen | [13 empty] | manual
BodyPCTSeq (Adult)
BodyPCTSeq (Adult) | DynSerio |  |  | 100 | 150 | 150 |  |  |  | 3 | Off | CARE Dose | 265.19 |  | 25% |  |  | 0.5 |  | 2 | 0 | 0 | 10 | 32x1.2mm | [15 empty] | Off
|  | DynSerio 10.0 B30f | [22 empty] | 10 |  | 210 | B30f medium smooth | Abdomen | [13 empty] | manual
BodyPCTSingleSeq (Adult)
BodyPCTSingleSeq (Adult) | DynSerio |  |  | 100 | 150 | 150 |  |  |  | 3 | Off | CARE Dose | 236.85 |  | 25% |  |  | 0.5 |  | 2 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | DynSerio 10.0 B30s | [22 empty] | 10 |  | 42 | B30s medium smooth | Abdomen | [13 empty] | manual
BodyVPCT (Adult)
BodyVPCT (Adult) | DynMulti4D |  |  | 100 |  | 150 |  |  |  | 3 | Off | CARE Dose | 146.58 |  | 25% |  |  | 0.285 |  | 2 | 0.5 |  | 5 | 32x1.2mm
|  | VPCT DynMulti4D 5.0 B20f | [22 empty] | 5 | 3 | 864 | B20f smooth | Abdomen
|  | VPCT DynMulti4D 3.0 B20f | [22 empty] | 3 | 2 | 1323 | B20f smooth | Abdomen
|  | DynMulti4D 1.5 B20f | [22 empty] | 1.5 | 1 | 2673 | B20f smooth | Abdomen
BodyVPCT_long (Adult)
BodyVPCT_long (Adult) | DynMulti4D_L |  |  | 100 |  | 150 |  |  |  | 3 | Off | CARE Dose | 222.29 |  | 25% |  |  | 0.285 |  | 2 | 1.02 |  | 5 | 32x1.2mm
|  | VPCT DynMulti4D_L 5.0 B20f | [22 empty] | 5 | 3 | 1323 | B20f smooth | Abdomen
|  | VPCT DynMulti4D_L 3.0 B20f | [22 empty] | 3 | 2 | 1998 | B20f smooth | Abdomen
|  | DynMulti4D_L 1.5 B20f | [22 empty] | 1.5 | 1 | 4023 | B20f smooth | Abdomen
DynMultiECG (Adult)
DynMultiECG (Adult) | DynMulti |  | 80 | 80 |  | 100 |  |  | Off | 3 | Off | CARE Dose | 169.9 |  | 25% |  |  | 0.285 |  | 2 | 0 |  | 3 | 32x1.2mm | [15 empty] | Normal | 70% | 70%
|  | DynMulti 3.0 B30f 70% | [22 empty] | 3 | 1.5 | 720 | B30f medium smooth | Mediastinum | [13 empty] | manual | 70%
DynSerioECGSeq (Adult)
DynSerioECGSeq (Adult) | DynSerio |  | 80 | 80 | 100 | 100 |  |  | Off | 3 | Off | CARE Dose | 35 |  | 25% |  |  | 0.285 |  | 2 | 0 | 0 | 3 | 32x1.2mm | [13 empty] | 70% | 70% | Off
|  | DynSerio 3.0 B30f 70% | [22 empty] | 3 |  | 720 | B30f medium smooth | Mediastinum | [13 empty] | manual | 70%
ExtrMultiPhase_VAngioPlus (Adult)
ExtrMultiPhase_VAngioPlus (Adult) | DynMulti4D_XL |  |  | 100 |  | 100 |  |  |  | 3 | Off | CARE Dose | 41.49 |  | 25% |  |  | 0.285 |  | 2 | 1.48 |  | 5 | 32x1.2mm
|  | DynMulti4D_XL 5.0 B30f | [22 empty] | 5 | 3 | 1008 | B30f medium smooth | CT Angio
|  | DynMulti4D_XL 3.0 B20f | [22 empty] | 3 | 2 | 1520 | B20f smooth | CT Angio
|  | DynMulti4D_XL 1.5 B20f | [22 empty] | 1.5 | 1 | 3048 | B20f smooth | CT Angio
Flash_HeadTrauma (Adult)
Flash_HeadTrauma (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.29 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head
| Fl_Head |  | 120 | 120 | 210 |  | 210 |  | On | 3 | On | CARE Dose4D | 32.14 | 5s | 25% |  |  | 0.285 |  | 2 | 1.55 |  | 5 | 128x0.6mm
|  | Fl_Head 5.0 H31f | [22 empty] | 5 | 5 | 25 | H31f medium smooth + | Cerebrum
|  | Fl_Head 5.0 H60f | [22 empty] | 5 | 5 | 25 | H60f sharp FR | Bone
Flash_Trauma (Adult)
Flash_Trauma (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 365 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Fl_Trauma |  | 140 | 140 | 140 |  | 140 |  | On | 3 | On | CARE Dose4D | 15.43 | 30s | 25% |  |  | 0.285 |  | 2 | 2.1 |  | 7 | 128x0.6mm
|  | Fl_Trauma 7.0 B30f | [22 empty] | 7 | 7 | 108 | B30f medium smooth | Abdomen
|  | Fl_Trauma 1.5 B30f | [22 empty] | 1.5 | 1 | 751 | B30f medium smooth | Abdomen
|  | Fl_Trauma 1.5 B70f | [22 empty] | 1.5 | 1 | 751 | B70f very sharp | Bone
HeadTrauma (Adult)
HeadTrauma (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.29 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head
| Head |  | 120 | 120 | 390 | 390 |  |  | On | 3 | On | CARE Dose4D | 55.12 | 9s | 25% |  |  | 1 |  | 2 | 0.55 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | Head 5.0 H31s | [22 empty] | 5 | 5 | 25 | H31s medium smooth + | Cerebrum | [13 empty] | manual
|  | Head 5.0 H60s | [22 empty] | 5 | 5 | 25 | H60s sharp FR | Bone | [13 empty] | manual
HeadTraumaSeq (Adult)
HeadTraumaSeq (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.29 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head
| HeadSeq |  | 120 | 120 | 390 | 390 |  |  | On | 3 | On | CARE Dose4D | 62.46 |  | 25% |  |  | 1 |  | 2 | 0 | 34 | 5 | 32x1.2mm | [15 empty] | Off
|  | HeadSeq 5.0 H31s | [22 empty] | 5 |  | 27 | H31s medium smooth + | Base Orbita | [13 empty] | manual
|  | HeadSeq 5.0 H60s | [22 empty] | 5 |  | 27 | H60s sharp FR | Bone | [13 empty] | manual
NeuroMultiPhase_VAngio (Adult)
NeuroMultiPhase_VAngio (Adult) | DynMulti4D_L |  | 80 | 80 |  | 150 |  |  | Off | 3 | Off | CARE Dose | 297.06 |  | 25% |  |  | 0.285 |  | 2 | 1.02 |  | 5 | 32x1.2mm
|  | DynMulti4D_L 5.0 H30f | [22 empty] | 5 | 3 | 1470 | H30f medium smooth | Cerebrum
|  | DynMulti4D_L 3.0 H30f | [22 empty] | 3 | 2 | 2220 | H30f medium smooth | Cerebrum
|  | DynMulti4D_L 1.5 H30f | [22 empty] | 1.5 | 1 | 4470 | H30f medium smooth | Cerebrum
NeuroPCT (Adult)
NeuroPCT (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.29 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head
| Head |  | 120 | 120 | 390 | 390 |  |  | On | 3 | On | CARE Dose4D | 71.99 | 76s | 25% |  |  | 1 |  | 2 | 0.55 |  | 5 | 20x0.6mm | [15 empty] | Off
|  | Head 5.0 H31s | [22 empty] | 5 | 5 | 25 | H31s medium smooth + | Base Orbita | [13 empty] | manual
| DynMulti |  | 80 | 80 |  | 200 |  |  | Off | 3 | Off | CARE Dose | 405 |  | 25% |  |  | 1 |  | 2 | 0 |  | 10 | 32x1.2mm | [15 empty] | Off
|  | PCT DynMulti 10.0 H30s | [22 empty] | 10 | 7 | 195 | H30s medium smooth | Cerebrum | [13 empty] | manual
NeuroPCTSingle (Adult)
NeuroPCTSingle (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.29 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head
| Head |  | 120 | 120 | 390 | 390 |  |  | On | 3 | On | CARE Dose4D | 71.99 | 76s | 25% |  |  | 1 |  | 2 | 0.55 |  | 5 | 20x0.6mm | [15 empty] | Off
|  | Head 5.0 H31s | [22 empty] | 5 | 5 | 25 | H31s medium smooth + | Base Orbita | [13 empty] | manual
| DynMulti |  | 80 | 80 |  | 200 |  |  | Off | 3 | Off | CARE Dose | 361.73 |  | 25% |  |  | 1 |  | 2 | 0 |  | 10 | 1x10.0mm | [15 empty] | Off
|  | PCT DynMulti 10.0 H30s | [22 empty] | 10 | 10 | 39 | H30s medium smooth | Cerebrum | [13 empty] | manual
NeuroVPCT (Adult)
NeuroVPCT (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.29 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head
| Head |  | 120 | 120 | 320 | 320 |  |  | On | 3 | On | CARE Dose4D | 58.93 | 76s | 25% |  |  | 1 |  | 2 | 0.55 |  | 5 | 20x0.6mm | [15 empty] | Off
|  | Head 5.0 H31s | [22 empty] | 5 | 5 | 25 | H31s medium smooth + | Base Orbita | [13 empty] | manual
|  | Head 1.0 H23s | [22 empty] | 1 | 0.7 | 172 | H23s smooth PBV FR | CT Angio | [13 empty] | manual
| DynMulti4D |  | 80 | 80 |  | 200 |  |  | Off | 3 | Off | CARE Dose | 259.69 |  | 25% |  |  | 0.285 |  | 2 | 0.5 |  | 10 | 32x1.2mm
|  | VPCT DynMulti4D 10.0 H20f | [22 empty] | 10 | 4 | 690 | H20f smooth | Cerebrum
|  | VPCT DynMulti4D 5.0 H20f | [22 empty] | 5 | 3 | 960 | H20f smooth | Cerebrum
|  | VPCT DynMulti4D 1.5 H20f | [22 empty] | 1.5 | 1 | 2970 | H20f smooth | Cerebrum
| HeadAngio |  | 100 | 100 | 175 | 175 |  |  | On | 11 | On | CARE Dose4D | 16.89 | 5s | 25% |  |  | 0.5 |  | 2 | 0.8 |  | 1 | 128x0.6mm | [15 empty] | Off
|  | HeadAngio 1.0 H22f | [22 empty] | 1 | 0.7 | 172 | H22f smooth FR | CT Angio | [13 empty] | manual
PolyTrauma (Adult)
PolyTrauma (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.29 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head
| Head |  | 120 | 120 | 390 | 390 |  |  | On | 3 | On | CARE Dose4D | 60.77 | 7s | 25% |  |  | 0.5 |  | 2 | 0.55 |  | 6 | 32x1.2mm | [15 empty] | Off
|  | Head 6.0 H31f | [22 empty] | 6 | 6 | 21 | H31f medium smooth + | Base Orbita | [13 empty] | manual
| Neck |  | 120 | 120 | 165 | 165 |  |  | On | 3 | On | CARE Dose4D | 11.39 | 10s | 25% |  |  | 0.5 |  | 5 | 0.8 |  | 5 | 32x1.2mm | [15 empty] | Off
|  | Neck 5.0 B31f | [22 empty] | 5 | 5 | 29 | B31f medium smooth + | Larynx | [13 empty] | manual
| Topogram | Topogram 0.6 T20f |  | 120 |  | 275 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| ThorAbd |  | 120 | 120 | 210 | 210 |  |  | On | 3 | On | CARE Dose4D | 14.51 | 30s | 25% |  |  | 0.5 |  | 2 | 1.4 |  | 7 | 32x1.2mm | [15 empty] | Off
|  | Tho ThorAbd 7.0 B31f | [22 empty] | 7 | 7 | 29 | B31f medium smooth + | Mediastinum | [13 empty] | manual
|  | Lun ThorAbd 7.0 B80f | [22 empty] | 7 | 7 | 29 | B80f ultra sharp | Lung | [13 empty] | manual
|  | Abd ThorAbd 7.0 B30f | [22 empty] | 7 | 7 | 58 | B30f medium smooth | Abdomen | [13 empty] | manual
|  | MPR ThorAbd 1.5 B30f | [22 empty] | 1.5 | 1 | 601 | B30f medium smooth | Mediastinum | [13 empty] | manual
SELECT_PROTOCOL (Adult)
SELECT__PROTOCOL (Adult)
Trauma (Adult)
Trauma (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 365 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Trauma |  | 120 | 120 | 180 | 180 |  |  | On | 3 | On | CARE Dose4D | 12.42 | 30s | 25% |  |  | 0.5 |  | 2 | 1.5 |  | 7 | 32x1.2mm | [15 empty] | Off
|  | Trauma 7.0 B30f | [22 empty] | 7 | 7 | 108 | B30f medium smooth | Abdomen | [13 empty] | manual
|  | Trauma 1.5 B30f | [22 empty] | 1.5 | 1 | 751 | B30f medium smooth | Abdomen | [13 empty] | manual
|  | Trauma 1.5 B70f | [22 empty] | 1.5 | 1 | 751 | B70f very sharp | Bone | [13 empty] | manual
TraumaThin (Adult)
TraumaThin (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 365 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| TraumaThin |  | 120 | 120 | 180 | 180 |  |  | On | 3 | On | CARE Dose4D | 12.14 | 30s | 25% |  |  | 0.5 |  | 2 | 1.5 |  | 7 | 128x0.6mm | [15 empty] | Off
|  | TraumaThin 7.0 B30f | [22 empty] | 7 | 7 | 108 | B30f medium smooth | Abdomen | [13 empty] | manual
|  | TraumaThin 0.6 B30f | [22 empty] | 0.6 | 0.3 | 2501 | B30f medium smooth | Abdomen | [13 empty] | manual
|  | TraumaThin 0.6 B70f | [22 empty] | 0.6 | 0.3 | 2501 | B70f very sharp | Bone | [13 empty] | manual
UpperExtremities
4000_ang_UppExt (Adult)
4000_ang_UppExt (Adult) | Topogram PA | Topogram |  | 120 |  | 471 |  |  |  | 3 | Off | CARE Dose | 0.24 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| Topogram LAT | Topogram |  | 120 |  | 471 |  |  |  | 3 | Off | CARE Dose | 0.24 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| PreMonitoring |  | 120 | 120 | 250 | 50 |  |  | Off | 3 | Off | CARE Dose4D | 3.08 |  | 25% |  |  | 0.5 |  | 2 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | PreMonitoring 10.0 B70s | [22 empty] | 10 |  | 1 | B70s very sharp | Abdomen | [13 empty] | manual
| Monitoring |  | 120 | 120 | 250 | 50 |  |  | Off | 3 | Off | CARE Dose4D | 123.28 |  | 25% |  |  | 0.5 |  | 10 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | Monitoring 10.0 B70s | [22 empty] | 10 |  | 40 | B70s very sharp | Abdomen | [13 empty] | manual
| art UppExt |  | 120 | 120 | 180 | 165 |  |  | Off | 3 | On | CARE Dose4D | 12.15 | 30s | 25% | 60 |  | 0.5 |  | 2 | 0.8 |  | 2 | 128x0.6mm | [15 empty] | Off
|  | art UppExt u000 | [22 empty] | 2 | 2 | 232 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | art UppExt cor | [22 empty] | 2 | 2 | 181 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | art UppExt sag | [22 empty] | 2 | 2 | 181 | I30f medium smooth | Abdomen | [13 empty] | manual
|  | art UppExt cor mip | [22 empty] | 20 | 5 | 73 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | art UppExt u000 :-D | [22 empty] | 0.6 | 0.6 | 771 | I26f medium smooth ASA | Abdomen |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
| art Arm |  | 120 | 120 | 180 | 165 |  |  | Off | 3 | On | CARE Dose4D | 12.15 | 30s | 25% | 60 |  | 0.5 |  | 2 | 0.8 |  | 2 | 128x0.6mm | [15 empty] | Off
|  | art Arm u000 | [22 empty] | 2 | 2 | 232 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | art Arm cor | [22 empty] | 2 | 2 | 181 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | art Arm sag | [22 empty] | 2 | 2 | 181 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | art Arm cor mip | [22 empty] | 20 | 5 | 73 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | art Arm u000 :-D | [22 empty] | 0.6 | 0.6 | 771 | I26f medium smooth ASA | Abdomen |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
| ven Arm |  | 120 | 120 | 180 | 165 |  |  | Off | 3 | On | CARE Dose4D | 12.15 | 30s | 25% | 60 |  | 0.5 |  | 9 | 0.8 |  | 2 | 128x0.6mm | [15 empty] | Off
|  | ven Arm u000 | [22 empty] | 2 | 2 | 232 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | ven Arm cor | [22 empty] | 2 | 2 | 181 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | ven Arm sag | [22 empty] | 2 | 2 | 181 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | ven Arm cor mip | [22 empty] | 20 | 5 | 73 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | ven Arm u000 :-D | [22 empty] | 0.6 | 0.6 | 771 | I26f medium smooth ASA | Abdomen |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
4010_Cla_uni_sim_wo_cst (Adult)
4010_Cla_uni_sim_wo_cst (Adult) | Topogram PA | Topogram PA 0.6 T20s |  | 120 |  | 164 |  |  |  | 3 | Off | CARE Dose | 0.24 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| Topogram LAT | Topogram LAT 0.6 T20s |  | 120 |  | 164 |  |  |  | 3 | Off | CARE Dose | 0.24 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| Clavicle simple |  | 120 | 140 | 250 | 165 |  |  | Off | 3 | On | CARE Dose4D | 24.55 | 30s | 25% | 60 |  | 1 |  | 2 | 1 |  | 0.6 | 128x0.6mm | [15 empty] | Off
|  | Cla bon u010 :-D | [22 empty] | 0.6 | 0.6 | 771 | B70s very sharp | Bone |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
|  | Cla bon u010 | [22 empty] | 2 | 2 | 76 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Cla :-D | [22 empty] | 0.6 | 0.6 | 251 | B26s medium smooth ASA | Shoulder |  |  |  | ROUTER | GEPACS | [8 empty] | manual
|  | Cla | [22 empty] | 2 | 2 | 76 | B40s medium | Soft Tissue |  |  |  | GEPACS | [9 empty] | manual
4020_Cla_uni_com_wo_cst (Adult)
4020_Cla_uni_com_wo_cst (Adult) | Topogram PA | Topogram PA 0.6 T20s |  | 120 |  | 164 |  |  |  | 3 | Off | CARE Dose | 0.24 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| Topogram LAT | Topogram LAT 0.6 T20s |  | 120 |  | 164 |  |  |  | 3 | Off | CARE Dose | 0.24 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| Clavicle com |  | 120 | 140 | 250 | 165 |  |  | Off | 3 | On | CARE Dose4D | 24.55 | 30s | 25% | 60 |  | 1 |  | 2 | 1 |  | 0.6 | 128x0.6mm | [15 empty] | Off
|  | Cla bon u020 :-D | [22 empty] | 0.6 | 0.6 | 771 | B70s very sharp | Bone |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
|  | Cla bon u020 | [22 empty] | 2 | 2 | 76 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Cla :-D | [22 empty] | 0.6 | 0.6 | 251 | B26s medium smooth ASA | Shoulder |  |  |  | ROUTER | GEPACS | [8 empty] | manual
|  | Cla | [22 empty] | 2 | 2 | 76 | B40s medium | Soft Tissue |  |  |  | GEPACS | [9 empty] | manual
4030_Cla_bil_sim_wo_cst (Adult)
4030_Cla_bil_sim_wo_cst (Adult) | Topogram PA | Topogram PA 0.6 T20s |  | 120 |  | 164 |  |  |  | 3 | Off | CARE Dose | 0.24 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| Topogram LAT | Topogram LAT 0.6 T20s |  | 120 |  | 164 |  |  |  | 3 | Off | CARE Dose | 0.24 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| Cla bil sim |  | 120 | 140 | 250 | 165 |  |  | Off | 3 | On | CARE Dose4D | 24.55 | 30s | 25% | 60 |  | 1 |  | 2 | 1 |  | 0.6 | 128x0.6mm | [15 empty] | Off
|  | Cla bon u030 :-D | [22 empty] | 0.6 | 0.6 | 771 | B70s very sharp | Bone |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
|  | Cla bon u030 | [22 empty] | 2 | 2 | 76 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Cla :-D | [22 empty] | 0.6 | 0.6 | 251 | B26s medium smooth ASA | Shoulder |  |  |  | ROUTER | GEPACS | [8 empty] | manual
|  | Cla | [22 empty] | 2 | 2 | 76 | B40s medium | Soft Tissue |  |  |  | GEPACS | [9 empty] | manual
4040_Cla_bil_com_wo_cst (Adult)
4040_Cla_bil_com_wo_cst (Adult) | Topogram PA | Topogram PA 0.6 T20s |  | 120 |  | 164 |  |  |  | 3 | Off | CARE Dose | 0.24 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| Topogram LAT | Topogram LAT 0.6 T20s |  | 120 |  | 164 |  |  |  | 3 | Off | CARE Dose | 0.24 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| Cla bil com |  | 120 | 140 | 250 | 165 |  |  | Off | 3 | On | CARE Dose4D | 24.55 | 30s | 25% | 60 |  | 1 |  | 2 | 1 |  | 0.6 | 128x0.6mm | [15 empty] | Off
|  | Cla bon u040 :-D | [22 empty] | 0.6 | 0.6 | 771 | B70s very sharp | Bone |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
|  | Cla bon u040 | [22 empty] | 2 | 2 | 76 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Cla :-D | [22 empty] | 0.6 | 0.6 | 251 | B26s medium smooth ASA | Shoulder |  |  |  | ROUTER | GEPACS | [8 empty] | manual
|  | Cla | [22 empty] | 2 | 2 | 76 | B40s medium | Soft Tissue |  |  |  | GEPACS | [9 empty] | manual
4050_Elb_sim_wo_cst (Adult)
4050_Elb_sim_wo_cst (Adult) | Topogram PA | Topogram PA 0.6 T20s |  | 120 |  | 164 |  |  |  | 3 | Off | CARE Dose | 0.24 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| Topogram LAT | Topogram LAT 0.6 T20s |  | 120 |  | 164 |  |  |  | 3 | Off | CARE Dose | 0.24 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| Elbow simple |  | 120 | 120 | 250 | 250 |  |  | Off | 3 | Off | CARE Dose4D | 16.87 | 30s | 25% | 60 |  | 1 |  | 2 | 1 |  | 0.6 | 128x0.6mm | [15 empty] | Off
|  | Elb bon u050 :-D | [22 empty] | 0.6 | 0.6 | 771 | B70s very sharp | Bone |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
|  | Elb bon u050 | [22 empty] | 2 | 2 | 76 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Elb :-D | [22 empty] | 0.6 | 0.6 | 251 | B26s medium smooth ASA | Bone |  |  |  | ROUTER | GEPACS | [8 empty] | manual
|  | Elb | [22 empty] | 2 | 2 | 76 | B40s medium | Soft Tissue |  |  |  | GEPACS | [9 empty] | manual
4060_Elb_com_wo_cst (Adult)
4060_Elb_com_wo_cst (Adult) | Topogram PA | Topogram PA 0.6 T20s |  | 120 |  | 164 |  |  |  | 3 | Off | CARE Dose | 0.24 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| Topogram LAT | Topogram LAT 0.6 T20s |  | 120 |  | 164 |  |  |  | 3 | Off | CARE Dose | 0.24 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| Elbow complex |  | 120 | 120 | 250 | 250 |  |  | Off | 3 | Off | CARE Dose4D | 16.87 | 30s | 25% | 60 |  | 1 |  | 2 | 1 |  | 0.6 | 128x0.6mm | [15 empty] | Off
|  | Elb bon u060 :-D | [22 empty] | 0.6 | 0.6 | 771 | B70s very sharp | Bone |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
|  | Elb bon u060 | [22 empty] | 2 | 2 | 76 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Elb :-D | [22 empty] | 0.6 | 0.6 | 251 | B26s medium smooth ASA | Bone |  |  |  | ROUTER | GEPACS | [8 empty] | manual
|  | Elb | [22 empty] | 2 | 2 | 76 | B40s medium | Soft Tissue |  |  |  | GEPACS | [9 empty] | manual
4065_Elb_gou_wo_cst (Adult)
4065_Elb_gou_wo_cst (Adult) | PA Topo | Topogram |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body |  |  |  | GEPACS | TRACKER
| LAT Topo | Topogram |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body |  |  |  | GEPACS | TRACKER
| Elb |  |  | 100 | 124 |  | 250 | 250 |  | 3 | On | CARE Dose | 23.76 | 30s | 25% | 60 |  | 1 |  | 2 | 1.2 |  | 2 | 40x0.6mm | [15 empty] | Off
|  | Elb bon | [22 empty] | 2 | 2 | 120 | B70s very sharp | 2000/500 MSK |  |  |  | GEPACS | [9 empty] | manual
|  | Elb u065 :-D | [22 empty] | 0.6 | 0.3 | 797 | D30s medium smooth | 400/40 MSK/Std |  |  |  | ROUTER | GEPACS | MMWP1 | [7 empty] | manual
|  | Elb bon :-D | [22 empty] | 0.6 | 0.3 | 797 | B70s very sharp | Bone |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
4070_Hum_sim_wo_cst (Adult)
4070_Hum_sim_wo_cst (Adult) | Topogram PA | Topogram PA 0.6 T20s |  | 120 |  | 318 |  |  |  | 3 | Off | CARE Dose | 0.24 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| Topogram LAT | Topogram LAT 0.6 T20s |  | 120 |  | 318 |  |  |  | 3 | Off | CARE Dose | 0.24 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| Humerus simple |  | 120 | 140 | 250 | 165 |  |  | Off | 3 | On | CARE Dose4D | 24.55 | 30s | 25% | 60 |  | 1 |  | 2 | 1 |  | 0.6 | 128x0.6mm | [15 empty] | Off
|  | Hum bon u070 :-D | [22 empty] | 0.6 | 0.6 | 771 | B70s very sharp | Bone |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
|  | Hum bon u070 | [22 empty] | 2 | 2 | 76 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Hum :-D | [22 empty] | 0.6 | 0.6 | 251 | B26s medium smooth ASA | Shoulder |  |  |  | ROUTER | GEPACS | [8 empty] | manual
|  | Hum | [22 empty] | 2 | 2 | 76 | B40s medium | Soft Tissue |  |  |  | GEPACS | [9 empty] | manual
4080_Hum_com_wo_cst (Adult)
4080_Hum_com_wo_cst (Adult) | Topogram PA | Topogram PA 0.6 T20s |  | 120 |  | 318 |  |  |  | 3 | Off | CARE Dose | 0.24 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| Topogram LAT | Topogram LAT 0.6 T20s |  | 120 |  | 318 |  |  |  | 3 | Off | CARE Dose | 0.24 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| Humerus com |  | 120 | 140 | 250 | 165 |  |  | Off | 3 | On | CARE Dose4D | 24.55 | 30s | 25% | 60 |  | 1 |  | 2 | 1 |  | 0.6 | 128x0.6mm | [15 empty] | Off
|  | Hum bon u080 :-D | [22 empty] | 0.6 | 0.6 | 771 | B70s very sharp | Bone |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
|  | Hum bon u080 | [22 empty] | 2 | 2 | 76 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Hum :-D | [22 empty] | 0.6 | 0.6 | 251 | B26s medium smooth ASA | Shoulder |  |  |  | ROUTER | GEPACS | [8 empty] | manual
|  | Hum | [22 empty] | 2 | 2 | 76 | B40s medium | Soft Tissue |  |  |  | GEPACS | [9 empty] | manual
4090_RadUln_sim_wo_cst (Adult)
4090_RadUln_sim_wo_cst (Adult) | Topogram PA | Topogram PA 0.6 T20s |  | 120 |  | 164 |  |  |  | 3 | Off | CARE Dose | 0.24 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| Topogram LAT | Topogram LAT 0.6 T20s |  | 120 |  | 164 |  |  |  | 3 | Off | CARE Dose | 0.24 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| RadUln sim |  | 120 | 120 | 250 | 250 |  |  | Off | 3 | Off | CARE Dose4D | 16.87 | 30s | 25% | 60 |  | 1 |  | 2 | 1 |  | 0.6 | 128x0.6mm | [15 empty] | Off
|  | RadUln bon u090 :-D | [22 empty] | 0.6 | 0.6 | 771 | B70s very sharp | Bone |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
|  | RadUln bon u090 | [22 empty] | 2 | 2 | 76 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | RadUln :-D | [22 empty] | 0.6 | 0.6 | 251 | B26s medium smooth ASA | Bone |  |  |  | ROUTER | GEPACS | [8 empty] | manual
|  | RadUln | [22 empty] | 2 | 2 | 76 | B40s medium | Soft Tissue |  |  |  | GEPACS | [9 empty] | manual
4100_RadUln_com_wo_cst (Adult)
4100_RadUln_com_wo_cst (Adult) | Topogram PA | Topogram PA 0.6 T20s |  | 120 |  | 164 |  |  |  | 3 | Off | CARE Dose | 0.24 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| Topogram LAT | Topogram LAT 0.6 T20s |  | 120 |  | 164 |  |  |  | 3 | Off | CARE Dose | 0.24 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| RadUln com |  | 120 | 120 | 250 | 250 |  |  | Off | 3 | Off | CARE Dose4D | 16.87 | 30s | 25% | 60 |  | 1 |  | 2 | 1 |  | 0.6 | 128x0.6mm | [15 empty] | Off
|  | RadUln bon u100 :-D | [22 empty] | 0.6 | 0.6 | 771 | B70s very sharp | Bone |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
|  | RadUln bon u100 | [22 empty] | 2 | 2 | 76 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | RadUln :-D | [22 empty] | 0.6 | 0.6 | 251 | B26s medium smooth ASA | Bone |  |  |  | ROUTER | GEPACS | [8 empty] | manual
|  | RadUln | [22 empty] | 2 | 2 | 76 | B40s medium | Soft Tissue |  |  |  | GEPACS | [9 empty] | manual
4110_Rib_uni_wo_cst (Adult)
4110_Rib_uni_wo_cst (Adult) | Topogram PA | Topogram PA 0.6 T20s |  | 100 |  | 318 |  |  |  | 3 | Off | CARE Dose | 0.15 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| Topogram LAT | Topogram LAT 0.6 T20s |  | 100 |  | 318 |  |  |  | 3 | Off | CARE Dose | 0.15 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| Rib uni |  | 120 | 120 | 250 | 165 |  |  | Off | 3 | On | CARE Dose4D | 16.87 | 30s | 25% | 60 |  | 1 |  | 2 | 1 |  | 0.6 | 128x0.6mm | [15 empty] | Off
|  | Rib bon u110 :-D | [22 empty] | 0.6 | 0.6 | 771 | B70s very sharp | Bone |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
|  | Rib bon u110 | [22 empty] | 2 | 2 | 76 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Rib :-D | [22 empty] | 0.6 | 0.6 | 251 | B26s medium smooth ASA | Shoulder |  |  |  | ROUTER | GEPACS | [8 empty] | manual
|  | Rib | [22 empty] | 2 | 2 | 76 | B40s medium | Soft Tissue |  |  |  | GEPACS | [9 empty] | manual
|  | Rib cor bon | [22 empty] | 2 | 2 | 181 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Rib sag bon | [22 empty] | 2 | 2 | 181 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
4120_Rib_bil_wo_cst (Adult)
4120_Rib_bil_wo_cst (Adult) | Topogram PA | Topogram PA 0.6 T20s |  | 100 |  | 318 |  |  |  | 3 | Off | CARE Dose | 0.15 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| Topogram LAT | Topogram LAT 0.6 T20s |  | 100 |  | 318 |  |  |  | 3 | Off | CARE Dose | 0.15 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| Rib bi |  | 120 | 120 | 250 | 165 |  |  | Off | 3 | On | CARE Dose4D | 16.87 | 30s | 25% | 60 |  | 1 |  | 2 | 1 |  | 0.6 | 128x0.6mm | [15 empty] | Off
|  | Rib bon u120 :-D | [22 empty] | 0.6 | 0.6 | 771 | B70s very sharp | Bone |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
|  | Rib bon u120 | [22 empty] | 2 | 2 | 76 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Rib :-D | [22 empty] | 0.6 | 0.6 | 251 | B26s medium smooth ASA | Shoulder |  |  |  | ROUTER | GEPACS | [8 empty] | manual
|  | Rib | [22 empty] | 2 | 2 | 76 | B40s medium | Soft Tissue |  |  |  | GEPACS | [9 empty] | manual
|  | Rib cor bon | [22 empty] | 2 | 2 | 181 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Rib sag bon | [22 empty] | 2 | 2 | 181 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
4130_Sca_sim_wo_cst (Adult)
4130_Sca_sim_wo_cst (Adult) | Topogram PA | Topogram PA 0.6 T20s |  | 120 |  | 318 |  |  |  | 3 | Off | CARE Dose | 0.24 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| Topogram LAT | Topogram LAT 0.6 T20s |  | 120 |  | 318 |  |  |  | 3 | Off | CARE Dose | 0.24 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| Scapula sim |  | 120 | 140 | 250 | 165 |  |  | Off | 3 | On | CARE Dose4D | 24.55 | 30s | 25% | 60 |  | 1 |  | 2 | 1 |  | 0.6 | 128x0.6mm | [15 empty] | Off
|  | Sca bon u130 :-D | [22 empty] | 0.6 | 0.6 | 771 | B70s very sharp | Bone |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
|  | Sca bon u130 | [22 empty] | 2 | 2 | 76 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Sca :-D | [22 empty] | 0.6 | 0.6 | 251 | B26s medium smooth ASA | Shoulder |  |  |  | ROUTER | GEPACS | [8 empty] | manual
|  | Sca | [22 empty] | 2 | 2 | 76 | B40s medium | Soft Tissue |  |  |  | GEPACS | [9 empty] | manual
4140_Sca_com_wo_cst (Adult)
4140_Sca_com_wo_cst (Adult) | Topogram PA | Topogram PA 0.6 T20s |  | 120 |  | 318 |  |  |  | 3 | Off | CARE Dose | 0.24 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| Topogram LAT | Topogram LAT 0.6 T20s |  | 120 |  | 318 |  |  |  | 3 | Off | CARE Dose | 0.24 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| Scapula com |  | 120 | 140 | 250 | 165 |  |  | Off | 3 | On | CARE Dose4D | 24.55 | 30s | 25% | 60 |  | 1 |  | 2 | 1 |  | 0.6 | 128x0.6mm | [15 empty] | Off
|  | Sca bon u140 :-D | [22 empty] | 0.6 | 0.6 | 771 | B70s very sharp | Bone |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
|  | Sca bon u140 | [22 empty] | 2 | 2 | 76 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Sca :-D | [22 empty] | 0.6 | 0.6 | 251 | B26s medium smooth ASA | Shoulder |  |  |  | ROUTER | GEPACS | [8 empty] | manual
|  | Sca | [22 empty] | 2 | 2 | 76 | B40s medium | Soft Tissue |  |  |  | GEPACS | [9 empty] | manual
4150_Sho_sim_wo_cst (Adult)
4150_Sho_sim_wo_cst (Adult) | Topogram PA | Topogram PA 0.6 T20s |  | 120 |  | 164 |  |  |  | 3 | Off | CARE Dose | 0.24 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| Topogram LAT | Topogram LAT 0.6 T20s |  | 120 |  | 164 |  |  |  | 3 | Off | CARE Dose | 0.24 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| Shoulder |  | 120 | 140 | 250 | 165 |  |  | Off | 3 | On | CARE Dose4D | 24.55 | 30s | 25% | 60 |  | 1 |  | 2 | 1 |  | 0.6 | 128x0.6mm | [15 empty] | Off
|  | Sho bon u150 :-D | [22 empty] | 0.6 | 0.6 | 771 | B70s very sharp | Bone |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
|  | Sho bon u150 | [22 empty] | 2 | 2 | 76 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Sho :-D | [22 empty] | 0.6 | 0.6 | 251 | B26s medium smooth ASA | Shoulder |  |  |  | ROUTER | GEPACS | [8 empty] | manual
|  | Sho | [22 empty] | 2 | 2 | 76 | B40s medium | Soft Tissue |  |  |  | GEPACS | [9 empty] | manual
4160_Sho_com_wo_cst (Adult)
4160_Sho_com_wo_cst (Adult) | Topogram PA | Topogram PA 0.6 T20s |  | 120 |  | 164 |  |  |  | 3 | Off | CARE Dose | 0.24 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| Topogram LAT | Topogram LAT 0.6 T20s |  | 120 |  | 164 |  |  |  | 3 | Off | CARE Dose | 0.24 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| Shoulder |  | 120 | 140 | 250 | 165 |  |  | Off | 3 | On | CARE Dose4D | 24.55 | 30s | 25% | 60 |  | 1 |  | 2 | 1 |  | 0.6 | 128x0.6mm | [15 empty] | Off
|  | Sho bon u160 :-D | [22 empty] | 0.6 | 0.6 | 771 | B70s very sharp | Bone |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
|  | Sho bon u160 | [22 empty] | 2 | 2 | 76 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Sho :-D | [22 empty] | 0.6 | 0.6 | 251 | B26s medium smooth ASA | Shoulder |  |  |  | ROUTER | GEPACS | [8 empty] | manual
|  | Sho | [22 empty] | 2 | 2 | 76 | B40s medium | Soft Tissue |  |  |  | GEPACS | MCH3DLabAW | [8 empty] | manual
4170_Ste_joi_sim_wo_cst (Adult)
4170_Ste_joi_sim_wo_cst (Adult) | Topogram PA | Topogram PA 0.6 T20s |  | 120 |  | 164 |  |  |  | 3 | Off | CARE Dose | 0.24 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| Topogram LAT | Topogram LAT 0.6 T20s |  | 120 |  | 164 |  |  |  | 3 | Off | CARE Dose | 0.24 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| Ste joi sim |  | 120 | 140 | 250 | 165 |  |  | Off | 3 | On | CARE Dose4D | 24.55 | 30s | 25% | 60 |  | 1 |  | 2 | 1 |  | 0.6 | 128x0.6mm | [15 empty] | Off
|  | Ste joi bon u170 :-D | [22 empty] | 0.6 | 0.6 | 771 | B70s very sharp | Bone |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
|  | Ste joi bon u170 | [22 empty] | 2 | 2 | 76 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Ste joi :-D | [22 empty] | 0.6 | 0.6 | 251 | B26s medium smooth ASA | Shoulder |  |  |  | ROUTER | GEPACS | [8 empty] | manual
|  | Ste joi | [22 empty] | 2 | 2 | 76 | B40s medium | Soft Tissue |  |  |  | GEPACS | [9 empty] | manual
4180_Ste_joi_com_wo_cst (Adult)
4180_Ste_joi_com_wo_cst (Adult) | Topogram PA | Topogram PA 0.6 T20s |  | 120 |  | 164 |  |  |  | 3 | Off | CARE Dose | 0.24 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| Topogram LAT | Topogram LAT 0.6 T20s |  | 120 |  | 164 |  |  |  | 3 | Off | CARE Dose | 0.24 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| Ste joi complex |  | 120 | 140 | 250 | 165 |  |  | Off | 3 | On | CARE Dose4D | 24.55 | 30s | 25% | 60 |  | 1 |  | 2 | 1 |  | 0.6 | 128x0.6mm | [15 empty] | Off
|  | Ste joi bon u180 :-D | [22 empty] | 0.6 | 0.6 | 771 | B70s very sharp | Bone |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
|  | Ste joi bon u180 | [22 empty] | 2 | 2 | 76 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Ste joi :-D | [22 empty] | 0.6 | 0.6 | 251 | B26s medium smooth ASA | Shoulder |  |  |  | ROUTER | GEPACS | [8 empty] | manual
|  | Ste joi | [22 empty] | 2 | 2 | 76 | B40s medium | Soft Tissue |  |  |  | GEPACS | [9 empty] | manual
4190_Ste_sim_wo_cst (Adult)
4190_Ste_sim_wo_cst (Adult) | Topogram PA | Topogram PA 0.6 T20s |  | 100 |  | 318 |  |  |  | 3 | Off | CARE Dose | 0.15 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| Topogram LAT | Topogram LAT 0.6 T20s |  | 100 |  | 318 |  |  |  | 3 | Off | CARE Dose | 0.15 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| Sternum simple |  | 120 | 140 | 250 | 165 |  |  | Off | 3 | On | CARE Dose4D | 24.55 | 30s | 25% | 60 |  | 1 |  | 2 | 1 |  | 0.6 | 128x0.6mm | [15 empty] | Off
|  | Ste bon u190 :-D | [22 empty] | 0.6 | 0.6 | 771 | B70s very sharp | Bone |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
|  | Ste bon u190 | [22 empty] | 2 | 2 | 76 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Ste :-D | [22 empty] | 0.6 | 0.6 | 251 | B26s medium smooth ASA | Shoulder |  |  |  | ROUTER | GEPACS | [8 empty] | manual
|  | Ste | [22 empty] | 2 | 2 | 76 | B40s medium | Soft Tissue |  |  |  | GEPACS | [9 empty] | manual
|  | Ste obl cor bon | [22 empty] | 0.6 | 0.6 | 601 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Ste obl sag bon | [22 empty] | 0.6 | 0.6 | 601 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
4200_Ste_com_wo_cst (Adult)
4200_Ste_com_wo_cst (Adult) | Topogram PA | Topogram PA 0.6 T20s |  | 100 |  | 318 |  |  |  | 3 | Off | CARE Dose | 0.15 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| Topogram LAT | Topogram LAT 0.6 T20s |  | 100 |  | 318 |  |  |  | 3 | Off | CARE Dose | 0.15 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| Sternum com |  | 120 | 140 | 250 | 165 |  |  | Off | 3 | On | CARE Dose4D | 24.55 | 30s | 25% | 60 |  | 1 |  | 2 | 1 |  | 0.6 | 128x0.6mm | [15 empty] | Off
|  | Ste bon u200 :-D | [22 empty] | 0.6 | 0.6 | 771 | B70s very sharp | Bone |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
|  | Ste bon u200 | [22 empty] | 2 | 2 | 76 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Ste :-D | [22 empty] | 0.6 | 0.6 | 251 | B26s medium smooth ASA | Shoulder |  |  |  | ROUTER | GEPACS | [8 empty] | manual
|  | Ste | [22 empty] | 2 | 2 | 76 | B40s medium | Soft Tissue |  |  |  | GEPACS | [9 empty] | manual
|  | Ste obl cor bon | [22 empty] | 0.6 | 0.6 | 601 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Ste obl sag bon | [22 empty] | 0.6 | 0.6 | 601 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
4210_WriHan_sim_wo_cst (Adult)
4210_WriHan_sim_wo_cst (Adult) | Topogram PA | Topogram PA 0.6 T20s |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| Topogram LAT | Topogram LAT 0.6 T20s |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| WriHan sim |  |  | 120 | 180 | 250 |  |  |  | 3 | Off | CARE Dose | 24.29 | 30s | 25% | 60 |  | 1 |  | 2 | 0.8 |  | 0.4 | 16x0.3mm | [15 empty] | Off
|  | WriHan bon u210 :-D | [22 empty] | 0.4 | 0.2 | 301 | U70u sharp | 2000/500 MSK |  | UHR |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
|  | WriHan bon u210 | [22 empty] | 2 | 2 | 31 | U70u sharp | Bone |  | UHR |  | GEPACS | [9 empty] | manual
|  | Wri :-D | [22 empty] | 0.6 | 0.6 | 101 | U30u medium smooth | Soft Tissue |  | UHR |  | ROUTER | GEPACS | [8 empty] | manual
|  | Wri | [22 empty] | 2 | 2 | 31 | U30u medium smooth | Soft Tissue |  | UHR |  | GEPACS | [9 empty] | manual
4220_WriHan_com_wo_cst (Adult)
4220_WriHan_com_wo_cst (Adult) | Topogram PA | Topogram PA 0.6 T20s |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| Topogram LAT | Topogram LAT 0.6 T20s |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| WriHan com |  |  | 120 | 180 | 250 |  |  |  | 3 | Off | CARE Dose | 24.29 | 30s | 25% | 60 |  | 1 |  | 2 | 0.8 |  | 0.4 | 16x0.3mm | [15 empty] | Off
|  | WriHan bon u220 :-D | [22 empty] | 0.4 | 0.2 | 301 | U70u sharp | 2000/500 MSK |  | UHR |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
|  | WriHan bon u220 | [22 empty] | 2 | 2 | 31 | U70u sharp | Bone |  | UHR |  | GEPACS | [9 empty] | manual
|  | Wri :-D | [22 empty] | 0.6 | 0.6 | 101 | U30u medium smooth | Soft Tissue |  | UHR |  | ROUTER | GEPACS | [8 empty] | manual
|  | Wri | [22 empty] | 2 | 2 | 31 | U30u medium smooth | Soft Tissue |  | UHR |  | GEPACS | [9 empty] | manual
4230_WriHan_gou_wo_cst (Adult)
4230_WriHan_gou_wo_cst (Adult) | PA Topo | PA Topo 0.6 T20s |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body |  |  |  |  | ROEMIR001a | TRACKER
| LAT Topo | LAT Topo 0.6 T20s |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body |  |  |  |  | ROEMIR001a | TRACKER
| DE UE |  |  | 80 | 249 |  | 340 | 170 |  | 3 | On | CARE Dose | 15.73 | 30s | 25% | 60 |  | 1 |  | 2 | 1.2 |  | 2 | 40x0.6mm | [15 empty] | Off
|  | WriHan u230 | [22 empty] | 2 | 2 | 120 | B70s very sharp | 2000/500 MSK |  |  |  | GEPACS | [9 empty] | manual
|  | WriHan u230 :-D | [22 empty] | 0.6 | 0.3 | 797 | D30s medium smooth | 400/40 MSK/Std |  |  |  | ROUTER | GEPACS | MMWP1 | [7 empty] | manual
|  | WriHan u230 :-D | [22 empty] | 0.6 | 0.3 | 797 | B70s very sharp | 400/40 MSK/Std |  |  |  | ROUTER | GEPACS | [8 empty] | manual
4240_Wri_sim_wo_cst (Adult)
4240_Wri_sim_wo_cst (Adult) | Topogram PA | Topogram PA 0.6 T20s |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| Topogram LAT | Topogram LAT 0.6 T20s |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| Wri sim |  |  | 120 | 180 | 250 |  |  |  | 3 | Off | CARE Dose | 24.29 | 30s | 25% | 60 |  | 1 |  | 2 | 0.8 |  | 0.4 | 16x0.3mm | [15 empty] | Off
|  | Wri bon u240 :-D | [22 empty] | 0.4 | 0.2 | 301 | U70u sharp | 2000/500 MSK |  | UHR |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
|  | Wri bon u240 | [22 empty] | 2 | 2 | 31 | U70u sharp | Bone |  | UHR |  | GEPACS | [9 empty] | manual
|  | Wri :-D | [22 empty] | 0.6 | 0.6 | 101 | U30u medium smooth | Soft Tissue |  | UHR |  | ROUTER | GEPACS | [8 empty] | manual
|  | Wri | [22 empty] | 2 | 2 | 31 | U30u medium smooth | Soft Tissue |  | UHR |  | GEPACS | [9 empty] | manual
4250_Wri_com_wo_cst (Adult)
4250_Wri_com_wo_cst (Adult) | Topogram PA | Topogram PA 0.6 T20s |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| Topogram LAT | Topogram LAT 0.6 T20s |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| Wri complex |  |  | 120 | 180 | 250 |  |  |  | 3 | Off | CARE Dose | 24.29 | 30s | 25% | 60 |  | 1 |  | 2 | 0.8 |  | 0.4 | 16x0.3mm | [15 empty] | Off
|  | Wri bon u250 :-D | [22 empty] | 0.4 | 0.2 | 301 | U70u sharp | 2000/500 MSK |  | UHR |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
|  | Wri bon u250 | [22 empty] | 2 | 2 | 31 | U70u sharp | Bone |  | UHR |  | GEPACS | [9 empty] | manual
|  | Wri :-D | [22 empty] | 0.6 | 0.6 | 101 | U30u medium smooth | Soft Tissue |  | UHR |  | ROUTER | GEPACS | [8 empty] | manual
|  | Wri | [22 empty] | 2 | 2 | 31 | U30u medium smooth | Soft Tissue |  | UHR |  | GEPACS | [9 empty] | manual
4260_Wri_DRUJ_wo_cst (Adult)
4260_Wri_DRUJ_wo_cst (Adult) | Topogram PA | Topogram PA 0.6 T20s |  | 120 |  | 164 |  |  |  | 3 | Off | CARE Dose | 0.24 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| Topogram LAT | Topogram LAT 0.6 T20s |  | 120 |  | 164 |  |  |  | 3 | Off | CARE Dose | 0.24 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| Wrist pronate |  |  | 120 | 180 | 250 |  |  |  | 3 | Off | CARE Dose | 24.29 | 30s | 25% | 60 |  | 1 |  | 2 | 0.8 |  | 0.4 | 16x0.3mm | [15 empty] | Off
|  | Wri bon u260 :-D | [22 empty] | 0.4 | 0.2 | 301 | U70u sharp | 2000/500 MSK |  | UHR |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
|  | Wri bon u260 | [22 empty] | 2 | 2 | 31 | U70u sharp | Bone |  | UHR |  | GEPACS | [9 empty] | manual
|  | Wri :-D | [22 empty] | 0.6 | 0.6 | 101 | U30u medium smooth | Soft Tissue |  | UHR |  | ROUTER | GEPACS | [8 empty] | manual
|  | Wri | [22 empty] | 2 | 2 | 31 | U30u medium smooth | Soft Tissue |  | UHR |  | GEPACS | [9 empty] | manual
| Wrist neutral |  |  | 120 | 180 | 250 |  |  |  | 3 | Off | CARE Dose | 24.29 | 30s | 25% | 60 |  | 1 |  | 2 | 0.8 |  | 0.4 | 16x0.3mm | [15 empty] | Off
|  | neu Wri bon u260 :-D | [22 empty] | 0.4 | 0.2 | 301 | U70u sharp | 2000/500 MSK |  | UHR |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
|  | neu Wri bon u260 | [22 empty] | 2 | 2 | 31 | U70u sharp | Bone |  | UHR |  | GEPACS | [9 empty] | manual
| Wrist supinate |  |  | 120 | 180 | 250 |  |  |  | 3 | Off | CARE Dose | 24.29 | 30s | 25% | 60 |  | 1 |  | 2 | 0.8 |  | 0.4 | 16x0.3mm | [15 empty] | Off
|  | pro Wri bon u260 :-D | [22 empty] | 0.4 | 0.2 | 301 | U70u sharp | 2000/500 MSK |  | UHR |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
|  | pro Wri bon u260 | [22 empty] | 2 | 2 | 31 | U70u sharp | Bone |  | UHR |  | GEPACS | [9 empty] | manual
Extremity (Adult)
Extremity (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Extremity |  |  | 120 | 100 | 100 |  |  |  | 3 | On | CARE Dose | 6.75 | 8s | 25% |  |  | 1 |  | 2 | 0.8 |  | 4 | 128x0.6mm | [15 empty] | Off
|  | Extremity 4.0 B60s | [22 empty] | 4 | 4 | 38 | B60s sharp | Extremity | [13 empty] | manual
|  | Extremity 1.0 B60s | [22 empty] | 1 | 0.7 | 215 | B60s sharp | Bone | [13 empty] | manual
ExtrRoutineZUHR (Adult)
ExtrRoutineZUHR (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| ExtrUHR |  |  | 120 | 180 | 180 |  |  |  | 3 | On | CARE Dose | 17.65 | 20s | 25% |  |  | 1 |  | 2 | 0.85 |  | 2 | 16x0.3mm | [15 empty] | Off
|  | ExtrUHR 2.0 U90u | [22 empty] | 2 | 2 | 31 | U90u ultra sharp | Extremity | [13 empty] | manual
|  | ExtrUHR 0.4 U90u | [22 empty] | 0.4 | 0.2 | 301 | U90u ultra sharp | Bone | [13 empty] | manual
ExtrUHR (Adult)
ExtrUHR (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| ExtrUHR |  |  | 120 | 120 | 120 |  |  |  | 3 | On | CARE Dose | 11.77 | 20s | 25% |  |  | 1 |  | 2 | 0.85 |  | 2 | 16x0.6mm | [15 empty] | Off
|  | ExtrUHR 2.0 U90u | [22 empty] | 2 | 2 | 31 | U90u ultra sharp | Extremity | [13 empty] | manual
|  | ExtrUHR 0.6 U90u | [22 empty] | 0.6 | 0.4 | 151 | U90u ultra sharp | Bone | [13 empty] | manual
ExtrZUHR_IR (Adult)
ExtrZUHR_IR (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| ExtrUHR |  |  | 120 | 126 | 126 |  |  |  | 3 | On | CARE Dose | 12.46 | 20s | 25% |  |  | 1 |  | 2 | 0.85 |  | 2 | 16x0.3mm | [15 empty] | Off
|  | ExtrUHR 2.0 V90u 3 | [22 empty] | 2 | 2 | 31 | V90u ultra sharp | Extremity | [13 empty] | manual
|  | ExtrUHR 0.4 V90u 3 | [22 empty] | 0.4 | 0.2 | 301 | V90u ultra sharp | Bone | [13 empty] | manual
WristZUHR (Adult)
WristZUHR (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| WristUHR |  |  | 120 | 180 | 180 |  |  |  | 3 | On | CARE Dose | 17.65 | 20s | 25% |  |  | 1 |  | 2 | 0.85 |  | 2 | 16x0.3mm | [15 empty] | Off
|  | WristUHR 2.0 U90u | [22 empty] | 2 | 2 | 31 | U90u ultra sharp | Extremity | [13 empty] | manual
|  | WristUHR 0.4 U90u | [22 empty] | 0.4 | 0.2 | 301 | U90u ultra sharp | Bone | [13 empty] | manual
LowerExtremities
9000_ang_Leg_run (Adult)
9000_ang_Leg_run (Adult) | Topogram | Topogram |  | 100 |  | 994 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 994 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER
| PreMonitoring |  | 120 | 120 | 350 | 50 |  |  | Off | 3 | Off | CARE Dose4D | 3.08 |  | 25% |  |  | 0.5 |  | 2 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | PreMonitoring | [22 empty] | 10 |  | 1 | B40s medium | 400/40 Abd | [13 empty] | manual
| Monitoring |  | 120 | 120 | 350 | 50 |  |  | Off | 3 | Off | CARE Dose4D | 123.28 |  | 25% |  |  | 0.5 |  | 10 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | Monitoring | [22 empty] | 10 |  | 40 | B40s medium | 400/40 Abd | [13 empty] | manual
| Leg runoff |  | 120 | 120 | 180 | 210 |  |  | On | 8 | On | CARE Dose4D | 12.15 | 30s | 25% | 60 |  | 0.5 |  | 6 | 0.8 |  | 2 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | art Leg l000 | [22 empty] | 2 | 2 | 120 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art Leg l000 :-D | [22 empty] | 0.6 | 0.6 | 400 | I26f medium smooth ASA | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
|  | art Leg cor | [22 empty] | 0.6 | 0.6 | 634 | I26f medium smooth ASA | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
| 2 run if needed |  | 120 | 120 | 180 | 210 |  |  | On | 8 | On | CARE Dose4D | 12.15 | 30s | 25% | 60 |  | 0.5 |  | 2 | 0.8 |  | 2 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | art Leg l000 | [22 empty] | 2 | 2 | 120 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | art Leg l000 :-D | [22 empty] | 0.6 | 0.6 | 400 | I26f medium smooth ASA | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
9010_Ank_sim_wo_cst (Adult)
9010_Ank_sim_wo_cst (Adult) | Topogram PA | Topogram PA 0.6 T20s |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.09 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| Topogram LAT | Topogram LAT 0.6 T20s |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.09 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| Ankle simple |  |  | 120 | 120 | 250 |  |  |  | 3 | Off | CARE Dose | 24.04 | 30s | 25% | 60 |  | 1 |  | 2 | 1 |  | 0.6 | 16x0.3mm | [15 empty] | Off
|  | Ank bon l010 :-D | [22 empty] | 0.6 | 0.4 | 587 | U70u sharp | Bone |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
|  | Ank bon l010 | [22 empty] | 2 | 2 | 117 | U70u sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Ank :-D | [22 empty] | 0.6 | 0.4 | 584 | U30u medium smooth | Soft Tissue |  |  |  | ROUTER | GEPACS | [8 empty] | manual
|  | Ank | [22 empty] | 2 | 2 | 117 | U30u medium smooth | Soft Tissue |  |  |  | GEPACS | [9 empty] | manual
9020_Ank_com_wo_cst (Adult)
9020_Ank_com_wo_cst (Adult) | Topogram PA | Topogram PA 0.6 T20s |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.09 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body |  | right |  |  |  | TRACKER
| Topogram LAT | Topogram LAT 0.6 T20s |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.09 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body |  | right |  |  |  | TRACKER
| Ankle complex |  |  | 120 | 120 | 250 |  |  |  | 3 | Off | CARE Dose | 24.04 | 30s | 25% | 60 |  | 1 |  | 2 | 1 |  | 0.6 | 16x0.3mm | [15 empty] | Off
|  | Ank bon l020 :-D | [22 empty] | 0.6 | 0.4 | 587 | U70u sharp | Bone |  | right |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
|  | Ank bon l020 | [22 empty] | 2 | 2 | 117 | U70u sharp | Bone |  | right |  | GEPACS | [9 empty] | manual
|  | Ank :-D | [22 empty] | 0.6 | 0.4 | 584 | U30u medium smooth | Soft Tissue |  | right |  | ROUTER | GEPACS | [8 empty] | manual
|  | Ank | [22 empty] | 2 | 2 | 117 | U30u medium smooth | Soft Tissue |  | right |  | GEPACS | [9 empty] | manual
9030_AnkFoo_gou_wo_cst (Adult)
9030_AnkFoo_gou_wo_cst (Adult) | PA Topo | PA Topo 0.6 T20s |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| LAT Topo | LAT Topo 0.6 T20s |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| DE LE |  |  | 80 | 249 |  | 340 | 170 |  | 3 | On | CARE Dose | 15.73 | 30s | 25% | 60 |  | 1 |  | 2 | 1.2 |  | 2 | 40x0.6mm | [15 empty] | Off
|  | AnkFoo bon | [22 empty] | 2 | 2 | 120 | B70s very sharp | 2000/500 MSK |  |  |  | GEPACS | [9 empty] | manual
|  | AnkFoo l030 #PP :-D | [22 empty] | 0.6 | 0.3 | 797 | D30s medium smooth | 400/40 MSK/Std |  |  |  | ROUTER | GEPACS | MMWP1 | [7 empty] | manual
|  | AnkFoo l030 bon :-D | [22 empty] | 0.6 | 0.3 | 797 | B70s very sharp | Bone |  |  |  | ROUTER | GEPACS | [8 empty] | manual
9040_Fem_sim_wo_cst (Adult)
9040_Fem_sim_wo_cst (Adult) | Topogram PA | Topogram PA 0.6 T20s |  | 120 |  | 318 |  |  |  | 3 | Off | CARE Dose | 0.24 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| Topogram LAT | Topogram LAT 0.6 T20s |  | 120 |  | 318 |  |  |  | 3 | Off | CARE Dose | 0.24 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| Femur simple |  | 120 | 140 | 250 | 165 |  |  | Off | 3 | On | CARE Dose4D | 24.55 | 30s | 25% | 60 |  | 1 |  | 2 | 1 |  | 0.6 | 128x0.6mm | [15 empty] | Off
|  | Fem bon l040 :-D | [22 empty] | 0.6 | 0.6 | 771 | B70s very sharp | Bone |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
|  | Fem bon l040 | [22 empty] | 2 | 2 | 76 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Fem :-D | [22 empty] | 0.6 | 0.6 | 251 | B26s medium smooth ASA | Shoulder |  |  |  | ROUTER | GEPACS | [8 empty] | manual
|  | Fem | [22 empty] | 2 | 2 | 76 | B40s medium | Soft Tissue |  |  |  | GEPACS | [9 empty] | manual
|  | Fem obl cor bon | [22 empty] | 2 | 2 | 101 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Fem obl sag bon | [22 empty] | 2 | 2 | 101 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
9050_Fem_com_wo_cst (Adult)
9050_Fem_com_wo_cst (Adult) | Topogram PA | Topogram PA 0.6 T20s |  | 120 |  | 318 |  |  |  | 3 | Off | CARE Dose | 0.24 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| Topogram LAT | Topogram LAT 0.6 T20s |  | 120 |  | 318 |  |  |  | 3 | Off | CARE Dose | 0.24 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| Femur complex |  | 120 | 140 | 250 | 165 |  |  | Off | 3 | On | CARE Dose4D | 24.55 | 30s | 25% | 60 |  | 1 |  | 2 | 1 |  | 0.6 | 128x0.6mm | [15 empty] | Off
|  | Fem bon l050 :-D | [22 empty] | 0.6 | 0.6 | 771 | B70s very sharp | Bone |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
|  | Fem bon l050 | [22 empty] | 2 | 2 | 76 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Fem :-D | [22 empty] | 0.6 | 0.6 | 251 | B26s medium smooth ASA | Shoulder |  |  |  | ROUTER | GEPACS | [8 empty] | manual
|  | Fem | [22 empty] | 2 | 2 | 76 | B40s medium | Soft Tissue |  |  |  | GEPACS | [9 empty] | manual
|  | Fem obl cor bon | [22 empty] | 2 | 2 | 101 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Fem obl sag bon | [22 empty] | 2 | 2 | 101 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
9060_Foo_sim_wo_cst (Adult)
9060_Foo_sim_wo_cst (Adult) | Topogram PA | Topogram PA 0.6 T20s |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.09 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| Topogram LAT | Topogram LAT 0.6 T20s |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.09 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| Foot simple |  |  | 120 | 120 | 250 |  |  |  | 3 | Off | CARE Dose | 24.04 | 30s | 25% | 60 |  | 1 |  | 2 | 1 |  | 0.6 | 16x0.3mm | [15 empty] | Off
|  | Foo bon l060 :-D | [22 empty] | 0.6 | 0.4 | 587 | U70u sharp | Bone |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
|  | Foo bon l060 | [22 empty] | 2 | 2 | 117 | U70u sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Foo :-D | [22 empty] | 0.6 | 0.4 | 584 | U30u medium smooth | Soft Tissue |  |  |  | ROUTER | GEPACS | [8 empty] | manual
|  | Foo | [22 empty] | 2 | 2 | 117 | U30u medium smooth | Soft Tissue |  |  |  | GEPACS | [9 empty] | manual
9070_Foo_com_wo_cst (Adult)
9070_Foo_com_wo_cst (Adult) | Topogram PA | Topogram PA 0.6 T20s |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.09 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| Topogram LAT | Topogram LAT 0.6 T20s |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.09 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| Foot complex |  |  | 120 | 120 | 250 |  |  |  | 3 | Off | CARE Dose | 24.04 | 30s | 25% | 60 |  | 1 |  | 2 | 1 |  | 0.6 | 16x0.3mm | [15 empty] | Off
|  | Foo bon l070 :-D | [22 empty] | 0.6 | 0.4 | 587 | U70u sharp | Bone |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
|  | Foo bon l070 | [22 empty] | 2 | 2 | 117 | U70u sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Foo :-D | [22 empty] | 0.6 | 0.4 | 584 | U30u medium smooth | Soft Tissue |  |  |  | ROUTER | GEPACS | [8 empty] | manual
|  | Foo | [22 empty] | 2 | 2 | 117 | U30u medium smooth | Soft Tissue |  |  |  | GEPACS | [9 empty] | manual
9080_HipFemKne_ant_wo_cst (Adult)
9080_HipFemKne_ant_wo_cst (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body |  |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body |  |  |  | GEPACS |  | TRACKER
| HipFemKne |  | 120 | 120 | 240 | 165 |  |  | Off | 3 | On | CARE Dose4D | 16.2 | 30s | 25% | 60 |  | 1 |  | 2 | 0.8 |  | 0.6 | 128x0.6mm | [15 empty] | Off
|  | HipFemKne bon l080 :-D | [22 empty] | 0.6 | 0.4 | 1156 | B70s very sharp | Bone |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
|  | HipFemKne bon l080 | [22 empty] | 2 | 2 | 76 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | HipFemKne :-D | [22 empty] | 0.6 | 0.4 | 376 | B26s medium smooth ASA | Shoulder |  |  |  | ROUTER | GEPACS | [8 empty] | manual
|  | HipFemKne | [22 empty] | 2 | 2 | 76 | B40s medium | Soft Tissue |  |  |  | GEPACS | [9 empty] | manual
|  | HipFemKne obl cor bon | [22 empty] | 2 | 2 | 101 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | HipFemKne obl sag bon | [22 empty] | 2 | 2 | 101 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
9090_HipFemKneTibFibAnk_ant_wo_cst (Adult)
9090_HipFemKneTibFibAnk_ant_wo_cst (Adult) | Topogram | Topogram |  | 100 |  | 393 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body |  |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 393 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body |  |  |  | GEPACS |  | TRACKER
| HipFemKne |  | 120 | 120 | 240 | 165 |  |  | Off | 3 | On | CARE Dose4D | 16.2 | 30s | 25% | 60 |  | 1 |  | 2 | 0.8 |  | 0.6 | 128x0.6mm | [15 empty] | Off
|  | HipFemKneTibFibAnk bon l090 :-D | [22 empty] | 0.6 | 0.4 | 1156 | B70s very sharp | Bone |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
|  | HipFemKneTibFibAnk bon l090 | [22 empty] | 2 | 2 | 76 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | HipFemKneTibFibAnk :-D | [22 empty] | 0.6 | 0.4 | 376 | B26s medium smooth ASA | Shoulder |  |  |  | ROUTER | GEPACS | [8 empty] | manual
|  | HipFemKneTibFibAnk | [22 empty] | 2 | 2 | 76 | B40s medium | Soft Tissue |  |  |  | GEPACS | [9 empty] | manual
|  | HipFemKneTibFibAnk obl cor bon | [22 empty] | 2 | 2 | 101 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | HipFemKneTibFibAnk obl sag bon | [22 empty] | 2 | 2 | 101 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Hips ace bon l090 | [22 empty] | 2 | 2 | 76 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
9100_Kne_sim_wo_cst (Adult)
9100_Kne_sim_wo_cst (Adult) | Topogram PA | Topogram PA 0.6 T20s |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.09 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| Topogram LAT | Topogram LAT 0.6 T20s |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.09 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| Knee simple |  |  | 120 | 120 | 250 |  |  |  | 3 | Off | CARE Dose | 24.04 | 30s | 25% | 60 |  | 1 |  | 2 | 1 |  | 0.6 | 16x0.3mm | [15 empty] | Off
|  | Kne bon l100 :-D | [22 empty] | 0.6 | 0.4 | 587 | U70u sharp | Bone |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
|  | Kne bon l100 | [22 empty] | 2 | 2 | 117 | U70u sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Kne :-D | [22 empty] | 0.6 | 0.4 | 584 | U30u medium smooth | Soft Tissue |  |  |  | ROUTER | GEPACS | [8 empty] | manual
|  | Kne | [22 empty] | 2 | 2 | 117 | U30u medium smooth | Soft Tissue |  |  |  | GEPACS | [9 empty] | manual
|  | Kne obl cor bon | [22 empty] | 2 | 2 | 76 | U30u medium smooth | Soft Tissue |  |  |  | GEPACS | [9 empty] | manual
|  | Kne obl sag bon | [22 empty] | 2 | 2 | 76 | U30u medium smooth | Soft Tissue |  |  |  | GEPACS | [9 empty] | manual
9110_Kne_com_wo_cst (Adult)
9110_Kne_com_wo_cst (Adult) | Topogram PA | Topogram PA 0.6 T20s |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.09 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| Topogram LAT | Topogram LAT 0.6 T20s |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.09 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| Knee complex |  |  | 120 | 120 | 250 |  |  |  | 3 | Off | CARE Dose | 24.04 | 30s | 25% | 60 |  | 1 |  | 2 | 1 |  | 0.6 | 16x0.3mm | [15 empty] | Off
|  | Kne bon l110 :-D | [22 empty] | 0.6 | 0.4 | 587 | U70u sharp | Bone |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
|  | Kne bon l110 | [22 empty] | 2 | 2 | 117 | U70u sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Kne :-D | [22 empty] | 0.6 | 0.4 | 584 | U30u medium smooth | Soft Tissue |  |  |  | ROUTER | GEPACS | [8 empty] | manual
|  | Kne | [22 empty] | 2 | 2 | 117 | U30u medium smooth | Soft Tissue |  |  |  | GEPACS | [9 empty] | manual
|  | Kne obl cor bon | [22 empty] | 2 | 2 | 76 | U30u medium smooth | Soft Tissue |  |  |  | GEPACS | [9 empty] | manual
|  | Kne obl sag bon | [22 empty] | 2 | 2 | 76 | U30u medium smooth | Soft Tissue |  |  |  | GEPACS | [9 empty] | manual
9115_Kne_gou_wo_cst (Adult)
9115_Kne_gou_wo_cst (Adult) | PA Topo | PA Topo 0.6 T20s |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| LAT Topo | LAT Topo 0.6 T20s |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| Kne |  |  | 100 | 124 |  | 250 | 250 |  | 3 | On | CARE Dose | 23.76 | 30s | 25% | 60 |  | 1 |  | 2 | 1.2 |  | 2 | 40x0.6mm | [15 empty] | Off
|  | Kne bon | [22 empty] | 2 | 2 | 120 | B70s very sharp | 2000/500 MSK |  |  |  | GEPACS | [9 empty] | manual
|  | Kne l115 :-D | [22 empty] | 0.6 | 0.3 | 797 | D30s medium smooth | 400/40 MSK/Std |  |  |  | ROUTER | GEPACS | MMWP1 | [7 empty] | manual
|  | Kne bon :-D | [22 empty] | 0.6 | 0.3 | 797 | B70s very sharp | Bone |  |  |  | ROUTER | GEPACS | [8 empty] | manual
9120_Kne_TKA_com_wo_cst (Adult)
9120_Kne_TKA_com_wo_cst (Adult) | Topogram PA | Topogram PA 0.6 T20s |  | 120 |  | 164 |  |  |  | 3 | Off | CARE Dose | 0.24 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| Topogram LAT | Topogram LAT 0.6 T20s |  | 120 |  | 164 |  |  |  | 3 | Off | CARE Dose | 0.24 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| Knee TKA |  | 120 | 140 | 250 | 300 |  |  | Off | 3 | Off | CARE Dose4D | 29.46 | 30s | 25% | 60 |  | 1 |  | 2 | 1 |  | 0.6 | 128x0.6mm | [15 empty] | Off
|  | Kne bon l120 :-D | [22 empty] | 0.6 | 0.6 | 771 | B70s very sharp | Bone |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
|  | Kne bon l120 | [22 empty] | 2 | 2 | 76 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Kne :-D | [22 empty] | 0.6 | 0.6 | 251 | B26s medium smooth ASA | Shoulder |  |  |  | ROUTER | GEPACS | [8 empty] | manual
|  | Kne | [22 empty] | 2 | 2 | 76 | B40s medium | Soft Tissue |  |  |  | GEPACS | [9 empty] | manual
|  | Kne obl cor bon | [22 empty] | 2 | 2 | 101 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Kne obl sag bon | [22 empty] | 2 | 2 | 101 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
9130_KneTibFibAnk_ant_wo_cst (Adult)
9130_KneTibFibAnk_ant_wo_cst (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body |  |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body |  |  |  | GEPACS | ROEMIR001a | TRACKER
| KneTibFibAnk |  |  | 120 | 120 | 240 |  |  |  | 3 | Off | CARE Dose | 23.32 | 30s | 25% | 60 |  | 1 |  | 2 | 0.8 |  | 0.6 | 16x0.3mm | [15 empty] | Off
|  | KneTibFibAnk bon I130 :-D | [22 empty] | 0.6 | 0.4 | 587 | U70u sharp | Bone |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
|  | KneTibFibAnk bon I130 | [22 empty] | 2 | 2 | 117 | U70u sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | KneTibFibAnk :-D | [22 empty] | 0.6 | 0.4 | 584 | U30u medium smooth | Soft Tissue |  |  |  | ROUTER | GEPACS | [8 empty] | manual
|  | KneTibFibAnk | [22 empty] | 2 | 2 | 117 | U30u medium smooth | Soft Tissue |  |  |  | GEPACS | [9 empty] | manual
|  | KneTibFibAnk obl cor bon | [22 empty] | 2 | 2 | 76 | U70u sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | KneTibFibAnk obl sag bon | [22 empty] | 2 | 2 | 76 | U70u sharp | Soft Tissue |  |  |  | GEPACS | [9 empty] | manual
9140_Pel_sim_wo_cst (Adult)
9140_Pel_sim_wo_cst (Adult) | Topogram PA | Topogram PA 0.6 T20s |  | 120 |  | 318 |  |  |  | 3 | Off | CARE Dose | 0.24 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| Topogram LAT | Topogram LAT 0.6 T20s |  | 120 |  | 318 |  |  |  | 3 | Off | CARE Dose | 0.24 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| Pelvis simple |  | 120 | 140 | 250 | 165 |  |  | Off | 3 | On | CARE Dose4D | 24.55 | 30s | 25% | 60 |  | 1 |  | 2 | 1 |  | 0.6 | 128x0.6mm | [15 empty] | Off
|  | Pel bon l140 :-D | [22 empty] | 0.6 | 0.6 | 771 | B70s very sharp | Bone |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
|  | Pel bon l140 | [22 empty] | 2 | 2 | 76 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Pel :-D | [22 empty] | 0.6 | 0.6 | 251 | B26s medium smooth ASA | Shoulder |  |  |  | ROUTER | GEPACS | [8 empty] | manual
|  | Pel | [22 empty] | 2 | 2 | 76 | B40s medium | Soft Tissue |  |  |  | GEPACS | [9 empty] | manual
|  | Pel cor bon | [22 empty] | 2 | 2 | 101 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Pel sag bon | [22 empty] | 2 | 2 | 101 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
9150_Pel_com_wo_cst (Adult)
9150_Pel_com_wo_cst (Adult) | Topogram PA | Topogram PA 0.6 T20s |  | 120 |  | 318 |  |  |  | 3 | Off | CARE Dose | 0.24 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| Topogram LAT | Topogram LAT 0.6 T20s |  | 120 |  | 318 |  |  |  | 3 | Off | CARE Dose | 0.24 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| Pelvis complex |  | 120 | 140 | 250 | 165 |  |  | Off | 3 | On | CARE Dose4D | 24.55 | 30s | 25% | 60 |  | 1 |  | 2 | 1 |  | 0.6 | 128x0.6mm | [15 empty] | Off
|  | Pel bon l150 :-D | [22 empty] | 0.6 | 0.6 | 771 | B70s very sharp | Bone |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
|  | Pel bon l150 | [22 empty] | 2 | 2 | 1078 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Pel :-D | [22 empty] | 0.6 | 0.6 | 3591 | B26s medium smooth ASA | Shoulder |  |  |  | ROUTER | GEPACS | [8 empty] | manual
|  | Pel | [22 empty] | 2 | 2 | 1078 | B40s medium | Soft Tissue |  |  |  | GEPACS | [9 empty] | manual
|  | Pel cor bon | [22 empty] | 2 | 2 | 101 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Pel sag bon | [22 empty] | 2 | 2 | 101 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
9160_PelHip_sim_wo_cst (Adult)
9160_PelHip_sim_wo_cst (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body |  |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body |  |  |  | GEPACS |  | TRACKER
| PelHip sim |  | 120 | 120 | 240 | 165 |  |  | Off | 3 | On | CARE Dose4D | 16.19 | 30s | 25% | 60 |  | 1 |  | 2 | 1 |  | 0.6 | 128x0.6mm | [15 empty] | Off
|  | PelHip bon l160 :-D | [22 empty] | 0.6 | 0.6 | 771 | B70s very sharp | Bone |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
|  | PelHip bon l160 | [22 empty] | 2 | 2 | 76 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | PelHip :-D | [22 empty] | 0.6 | 0.6 | 251 | B26s medium smooth ASA | Shoulder |  |  |  | ROUTER | GEPACS | [8 empty] | manual
|  | PelHip | [22 empty] | 2 | 2 | 76 | B40s medium | Soft Tissue |  |  |  | GEPACS | [9 empty] | manual
|  | PelHip cor bon | [22 empty] | 2 | 2 | 101 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | PelHip sag bon | [22 empty] | 2 | 2 | 101 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
9170_PelHip_com_wo_cst (Adult)
9170_PelHip_com_wo_cst (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body |  |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body |  |  |  | GEPACS |  | TRACKER
| PelHip com |  | 120 | 120 | 240 | 165 |  |  | Off | 3 | On | CARE Dose4D | 16.19 | 30s | 25% | 60 |  | 1 |  | 2 | 1 |  | 0.6 | 128x0.6mm | [15 empty] | Off
|  | PelHip bon l170 :-D | [22 empty] | 0.6 | 0.6 | 771 | B70s very sharp | Bone |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
|  | PelHip bon l170 | [22 empty] | 2 | 2 | 76 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | PelHip :-D | [22 empty] | 0.6 | 0.6 | 251 | B26s medium smooth ASA | Shoulder |  |  |  | ROUTER | GEPACS | [8 empty] | manual
|  | PelHip | [22 empty] | 2 | 2 | 76 | B40s medium | Soft Tissue |  |  |  | GEPACS | [9 empty] | manual
|  | PelHip cor bon | [22 empty] | 2 | 2 | 101 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | PelHip sag bon | [22 empty] | 2 | 2 | 101 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
9175_PelHipFemKne_bio_wo_cst (Adult)
9175_PelHipFemKne_bio_wo_cst (Adult) | Topogram | Topogram |  | 100 |  | 393 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | TRACKER
| Topogram | Topogram |  | 100 |  | 393 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | TRACKER
| PelHipFem |  | 120 | 120 | 240 | 165 |  |  | Off | 3 | On | CARE Dose4D | 16.2 | 30s | 25% | 60 |  | 1 |  | 2 | 0.8 |  | 2 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | PelHipFem l175 | [22 empty] | 2 | 2 | 232 | I40s medium | Soft Tissue |  |  |  | GEPACS | [9 empty] | manual
|  | PelHipFem bon | [22 empty] | 2 | 2 | 232 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | PelHipFem obl cor bon | [22 empty] | 2 | 2 | 181 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | PelHipFem obl sag bon | [22 empty] | 2 | 2 | 181 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | PelHipFem l175 :-D | [22 empty] | 0.6 | 0.6 | 771 | I40s medium | Soft Tissue |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
|  | PelHipFem bon :-D | [22 empty] | 0.6 | 0.6 | 771 | B70s very sharp | Bone |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
| Fem con |  | 120 | 120 | 240 | 165 |  |  | Off | 3 | On | CARE Dose4D | 16.2 | 30s | 25% | 60 |  | 1 |  | 2 | 0.8 |  | 2 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | Fem con l175 | [22 empty] | 2 | 2 | 232 | I40s medium | Soft Tissue |  |  |  | GEPACS | [9 empty] | manual
|  | Fem con l175 :-D | [22 empty] | 0.6 | 0.6 | 771 | I40s medium | Soft Tissue |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
9180_Sac_joi_wo_cst (Adult)
9180_Sac_joi_wo_cst (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body |  |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body |  |  |  | GEPACS |  | TRACKER
| Sac joi |  | 120 | 120 | 240 | 165 |  |  | Off | 3 | On | CARE Dose4D | 16.19 | 30s | 25% | 60 |  | 1 |  | 2 | 1 |  | 0.6 | 128x0.6mm | [15 empty] | Off
|  | Sac joi bon I180 :-D | [22 empty] | 0.6 | 0.6 | 771 | B70s very sharp | Bone |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
|  | Sac joi bon I180 | [22 empty] | 2 | 2 | 76 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Sac joi :-D | [22 empty] | 0.6 | 0.6 | 251 | B26s medium smooth ASA | Shoulder |  |  |  | ROUTER | GEPACS | [8 empty] | manual
|  | Sac joi | [22 empty] | 2 | 2 | 76 | B40s medium | Soft Tissue |  |  |  | GEPACS | [9 empty] | manual
|  | Sac joi obl cor bon | [22 empty] | 2 | 2 | 101 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Sac joi sag bon | [22 empty] | 2 | 2 | 101 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
9185_Sac_joi_gou_wo_cst (Adult)
9185_Sac_joi_gou_wo_cst (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T81f sharp | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T81f sharp | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| DE Sac joi |  |  | 100 | 300 |  | 202 | 156 |  | 3 | On | CARE Dose4D | 23.31 | 30s | 25% | 60 |  | 0.5 |  | 2 | 0.6 |  | 2 | 32x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | Sac joi bon l185 | [22 empty] | 2 | 2 | 132 | B70f very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Sac joi 185 :-D | [22 empty] | 0.6 | 0.6 | 439 | D30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MMWP1 | [7 empty] | manual
|  | Sac joi bon l185 :-D | [22 empty] | 0.6 | 0.6 | 439 | B70f very sharp | Bone |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
|  | Sac joi 185 | [22 empty] | 2 | 2 | 132 | D30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | Sac joi obl cor bon | [22 empty] | 2 | 2 | 181 | B70f very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | Sac joi obl sag bon | [22 empty] | 2 | 2 | 181 | B70f very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
9190_TibFib_sim_wo_cst (Adult)
9190_TibFib_sim_wo_cst (Adult) | Topogram PA | Topogram PA 0.6 T20s |  | 120 |  | 318 |  |  |  | 3 | Off | CARE Dose | 0.24 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| Topogram LAT | Topogram LAT 0.6 T20s |  | 120 |  | 318 |  |  |  | 3 | Off | CARE Dose | 0.24 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| TibFib simple |  | 120 | 140 | 250 | 165 |  |  | Off | 3 | On | CARE Dose4D | 24.55 | 30s | 25% | 60 |  | 1 |  | 2 | 1 |  | 0.6 | 128x0.6mm | [15 empty] | Off
|  | TibFib bon l190 :-D | [22 empty] | 0.6 | 0.6 | 771 | B70s very sharp | Bone |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
|  | TibFib bon l190 | [22 empty] | 2 | 2 | 76 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | TibFib :-D | [22 empty] | 0.6 | 0.6 | 251 | B26s medium smooth ASA | Shoulder |  |  |  | ROUTER | GEPACS | [8 empty] | manual
|  | TibFib | [22 empty] | 2 | 2 | 76 | B40s medium | Soft Tissue |  |  |  | GEPACS | [9 empty] | manual
|  | TibFib obl cor bon | [22 empty] | 2 | 2 | 101 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | TibFib obl sag bon | [22 empty] | 2 | 2 | 101 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
9200_TibFIb_com_wo_cst (Adult)
9200_TibFIb_com_wo_cst (Adult) | Topogram PA | Topogram PA 0.6 T20s |  | 120 |  | 318 |  |  |  | 3 | Off | CARE Dose | 0.24 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| Topogram LAT | Topogram LAT 0.6 T20s |  | 120 |  | 318 |  |  |  | 3 | Off | CARE Dose | 0.24 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER
| TibFib complex |  | 120 | 140 | 250 | 165 |  |  | Off | 3 | On | CARE Dose4D | 24.55 | 30s | 25% | 60 |  | 1 |  | 2 | 1 |  | 0.6 | 128x0.6mm | [15 empty] | Off
|  | TibFib bon l200 :-D | [22 empty] | 0.6 | 0.6 | 771 | B70s very sharp | Bone |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual
|  | TibFib bon l200 | [22 empty] | 2 | 2 | 76 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | TibFib :-D | [22 empty] | 0.6 | 0.6 | 251 | B26s medium smooth ASA | Shoulder |  |  |  | ROUTER | GEPACS | [8 empty] | manual
|  | TibFib | [22 empty] | 2 | 2 | 76 | B40s medium | Soft Tissue |  |  |  | GEPACS | [9 empty] | manual
|  | TibFib obl cor bon | [22 empty] | 2 | 2 | 101 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | TibFib obl sag bon | [22 empty] | 2 | 2 | 101 | B70s very sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
Extremity (Adult)
Extremity (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Extremity |  |  | 120 | 100 | 100 |  |  |  | 3 | On | CARE Dose | 6.75 | 8s | 25% |  |  | 1 |  | 2 | 0.8 |  | 4 | 128x0.6mm | [15 empty] | Off
|  | Extremity 4.0 B60s | [22 empty] | 4 | 4 | 38 | B60s sharp | Extremity | [13 empty] | manual
|  | Extremity 1.0 B60s | [22 empty] | 1 | 0.7 | 215 | B60s sharp | Bone | [13 empty] | manual
ExtrRoutineZUHR (Adult)
ExtrRoutineZUHR (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| ExtrUHR |  |  | 120 | 180 | 180 |  |  |  | 3 | On | CARE Dose | 17.65 | 20s | 25% |  |  | 1 |  | 2 | 0.85 |  | 2 | 16x0.3mm | [15 empty] | Off
|  | ExtrUHR 2.0 U90u | [22 empty] | 2 | 2 | 31 | U90u ultra sharp | Extremity | [13 empty] | manual
|  | ExtrUHR 0.4 U90u | [22 empty] | 0.4 | 0.2 | 301 | U90u ultra sharp | Bone | [13 empty] | manual
ExtrUHR (Adult)
ExtrUHR (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| ExtrUHR |  |  | 120 | 120 | 120 |  |  |  | 3 | On | CARE Dose | 11.77 | 20s | 25% |  |  | 1 |  | 2 | 0.85 |  | 2 | 16x0.6mm | [15 empty] | Off
|  | ExtrUHR 2.0 U90u | [22 empty] | 2 | 2 | 31 | U90u ultra sharp | Extremity | [13 empty] | manual
|  | ExtrUHR 0.6 U90u | [22 empty] | 0.6 | 0.4 | 151 | U90u ultra sharp | Bone | [13 empty] | manual
ExtrZUHR_IR (Adult)
ExtrZUHR_IR (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| ExtrUHR |  |  | 120 | 126 | 126 |  |  |  | 3 | On | CARE Dose | 12.46 | 20s | 25% |  |  | 1 |  | 2 | 0.85 |  | 2 | 16x0.3mm | [15 empty] | Off
|  | ExtrUHR 2.0 V90u 3 | [22 empty] | 2 | 2 | 31 | V90u ultra sharp | Extremity | [13 empty] | manual
|  | ExtrUHR 0.4 V90u 3 | [22 empty] | 0.4 | 0.2 | 301 | V90u ultra sharp | Bone | [13 empty] | manual
FootUHR (Adult)
FootUHR (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| FootUHR |  |  | 120 | 120 | 120 |  |  |  | 3 | On | CARE Dose | 11.77 | 47s | 25% |  |  | 1 |  | 2 | 0.85 |  | 3 | 16x0.6mm | [15 empty] | Off
|  | FootUHR 3.0 U90u | [22 empty] | 3 | 3 | 51 | U90u ultra sharp | Extremity | [13 empty] | manual
|  | FootUHR 0.6 U90u | [22 empty] | 0.6 | 0.4 | 376 | U90u ultra sharp | Bone | [13 empty] | manual
KneeUHR (Adult)
KneeUHR (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| KneeUHR |  |  | 120 | 140 | 140 |  |  |  | 3 | On | CARE Dose | 13.73 | 47s | 25% |  |  | 1 |  | 2 | 0.85 |  | 3 | 16x0.6mm | [15 empty] | Off
|  | KneeUHR 3.0 U90u | [22 empty] | 3 | 3 | 51 | U90u ultra sharp | Extremity | [13 empty] | manual
|  | KneeUHR 0.6 U90u | [22 empty] | 0.6 | 0.4 | 376 | U90u ultra sharp | Bone | [13 empty] | manual
Cardiac
5XXX_ang_Hrt_cor_art_with3CTAModes (Adult)
5XXX_ang_Hrt_cor_art_with3CTAModes (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | ff800867-8571-4a15-997f-1d4ffbea68b7 |  |  | GEPACS | ROEMIR001a | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | ff800867-8571-4a15-997f-1d4ffbea68b7 |  |  | GEPACS | ROEMIR001a | TRACKER
| CaScSeq |  | 120 | 120 | 100 |  | 80 |  | Semi | 3 | On | CARE Dose4D | 3.75 |  | 25% |  |  | 0.285 |  | 2 | 0 | 34 | 3 | 32x1.2mm | [5 empty] | ff800867-8571-4a15-997f-1d4ffbea68b7 | [7 empty] | 70% | 70% | Off | 70% | 70%
|  | Hrt nfov c100 | [22 empty] | 3 |  | 103 | I30f medium smooth | Mediastinum |  |  |  | GEPACS | MCH3DLabAW | [8 empty] | manual | 70%
|  | Hrt | [22 empty] | 3 |  | 103 | I30f medium smooth | Mediastinum |  |  |  | GEPACS | [9 empty] | manual | 70%
| CorCTA_Flash |  | 120 | 100 | 290 |  | 290 |  | On | 9 | On | CARE Dose4D | 3.43 |  | 25% |  |  | 0.285 |  | 21 | 3.4 |  | 0.75 | 128x0.6mm | [5 empty] | ff800867-8571-4a15-997f-1d4ffbea68b7
|  | art Hrt c100 | [22 empty] | 0.75 | 0.5 | 398 | I26f medium smooth ASA | Cardiac |  |  |  | GEPACS | MCH3DLabAW | SYNGO_VIA | [8 empty] | 60%
| Cor_Seq |  | 100 | 100 | 290 |  | 290 |  | Semi | 11 | On | CARE Dose4D | 4.99 |  | 25% |  |  | 0.285 |  | 21 | 0 | 34 | 0.75 | 128x0.6mm | [5 empty] | ff800867-8571-4a15-997f-1d4ffbea68b7 | [7 empty] | 65% | 70% | Normal | 70% | 70%
|  | art Hrt c110 | [22 empty] | 0.75 |  | 413 | I26f medium smooth ASA | Cardiac |  | Best Diastolic |  | GEPACS | MCH3DLabAW | SYNGO_VIA | [7 empty] | BestDiast | 70%
| Cor_Spiral |  | 120 | 120 | 290 |  | 220 |  | On | 9 | On | CARE Dose4D | 42.64 |  | 25% |  |  | 0.285 |  | 21 | 0.23 |  | 0.75 | 128x0.6mm | [5 empty] | ff800867-8571-4a15-997f-1d4ffbea68b7 | [9 empty] | Advance | 70% | 70%
|  | art Hrt c110 | [22 empty] | 0.75 | 0.5 | 401 | I26f medium smooth ASA | Cardiac |  | Best Diastolic |  | MCH3DLabAW | GEPACS | SYNGO_VIA | [7 empty] | BestDiast | 70%
Cardiac_Seq_Mayo_Customized (Adult)
Cardiac_Seq_Mayo_Customized (Adult) | Topogram | Topogram 0.6 T20s |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 7d6f1732-78b9-4d4e-a156-4c4850e203f5 |  |  |  | ROEMIR001a | TRACKER
| DS_CorAdSeq |  | 120 | 100 | 370 |  | 320 |  | Off | 3 | Off | CARE Dose4D | 11.12 |  | 25% |  |  | 0.285 |  | 2 | 0 | 34 | 0.75 | 128x0.6mm | [5 empty] | 7d6f1732-78b9-4d4e-a156-4c4850e203f5 | [7 empty] | 50% | 70% | Normal | 60% | 70%
|  | DS_CorAdSeq 0.75 B26f 65% | [22 empty] | 0.75 |  | 197 | B26f medium smooth ASA | 500/60 Vascular | [13 empty] | manual | 65%
DSXXL_CoronaryCTA (Adult)
DSXXL_CoronaryCTA (Adult) | Topogram | Topogram 0.6 T20f |  | 140 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.2 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
DS_CaScore (Adult)
DS_CaScore (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| DS_CaSc |  | 120 | 120 | 80 |  | 80 |  | Semi | 3 | On | CARE Dose4D | 26.25 |  | 25% |  |  | 0.285 |  | 2 | 0.17 |  | 3 | 128x0.6mm | [15 empty] | Advance | 70% | 70%
|  | DS_CaSc 3.0 B35f BestDiast | [22 empty] | 3 | 1.5 | 101 | B35f HeartView medium CaSco | Mediastinum | [13 empty] | BestDiast | 70%
DS_CaScoreSeq (Adult)
DS_CaScoreSeq (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| DS_CaScSeq |  | 120 | 120 | 80 |  | 80 |  | Semi | 3 | On | CARE Dose4D | 3 |  | 25% |  |  | 0.285 |  | 2 | 0 | 34 | 3 | 32x1.2mm | [13 empty] | 70% | 70% | Off
|  | DS_CaScSeq 3.0 B35f 70% | [22 empty] | 3 |  | 91 | B35f HeartView medium CaSco | Mediastinum | [13 empty] | manual | 70%
DS_CoroCTA_AdaptSeqArrhythmicHR (Adult)
DS_CoroCTA_AdaptSeqArrhythmicHR (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| DS_CorAdSeq |  | 100 | 100 | 366 |  | 366 |  | On | 11 | On | CARE Dose4D | 47.37 |  | 25% |  |  | 0.285 |  | 2 | 0 | 34 | 3 | 128x0.6mm | [13 empty] | 250 ms | 450 ms | Normal | 300 ms | 400 ms
|  | DS_CorAdSeq 3.0 B30f BestSyst | [22 empty] | 3 |  | 46 | B30f medium smooth | Cardiac | [13 empty] | BestSyst | 350 ms
|  | DS_CorAdSeq 0.6 B26f BestSyst | [22 empty] | 0.6 |  | 344 | B26f medium smooth ASA | Cardiac | [13 empty] | BestSyst | 350 ms
DS_CoroCTA_AdaptSeq_Edge (Adult)
DS_CoroCTA_AdaptSeq_Edge (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| DS_CorAdSeq |  | 100 | 100 | 256 |  | 256 |  | On | 11 | On | CARE Dose4D | 33.15 |  | 25% |  |  | 0.285 |  | 2 | 0 | 34 | 3 | 128x0.6mm | [13 empty] | 70% | 70% | Normal | 70% | 70%
|  | DS_CorAdSeq 3.0 I30f 3 BestDiast | [22 empty] | 3 |  | 46 | I30f medium smooth | Cardiac | [13 empty] | BestDiast | 70%
|  | DS_CorAdSeq 0.75 I26f 3 BestDiast | [22 empty] | 0.75 |  | 344 | I26f medium smooth ASA | Cardiac | [13 empty] | BestDiast | 70%
|  | Edge DS_CorAdSeq 0.5 I46f 3 BestDiast | [22 empty] | 0.5 |  | 917 | I46f HeartView sharp ASA | Cardiac | [13 empty] | BestDiast | 70%
DS_CoroCTA_ArrhythmicHR (Adult)
DS_CoroCTA_ArrhythmicHR (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| DS_CorCTA |  | 100 | 100 | 366 |  | 366 |  | On | 11 | On | CARE Dose4D | 71.65 |  | 25% |  |  | 0.285 |  | 2 | 0.17 |  | 3 | 128x0.6mm | [15 empty] | Normal | 300 ms | 400 ms
|  | DS_CorCTA 3.0 B30f BestSyst | [22 empty] | 3 | 3 | 41 | B30f medium smooth | Cardiac | [13 empty] | BestSyst | 350 ms
|  | Coro DS_CorCTA 0.6 B26f BestSyst | [22 empty] | 0.6 | 0.3 | 401 | B26f medium smooth ASA | Cardiac | [13 empty] | BestSyst | 350 ms
DS_CoronaryCTA (Adult)
DS_CoronaryCTA (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| DS_CorCTA |  | 100 | 100 | 366 |  | 366 |  | On | 11 | On | CARE Dose4D | 71.65 |  | 25% |  |  | 0.285 |  | 2 | 0.17 |  | 3 | 128x0.6mm | [15 empty] | Normal | 70% | 70%
|  | DS_CorCTA 3.0 B30f BestDiast | [22 empty] | 3 | 3 | 41 | B30f medium smooth | Cardiac | [13 empty] | BestDiast | 70%
|  | Coro DS_CorCTA 0.6 B26f BestDiast | [22 empty] | 0.6 | 0.3 | 401 | B26f medium smooth ASA | Cardiac | [13 empty] | BestDiast | 70%
|  | Func DS_CorCTA 1.5 B26f 10 - 100 % Matrix 256 | [22 empty] | 1.5 | 1 | 1210 | B26f medium smooth ASA | Cardiac |  | Functional imaging 256 matrix | [11 empty] | manual |  | 10 % 20 % 30 % 40 % 50 % 60 % 70 % 80 % 90 % 100 %
DS_CoronaryCTABiSeg (Adult)
DS_CoronaryCTABiSeg (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| DS_CorCTABi |  | 100 | 100 | 366 |  | 366 |  | On | 11 | On | CARE Dose4D | 71.65 |  | 25% |  |  | 0.285 |  | 2 | 0.17 |  | 3 | 128x0.6mm | [15 empty] | Normal | 70% | 70%
|  | DS_CorCTABi 3.0 B30f BestDiast | [22 empty] | 3 | 3 | 41 | B30f medium smooth | Cardiac | [13 empty] | BestDiast | 70%
|  | Coro DS_CorCTABi 0.6 B26f BestDiast | [22 empty] | 0.6 | 0.3 | 401 | B26f medium smooth ASA | Cardiac | [13 empty] | BestDiast | 70%
|  | Func DS_CorCTABi 1.5 B26f 10 - 100 % Matrix 256 | [22 empty] | 1.5 | 1 | 1210 | B26f medium smooth ASA | Cardiac |  | Functional imaging 256 matrix | [11 empty] | manual |  | 10 % 20 % 30 % 40 % 50 % 60 % 70 % 80 % 90 % 100 %
DS_CoronaryCTA_AdaptSeq (Adult)
DS_CoronaryCTA_AdaptSeq (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| DS_CorAdSeq |  | 100 | 100 | 366 |  | 366 |  | On | 11 | On | CARE Dose4D | 47.37 |  | 25% |  |  | 0.285 |  | 2 | 0 | 34 | 3 | 128x0.6mm | [13 empty] | 70% | 70% | Normal | 70% | 70%
|  | DS_CorAdSeq 3.0 B30f BestDiast | [22 empty] | 3 |  | 46 | B30f medium smooth | Cardiac | [13 empty] | BestDiast | 70%
|  | DS_CorAdSeq 0.6 B26f BestDiast | [22 empty] | 0.6 |  | 344 | B26f medium smooth ASA | Cardiac | [13 empty] | BestDiast | 70%
DS_CoronaryCTA_AdaptSeq_Customized (Adult)
DS_CoronaryCTA_AdaptSeq_Customized (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body |  |  |  |  | ROEMIR001a | TRACKER
| DS_CorAdSeq |  | 120 | 120 | 320 |  | 320 |  | On | 11 | On | CARE Dose4D | 69.17 |  | 25% |  |  | 0.285 |  | 2 | 0 | 34 | 3 | 128x0.6mm | [13 empty] | 70% | 70% | Normal | 70% | 70%
|  | DS_CorAdSeq 3.0 B30f BestDiast | [22 empty] | 3 |  | 46 | B30f medium smooth | Cardiac | [13 empty] | BestDiast | 70%
|  | DS_CorAdSeq 0.6 B26f BestDiast | [22 empty] | 0.6 |  | 344 | B26f medium smooth ASA | Cardiac | [13 empty] | BestDiast | 70%
DS_CoronaryCTA_IR (Adult)
DS_CoronaryCTA_IR (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| DS_CorCTA |  | 100 | 100 | 256 |  | 256 |  | On | 11 | On | CARE Dose4D | 50.14 |  | 25% |  |  | 0.285 |  | 2 | 0.17 |  | 3 | 128x0.6mm | [15 empty] | Normal | 70% | 70%
|  | DS_CorCTA 3.0 I30f 3 BestDiast | [22 empty] | 3 | 3 | 41 | I30f medium smooth | Cardiac | [13 empty] | BestDiast | 70%
|  | Coro DS_CorCTA 0.6 I26f 3 BestDiast | [22 empty] | 0.6 | 0.3 | 401 | I26f medium smooth ASA | Cardiac | [13 empty] | BestDiast | 70%
|  | Func DS_CorCTA 1.5 I26f 3 10 - 100 % Matrix 256 | [22 empty] | 1.5 | 1 | 1210 | I26f medium smooth ASA | Cardiac |  | Functional imaging 256 matrix | [11 empty] | manual |  | 10 % 20 % 30 % 40 % 50 % 60 % 70 % 80 % 90 % 100 %
FAST_Flash_Cardio (Adult)
FAST_Flash_Cardio (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Fl_CaSc |  | 120 | 120 | 80 |  | 80 |  | Semi | 3 | On | CARE Dose4D | 1.31 |  | 25% |  |  | 0.285 |  | 2 | 3.4 |  | 3 | 128x0.6mm
|  | Fl_CaSc 3.0 B35f 70% | [22 empty] | 3 | 1.5 | 91 | B35f HeartView medium CaSco | Mediastinum | [14 empty] | 70%
| Fl_CorCTA |  | 100 | 100 | 366 |  | 366 |  | On | 11 | On | CARE Dose4D | 3.57 |  | 25% |  |  | 0.285 |  | 2 | 3.4 |  | 3 | 128x0.6mm
|  | Fl_CorCTA 3.0 B30f 60% | [22 empty] | 3 | 3 | 46 | B30f medium smooth | Cardiac | [14 empty] | 60%
|  | Fl_CorCTA 0.6 B26f 60% | [22 empty] | 0.6 | 0.3 | 451 | B26f medium smooth ASA | Cardiac | [14 empty] | 60%
| DS_CorAdSeq |  | 100 | 100 | 366 |  | 366 |  | On | 11 | On | CARE Dose4D | 47.37 |  | 25% |  |  | 0.285 |  | 2 | 0 | 34 | 3 | 128x0.6mm | [13 empty] | 70% | 70% | Normal | 70% | 70%
|  | DS_CorAdSeq 3.0 B30f BestDiast | [22 empty] | 3 |  | 46 | B30f medium smooth | Cardiac | [13 empty] | BestDiast | 70%
|  | DS_CorAdSeq 0.6 B26f BestDiast | [22 empty] | 0.6 |  | 344 | B26f medium smooth ASA | Cardiac | [13 empty] | BestDiast | 70%
Flash_Cardio_IR (Adult)
Flash_Cardio_IR (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Fl_CaSc |  | 120 | 120 | 80 |  | 80 |  | Semi | 3 | On | CARE Dose4D | 1.31 |  | 25% |  |  | 0.285 |  | 2 | 3.4 |  | 3 | 128x0.6mm
|  | Fl_CaSc 3.0 B35f 70% | [22 empty] | 3 | 1.5 | 91 | B35f HeartView medium CaSco | Mediastinum | [14 empty] | 70%
| Fl_CorCTA |  | 100 | 100 | 256 |  | 256 |  | On | 11 | On | CARE Dose4D | 2.5 |  | 25% |  |  | 0.285 |  | 2 | 3.4 |  | 3 | 128x0.6mm
|  | Fl_CorCTA 3.0 I30f 3 60% | [22 empty] | 3 | 3 | 46 | I30f medium smooth | Cardiac | [14 empty] | 60%
|  | Fl_CorCTA 0.6 I26f 3 60% | [22 empty] | 0.6 | 0.3 | 451 | I26f medium smooth ASA | Cardiac | [14 empty] | 60%
| DS_CorAdSeq |  | 100 | 100 | 256 |  | 256 |  | On | 11 | On | CARE Dose4D | 33.15 |  | 25% |  |  | 0.285 |  | 2 | 0 | 34 | 3 | 128x0.6mm | [13 empty] | 70% | 70% | Normal | 70% | 70%
|  | DS_CorAdSeq 3.0 I30f 3 BestDiast | [22 empty] | 3 |  | 46 | I30f medium smooth | Cardiac | [13 empty] | BestDiast | 70%
|  | DS_CorAdSeq 0.6 I26f 3 BestDiast | [22 empty] | 0.6 |  | 344 | I26f medium smooth ASA | Cardiac | [13 empty] | BestDiast | 70%
Flash_CaScore (Adult)
Flash_CaScore (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Fl_CaSc |  | 120 | 120 | 80 |  | 80 |  | Semi | 3 | On | CARE Dose4D | 1.31 |  | 25% |  |  | 0.285 |  | 2 | 3.4 |  | 3 | 128x0.6mm
|  | Fl_CaSc 3.0 B35f 70% | [22 empty] | 3 | 1.5 | 91 | B35f HeartView medium CaSco | Mediastinum | [14 empty] | 70%
Vascular
AngioRunOff (Adult)
AngioRunOff (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 365 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| AngioRunOff |  | 120 | 120 | 120 | 120 |  |  | On | 11 | On | CARE Dose4D | 8.11 | 27s | 25% |  |  | 0.5 |  | 2 | 0.6 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | AngioRunOff 5.0 B30f | [22 empty] | 5 | 5 | 201 | B30f medium smooth | Abdomen | [13 empty] | manual
|  | AngioRunOff 1.0 B26f | [22 empty] | 1 | 0.7 | 1429 | B26f medium smooth ASA | CT Angio | [13 empty] | manual
AngioRunOff_IR (Adult)
AngioRunOff_IR (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 365 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| AngioRunOff |  | 120 | 120 | 84 | 84 |  |  | On | 11 | On | CARE Dose4D | 5.69 | 27s | 25% |  |  | 0.5 |  | 2 | 0.6 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | AngioRunOff 5.0 I30f 3 | [22 empty] | 5 | 5 | 201 | I30f medium smooth | Abdomen | [13 empty] | manual
|  | AngioRunOff 1.0 I26f 3 | [22 empty] | 1 | 0.7 | 1429 | I26f medium smooth ASA | CT Angio | [13 empty] | manual
BodyAngioRoutine (Adult)
BodyAngioRoutine (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| BodyAngio |  | 120 | 120 | 120 | 120 |  |  | On | 11 | On | CARE Dose4D | 8.11 | 6s | 25% |  |  | 0.5 |  | 2 | 1.2 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | BodyAngio 5.0 B30f | [22 empty] | 5 | 5 | 81 | B30f medium smooth | Abdomen | [13 empty] | manual
|  | BodyAngio 0.75 B26f | [22 empty] | 0.75 | 0.5 | 801 | B26f medium smooth ASA | CT Angio | [13 empty] | manual
BodyAngio_IR (Adult)
BodyAngio_IR (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| BodyAngio |  | 120 | 120 | 84 | 84 |  |  | On | 11 | On | CARE Dose4D | 5.69 | 6s | 25% |  |  | 0.5 |  | 2 | 1.2 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | BodyAngio 5.0 I30f 3 | [22 empty] | 5 | 5 | 81 | I30f medium smooth | Abdomen | [13 empty] | manual
|  | BodyAngio 0.75 I26f 3 | [22 empty] | 0.75 | 0.5 | 801 | I26f medium smooth ASA | CT Angio | [13 empty] | manual
BodyAngio_Vol (Adult)
BodyAngio_Vol (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| BodyAngio |  | 120 | 120 | 120 | 120 |  |  | On | 11 | On | CARE Dose4D | 8.11 | 6s | 25% |  |  | 0.5 |  | 2 | 1.2 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | BodyAngio 5.0 B30f | [22 empty] | 5 | 5 | 81 | B30f medium smooth | Abdomen | [13 empty] | manual
|  | BodyAngio 3.0 MIP cor | [22 empty] | 3 | 3 | 34 | B30f medium smooth | CT Angio | [13 empty] | manual
CarotidAngioRoutine (Adult)
CarotidAngioRoutine (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| CarotidAngio |  | 120 | 120 | 120 | 120 |  |  | On | 11 | On | CARE Dose4D | 8.12 | 5s | 25% |  |  | 0.285 |  | 2 | 1.2 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | CarotidAngio 5.0 B30f | [22 empty] | 5 | 5 | 61 | B30f medium smooth | Larynx | [13 empty] | manual
|  | CarotidAngio 0.6 B26f | [22 empty] | 0.6 | 0.4 | 751 | B26f medium smooth ASA | CT Angio | [13 empty] | manual
CarotidAngio_IR (Adult)
CarotidAngio_IR (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| CarotidAngio |  | 120 | 120 | 84 | 84 |  |  | On | 11 | On | CARE Dose4D | 5.68 | 5s | 25% |  |  | 0.285 |  | 2 | 1.2 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | CarotidAngio 5.0 I30f 3 | [22 empty] | 5 | 5 | 61 | I30f medium smooth | Larynx | [13 empty] | manual
|  | CarotidAngio 0.6 I26f 3 | [22 empty] | 0.6 | 0.4 | 751 | I26f medium smooth ASA | CT Angio | [13 empty] | manual
CarotidAngio_Vol (Adult)
CarotidAngio_Vol (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| CarotidAngio |  | 120 | 120 | 120 | 120 |  |  | On | 11 | On | CARE Dose4D | 8.12 | 5s | 25% |  |  | 0.285 |  | 2 | 1.2 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | CarotidAngio 5.0 B30f | [22 empty] | 5 | 5 | 61 | B30f medium smooth | Larynx | [13 empty] | manual
|  | CarotidAngio 3.0 MIP cor | [22 empty] | 3 | 3 | 27 | B30f medium smooth | CT Angio | [13 empty] | manual
CarotidDSACT (Adult)
CarotidDSACT (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Non Contrast |  | 120 | 120 | 60 | 60 |  |  | On | 3 | On | CARE Dose4D | 4.06 | 5s | 25% |  |  | 0.285 |  | 2 | 1.2 |  | 0.6 | 128x0.6mm | [15 empty] | Off
|  | DSAnc Non Contrast 0.6 B26f | [22 empty] | 0.6 | 0.4 | 751 | B26f medium smooth ASA | CT Angio | [13 empty] | manual
| CarotidAngio |  | 120 | 120 | 120 | 120 |  |  | On | 11 | On | CARE Dose4D | 8.12 | 5s | 25% |  |  | 0.285 |  | 2 | 1.2 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | CarotidAngio 5.0 B30f | [22 empty] | 5 | 5 | 61 | B30f medium smooth | Larynx | [13 empty] | manual
|  | DSA CarotidAngio 0.6 B26f | [22 empty] | 0.6 | 0.4 | 751 | B26f medium smooth ASA | CT Angio | [13 empty] | manual
DSXXL_ChestPainECG (Adult)
DSXXL_ChestPainECG (Adult) | Topogram | Topogram 0.6 T20f |  | 140 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.2 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
DS_ChestECG_Abd (Adult)
DS_ChestECG_Abd (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 365 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| DS_ChestPain |  | 120 | 120 | 320 |  | 320 |  | On | 11 | On | CARE Dose4D | 104.63 |  | 25% |  |  | 0.285 |  | 2 | 0.17 |  | 3 | 128x0.6mm | [15 empty] | Advance | 70% | 70%
|  | Cor DS_ChestPain 3.0 B30f BestDiast | [22 empty] | 3 | 3 | 51 | B30f medium smooth | Cardiac | [13 empty] | BestDiast | 70%
|  | QCA_D DS_ChestPain 0.6 B26f BestDiast | [22 empty] | 0.6 | 0.3 | 501 | B26f medium smooth ASA | Cardiac | [13 empty] | BestDiast | 70%
|  | QCA_S DS_ChestPain 0.6 B26f BestSyst | [22 empty] | 0.6 | 0.3 | 501 | B26f medium smooth ASA | Cardiac | [13 empty] | BestSyst | 20%
|  | Tho DS_ChestPain 3.0 B30f BestDiast | [22 empty] | 3 | 3 | 84 | B30f medium smooth | Mediastinum | [13 empty] | BestDiast | 70%
|  | Lun DS_ChestPain 3.0 B70f BestDiast | [22 empty] | 3 | 3 | 84 | B70f very sharp | Lung | [13 empty] | BestDiast | 70%
|  | PE DS_ChestPain 1.0 B30f BestDiast | [22 empty] | 1 | 0.5 | 501 | B30f medium smooth | Mediastinum | [13 empty] | BestDiast | 70%
| BodyAngio |  | 120 | 120 | 120 | 120 |  |  | On | 11 | On | CARE Dose4D | 6.66 | 5s | 25% |  |  | 0.285 |  | 3 | 1.2 |  | 3 | 128x0.6mm | [15 empty] | Off
|  | BodyAngio 3.0 B30f | [22 empty] | 3 | 3 | 101 | B30f medium smooth | Abdomen | [13 empty] | manual
|  | BodyAngio 0.6 B26f | [22 empty] | 0.6 | 0.4 | 751 | B26f medium smooth ASA | CT Angio | [13 empty] | manual
DS_ChestPainECG (Adult)
DS_ChestPainECG (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| DS_ChestPain |  | 120 | 120 | 320 |  | 320 |  | On | 11 | On | CARE Dose4D | 104.63 |  | 25% |  |  | 0.285 |  | 2 | 0.17 |  | 3 | 128x0.6mm | [15 empty] | Advance | 70% | 70%
|  | Cor DS_ChestPain 3.0 B30f BestDiast | [22 empty] | 3 | 3 | 51 | B30f medium smooth | Cardiac | [13 empty] | BestDiast | 70%
|  | QCA_D DS_ChestPain 0.6 B26f BestDiast | [22 empty] | 0.6 | 0.3 | 501 | B26f medium smooth ASA | Cardiac | [13 empty] | BestDiast | 70%
|  | QCA_S DS_ChestPain 0.6 B26f BestSyst | [22 empty] | 0.6 | 0.3 | 501 | B26f medium smooth ASA | Cardiac | [13 empty] | BestSyst | 20%
|  | Tho DS_ChestPain 3.0 B30f BestDiast | [22 empty] | 3 | 3 | 84 | B30f medium smooth | Mediastinum | [13 empty] | BestDiast | 70%
|  | Lun DS_ChestPain 3.0 B70f BestDiast | [22 empty] | 3 | 3 | 84 | B70f very sharp | Lung | [13 empty] | BestDiast | 70%
|  | PE DS_ChestPain 1.0 B30f BestDiast | [22 empty] | 1 | 0.5 | 501 | B30f medium smooth | Mediastinum | [13 empty] | BestDiast | 70%
Embolism (Adult)
Embolism (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Embolism |  | 100 | 100 | 150 | 150 |  |  | On | 11 | On | CARE Dose4D | 6.18 | 5s | 25% |  |  | 0.285 |  | 2 | 1.2 |  | 3 | 128x0.6mm | [15 empty] | Off
|  | Embolism 3.0 B30f | [22 empty] | 3 | 3 | 101 | B30f medium smooth | Mediastinum | [13 empty] | manual
|  | Embolism 3.0 B70f | [22 empty] | 3 | 3 | 101 | B70f very sharp | Lung | [13 empty] | manual
|  | Embolism 1.0 B30f | [22 empty] | 1 | 0.5 | 601 | B30f medium smooth | CT Angio | [13 empty] | manual
Embolism_IR (Adult)
Embolism_IR (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Embolism |  | 100 | 100 | 105 | 105 |  |  | On | 11 | On | CARE Dose4D | 4.33 | 5s | 25% |  |  | 0.285 |  | 2 | 1.2 |  | 3 | 128x0.6mm | [15 empty] | Off
|  | Embolism 3.0 I30f 3 | [22 empty] | 3 | 3 | 101 | I30f medium smooth | Mediastinum | [13 empty] | manual
|  | Embolism 3.0 I70f 3 | [22 empty] | 3 | 3 | 101 | I70f very sharp ASA | Lung | [13 empty] | manual
|  | Embolism 1.0 I30f 3 | [22 empty] | 1 | 0.5 | 601 | I30f medium smooth | CT Angio | [13 empty] | manual
Flash_ChestPainECG (Adult)
Flash_ChestPainECG (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Fl_ChestPain |  | 120 | 120 | 320 |  | 320 |  | On | 11 | On | CARE Dose4D | 5.54 |  | 25% |  |  | 0.285 |  | 2 | 3.2 |  | 3 | 128x0.6mm
|  | Cor Fl_ChestPain 3.0 B30f | [22 empty] | 3 | 3 | 51 | B30f medium smooth | Cardiac
|  | QCA Fl_ChestPain 0.6 B26f 60% | [22 empty] | 0.6 | 0.3 | 501 | B26f medium smooth ASA | Cardiac | [14 empty] | 60%
|  | Tho Fl_ChestPain 3.0 B30f 60% | [22 empty] | 3 | 3 | 51 | B30f medium smooth | Mediastinum | [14 empty] | 60%
|  | Lun Fl_ChestPain 3.0 B70f 60% | [22 empty] | 3 | 3 | 84 | B70f very sharp | Lung | [14 empty] | 60%
|  | PE Fl_ChestPain 1.0 B30f 60% | [22 empty] | 1 | 0.5 | 501 | B30f medium smooth | Mediastinum | [14 empty] | 60%
| DS_CorAdSeq |  | 120 | 120 | 320 |  | 320 |  | On | 11 | On | CARE Dose4D | 69.17 |  | 25% |  |  | 0.285 |  | 2 | 0 | 34 | 3 | 128x0.6mm | [13 empty] | 70% | 70% | Normal | 70% | 70%
|  | Cor DS_CorAdSeq 3.0 B30f BestDiast | [22 empty] | 3 |  | 46 | B30f medium smooth | Cardiac | [13 empty] | BestDiast | 70%
|  | Cor DS_CorAdSeq 0.6 B26f BestDiast | [22 empty] | 0.6 |  | 344 | B26f medium smooth ASA | Cardiac | [13 empty] | BestDiast | 70%
Flash_Embolism (Adult)
Flash_Embolism (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Fl_Embolism |  | 100 | 100 | 172 |  | 172 |  | On | 11 | On | CARE Dose4D | 5.58 | 5s | 25% |  |  | 0.285 |  | 2 | 2.2 |  | 3 | 128x0.6mm
|  | Fl_Embolism 3.0 B30f | [22 empty] | 3 | 3 | 101 | B30f medium smooth | Mediastinum
|  | Fl_Embolism 3.0 B70f | [22 empty] | 3 | 3 | 101 | B70f very sharp | Lung
|  | Fl_Embolism 1.0 B30f | [22 empty] | 1 | 0.5 | 601 | B30f medium smooth | CT Angio
Flash_Embolism_IR (Adult)
Flash_Embolism_IR (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Fl_Embolism |  | 100 | 100 | 120 |  | 120 |  | On | 11 | On | CARE Dose4D | 3.98 | 5s | 25% |  |  | 0.285 |  | 2 | 2.2 |  | 3 | 128x0.6mm
|  | Fl_Embolism 3.0 I30f 3 | [22 empty] | 3 | 3 | 101 | I30f medium smooth | Mediastinum
|  | Fl_Embolism 3.0 I70f 3 | [22 empty] | 3 | 3 | 101 | I70f very sharp ASA | Lung
|  | Fl_Embolism 1.0 I30f 3 | [22 empty] | 1 | 0.5 | 601 | I30f medium smooth | CT Angio
Flash_ThorAngio (Adult)
Flash_ThorAngio (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Fl_ThorAngio |  | 100 | 100 | 160 |  | 160 |  | On | 11 | On | CARE Dose4D | 5.3 | 5s | 25% |  |  | 0.285 |  | 2 | 2.2 |  | 5 | 128x0.6mm
|  | Fl_ThorAngio 5.0 B30f | [22 empty] | 5 | 5 | 81 | B30f medium smooth | Mediastinum
|  | Fl_ThorAngio 5.0 B70f | [22 empty] | 5 | 5 | 81 | B70f very sharp | Lung
|  | Fl_ThorAngio 0.75 B26f | [22 empty] | 0.75 | 0.5 | 801 | B26f medium smooth ASA | CT Angio
Flash_ThorAngio_IR (Adult)
Flash_ThorAngio_IR (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Fl_ThorAngio |  | 100 | 100 | 112 |  | 112 |  | On | 11 | On | CARE Dose4D | 3.72 | 5s | 25% |  |  | 0.285 |  | 2 | 2.2 |  | 5 | 128x0.6mm
|  | Fl_ThorAngio 5.0 I30f 3 | [22 empty] | 5 | 5 | 81 | I30f medium smooth | Mediastinum
|  | Fl_ThorAngio 5.0 I70f 3 | [22 empty] | 5 | 5 | 81 | I70f very sharp ASA | Lung
|  | Fl_ThorAngio 0.75 I26f 3 | [22 empty] | 0.75 | 0.5 | 801 | I26f medium smooth ASA | CT Angio
HeadAngioRoutine (Adult)
HeadAngioRoutine (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.29 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head
| HeadAngio |  | 100 | 100 | 175 | 175 |  |  | On | 11 | On | CARE Dose4D | 16.91 | 5s | 25% |  |  | 0.5 |  | 2 | 1.2 |  | 4 | 128x0.6mm | [15 empty] | Off
|  | HeadAngio 4.0 H20f | [22 empty] | 4 | 4 | 31 | H20f smooth | Base Orbita | [13 empty] | manual
|  | HeadAngio 0.6 H10f | [22 empty] | 0.6 | 0.4 | 301 | H10f very smooth | CT Angio | [13 empty] | manual
HeadAngio_IR (Adult)
HeadAngio_IR (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.29 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head
| HeadAngio |  | 100 | 100 | 123 | 123 |  |  | On | 11 | On | CARE Dose4D | 11.92 | 5s | 25% |  |  | 0.5 |  | 2 | 1.2 |  | 4 | 128x0.6mm | [15 empty] | Off
|  | HeadAngio 4.0 J30f 3 | [22 empty] | 4 | 4 | 31 | J30f medium smooth | Base Orbita | [13 empty] | manual
|  | HeadAngio 0.6 J30f 3 | [22 empty] | 0.6 | 0.4 | 301 | J30f medium smooth | CT Angio | [13 empty] | manual
HeadAngio_Vol (Adult)
HeadAngio_Vol (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.29 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head
| HeadAngio |  | 100 | 100 | 175 | 175 |  |  | On | 11 | On | CARE Dose4D | 16.91 | 5s | 25% |  |  | 0.5 |  | 2 | 1.2 |  | 4 | 128x0.6mm | [15 empty] | Off
|  | HeadAngio 4.0 H20f | [22 empty] | 4 | 4 | 31 | H20f smooth | Base Orbita | [13 empty] | manual
|  | HeadAngio 4.0 MIP cor | [22 empty] | 4 | 4 | 21 | H10f very smooth | CT Angio | [13 empty] | manual
HeadAngio_XCARE (Adult)
HeadAngio_XCARE (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.29 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head
| HeadAngio |  | 100 | 100 | 175 | 175 |  |  | On | 11 | On | CARE Dose4D | 16.94 | 5s | 25% |  |  | 0.285 |  | 2 | 0.6 |  | 4 | 128x0.6mm | [15 empty] | Off
|  | HeadAngio 4.0 H20f | [22 empty] | 4 | 4 | 31 | H20f smooth | Base Orbita |  |  | X-CARE | [10 empty] | manual
|  | HeadAngio 0.6 H10f | [22 empty] | 0.6 | 0.4 | 301 | H10f very smooth | CT Angio |  |  | X-CARE | [10 empty] | manual
HeadAngio_XCARE_IR (Adult)
HeadAngio_XCARE_IR (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.29 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head
| HeadAngio |  | 100 | 100 | 123 | 123 |  |  | On | 11 | On | CARE Dose4D | 11.89 | 5s | 25% |  |  | 0.285 |  | 2 | 0.6 |  | 4 | 128x0.6mm | [15 empty] | Off
|  | HeadAngio 4.0 J30f 3 | [22 empty] | 4 | 4 | 31 | J30f medium smooth | Base Orbita |  |  | X-CARE | [10 empty] | manual
|  | HeadAngio 0.6 J30f 3 | [22 empty] | 0.6 | 0.4 | 301 | J30f medium smooth | CT Angio |  |  | X-CARE | [10 empty] | manual
NeuroDSACT (Adult)
NeuroDSACT (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.29 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head
| Non Contrast |  | 100 | 100 | 90 | 90 |  |  | On | 3 | On | CARE Dose4D | 8.7 | 5s | 25% |  |  | 0.5 |  | 2 | 1.2 |  | 0.6 | 128x0.6mm | [15 empty] | Off
|  | DSAnc Non Contrast 0.6 H10f | [22 empty] | 0.6 | 0.4 | 301 | H10f very smooth | CT Angio | [13 empty] | manual
| HeadAngio |  | 100 | 100 | 175 | 175 |  |  | On | 11 | On | CARE Dose4D | 16.91 | 5s | 25% |  |  | 0.5 |  | 2 | 1.2 |  | 4 | 128x0.6mm | [15 empty] | Off
|  | HeadAngio 4.0 H20f | [22 empty] | 4 | 4 | 31 | H20f smooth | Base Orbita | [13 empty] | manual
|  | DSA HeadAngio 0.6 H10f | [22 empty] | 0.6 | 0.4 | 301 | H10f very smooth | CT Angio | [13 empty] | manual
ThorAngioRoutine (Adult)
ThorAngioRoutine (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| ThorAngio |  | 80 | 80 | 224 | 224 |  |  | On | 11 | On | CARE Dose4D | 4.39 | 6s | 25% |  |  | 0.5 |  | 2 | 1.2 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | ThorAngio 5.0 B30f | [22 empty] | 5 | 5 | 81 | B30f medium smooth | Mediastinum | [13 empty] | manual
|  | ThorAngio 5.0 B70f | [22 empty] | 5 | 5 | 81 | B70f very sharp | Lung | [13 empty] | manual
|  | ThorAngio 0.75 B26f | [22 empty] | 0.75 | 0.5 | 801 | B26f medium smooth ASA | CT Angio | [13 empty] | manual
ThorAngio_IR (Adult)
ThorAngio_IR (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| ThorAngio |  | 80 | 80 | 157 | 157 |  |  | On | 11 | On | CARE Dose4D | 3.08 | 6s | 25% |  |  | 0.5 |  | 2 | 1.2 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | ThorAngio 5.0 I30f 3 | [22 empty] | 5 | 5 | 81 | I30f medium smooth | Mediastinum | [13 empty] | manual
|  | ThorAngio 5.0 I70f 3 | [22 empty] | 5 | 5 | 81 | I70f very sharp ASA | Lung | [13 empty] | manual
|  | ThorAngio 0.75 I26f 3 | [22 empty] | 0.75 | 0.5 | 801 | I26f medium smooth ASA | CT Angio | [13 empty] | manual
ThorAngio_Vol (Adult)
ThorAngio_Vol (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| ThorAngio |  | 80 | 80 | 224 | 224 |  |  | On | 11 | On | CARE Dose4D | 4.39 | 6s | 25% |  |  | 0.5 |  | 2 | 1.2 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | ThorAngio 5.0 B30f | [22 empty] | 5 | 5 | 81 | B30f medium smooth | Mediastinum | [13 empty] | manual
|  | ThorAngio 5.0 B70f | [22 empty] | 5 | 5 | 81 | B70f very sharp | Lung | [13 empty] | manual
|  | ThorAngio 3.0 MIP cor | [22 empty] | 3 | 3 | 29 | B30f medium smooth | CT Angio | [13 empty] | manual
ThorAngio_XCARE (Adult)
ThorAngio_XCARE (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| ThorAngio |  | 80 | 80 | 224 | 224 |  |  | On | 11 | On | CARE Dose4D | 4.39 | 6s | 25% |  |  | 0.285 |  | 2 | 0.6 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | ThorAngio 5.0 B30f | [22 empty] | 5 | 5 | 81 | B30f medium smooth | Mediastinum |  |  | X-CARE | [10 empty] | manual
|  | ThorAngio 5.0 B70f | [22 empty] | 5 | 5 | 81 | B70f very sharp | Lung |  |  | X-CARE | [10 empty] | manual
|  | ThorAngio 0.75 B26f | [22 empty] | 0.75 | 0.5 | 801 | B26f medium smooth ASA | CT Angio |  |  | X-CARE | [10 empty] | manual
ThorAngio_XCARE_IR (Adult)
ThorAngio_XCARE_IR (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| ThorAngio |  | 80 | 80 | 157 | 157 |  |  | On | 11 | On | CARE Dose4D | 3.08 | 6s | 25% |  |  | 0.285 |  | 2 | 0.6 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | ThorAngio 5.0 I30f 3 | [22 empty] | 5 | 5 | 81 | I30f medium smooth | Mediastinum |  |  | X-CARE | [10 empty] | manual
|  | ThorAngio 5.0 I70f 3 | [22 empty] | 5 | 5 | 81 | I70f very sharp ASA | Lung |  |  | X-CARE | [10 empty] | manual
|  | ThorAngio 0.75 I26f 3 | [22 empty] | 0.75 | 0.5 | 801 | I26f medium smooth ASA | CT Angio |  |  | X-CARE | [10 empty] | manual
ThorAngio_XCARE_SAFIRE01 (Adult)
ThorAngio_XCARE_SAFIRE01 (Adult) | Topogram | Topogram 0.6 T20f |  | 100 |  | 212 |  |  |  | 3 | Off | CARE Dose | 0.1 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body |  |  |  |  | ROEMIR001a | TRACKER
| Topogram | Topogram 0.6 T20f |  | 100 |  | 212 |  |  |  | 3 | Off | CARE Dose | 0.1 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body |  |  |  |  | ROEMIR001a | TRACKER
| ThorAngio |  | 120 | 100 | 90 | 90 |  |  | On | 8 | On | CARE Dose4D | 4.95 | 7s | 25% |  |  | 0.285 |  | 2 | 0.6 |  | 3 | 128x0.6mm | [15 empty] | Off
|  | ThorAngio 3.0 I30f 3 | [22 empty] | 3 | 2 | 201 | I30f medium smooth | Mediastinum |  |  | X-CARE | GEPACS | [9 empty] | manual
|  | ThorAngio 10.0 MIP cor | [22 empty] | 10 | 10 | 31 | I70f very sharp ASA | Lung |  |  | X-CARE | GEPACS | [9 empty] | manual
|  | ThorAngio 10.0 MIP sag | [22 empty] | 10 | 10 | 31 | I26f medium smooth ASA | CT Angio |  |  | X-CARE | GEPACS | [9 empty] | manual
|  | ThorAngio 0.6 I30f 3 | [22 empty] | 0.6 | 0.3 | 1334 | I30f medium smooth | Mediastinum |  |  | X-CARE | GEPACS | [9 empty] | manual
WholeBodyAngio (Adult)
WholeBodyAngio (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 365 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| WholeBody |  | 120 | 120 | 120 | 120 |  |  | On | 11 | On | CARE Dose4D | 8.11 | 11s | 25% |  |  | 0.5 |  | 2 | 1.2 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | WholeBody 5.0 B30f | [22 empty] | 5 | 5 | 161 | B30f medium smooth | Abdomen | [13 empty] | manual
|  | WholeBody 0.75 B26f | [22 empty] | 0.75 | 0.5 | 1601 | B26f medium smooth ASA | CT Angio | [13 empty] | manual
WholeBodyAngio_IR (Adult)
WholeBodyAngio_IR (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 365 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| WholeBody |  | 120 | 120 | 84 | 84 |  |  | On | 11 | On | CARE Dose4D | 5.69 | 11s | 25% |  |  | 0.5 |  | 2 | 1.2 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | WholeBody 5.0 I30f 3 | [22 empty] | 5 | 5 | 161 | I30f medium smooth | Abdomen | [13 empty] | manual
|  | WholeBody 0.75 I26f 3 | [22 empty] | 0.75 | 0.5 | 1601 | I26f medium smooth ASA | CT Angio | [13 empty] | manual
Private
10_Daily_Quality_Control (Adult)
10_Daily_Quality_Control (Adult) | Topogram LAT | Topogram LAT 0.6 T20s |  | 120 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.2 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body |  |  |  |  | ROEMIR001a | TRACKER
| HeadSeq |  | 120 | 120 | 340 | 400 |  |  | Off | 3 | Off | CARE Dose4D | 68 |  | 25% | 80 |  | 1 |  | 2 | 0 | 34 | 5 | 128x0.6mm | [15 empty] | Off
|  | HeadSeq 5.0 H45s | [22 empty] | 5 |  | 27 | H45s medium | 80/40 Brain Neuro | [13 empty] | manual
|  | HeadSeq 5.0 H70h | [22 empty] | 5 |  | 27 | H70h very sharp | 3700/600 Bone Neuro | [13 empty] | manual
ACRIN_Takahashi_2011_Customized (Adult)
ACRIN_Takahashi_2011_Customized (Adult) | Topogram | Topogram 0.6 T20s |  | 120 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.2 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  |  | ROEMIR001a | TRACKER
| 1st phase 2.5 |  |  | 120 |  | 50 |  |  |  | 11 | Off | CARE Dose | 69 |  | 25% |  |  | 0.285 |  | 2 | 0.5 |  | 5 | 32x1.2mm
|  | VPCT 1st phase 2.5 5.0 B31f | [22 empty] | 5 | 5 | 384 | B31f medium smooth + | 400/40 Abd
| 2nd phas 15 |  |  | 120 |  | 50 |  |  |  | 11 | Off | CARE Dose | 23 |  | 25% |  |  | 0.285 |  | 14 | 0.5 |  | 5 | 32x1.2mm
|  | VPCT 2nd phas 15 5.0 B31f | [22 empty] | 5 | 5 | 128 | B31f medium smooth + | Abdomen
CATPHAN_UHR_LENG_Customized (Adult)
CATPHAN_UHR_LENG_Customized (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body |  |  |  |  | ROEMIR001a
| ExtrUHR |  |  | 120 | 180 | 180 |  |  |  | 5 | On | CARE Dose | 17.65 | 22s | 25% |  |  | 1 |  | 2 | 0.85 |  | 2 | 16x0.3mm | [15 empty] | Off
|  | ExtrUHR 2.0 U90u | [22 empty] | 2 | 2 | 90 | U90u ultra sharp | Extremity | [13 empty] | manual
| ExtrUHR |  |  | 120 | 180 | 1278 |  |  |  | 5 | Off | CARE Dose | 92.31 | 22s | 25% |  |  | 1 |  | 2 | 0.85 |  | 0.4 | 16x0.3mm | [15 empty] | Off
|  | ExtrUHR 0.4 U30u | [22 empty] | 0.4 | 0.1 | 111 | U30u medium smooth | Extremity | [13 empty] | manual
|  | ExtrUHR 0.4 U40u | [22 empty] | 0.4 | 0.1 | 111 | U40u medium | Extremity | [13 empty] | manual
|  | ExtrUHR 0.4 U70u | [22 empty] | 0.4 | 0.1 | 111 | U70u sharp | Extremity | [13 empty] | manual
|  | ExtrUHR 0.4 U70u | [22 empty] | 0.4 | 0.1 | 111 | U70u sharp | Extremity | [13 empty] | manual
|  | ExtrUHR 0.4 U30u | [22 empty] | 0.4 | 0.1 | 111 | U30u medium smooth | Extremity | [13 empty] | manual
|  | ExtrUHR 0.4 U40u | [22 empty] | 0.4 | 0.1 | 111 | U40u medium | Extremity | [13 empty] | manual
|  | ExtrUHR 0.4 U70u | [22 empty] | 0.4 | 0.1 | 111 | U70u sharp | Extremity | [13 empty] | manual
|  | ExtrUHR 0.4 U70u | [22 empty] | 0.4 | 0.1 | 111 | U70u sharp | Extremity | [13 empty] | manual
CATPHAN_UHR_LENG_Customized_Customized (Adult)
CATPHAN_UHR_LENG_Customized_Customized (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body |  |  |  |  | ROEMIR001a
| ExtrUHR |  |  | 120 | 180 | 180 |  |  |  | 5 | On | CARE Dose | 17.65 | 22s | 25% |  |  | 1 |  | 2 | 0.85 |  | 2 | 16x0.3mm | [15 empty] | Off
|  | ExtrUHR 2.0 U90u | [22 empty] | 2 | 2 | 90 | U90u ultra sharp | Extremity | [13 empty] | manual
| ExtrUHR |  |  | 120 | 180 | 1278 |  |  |  | 5 | Off | CARE Dose | 92.31 | 22s | 25% |  |  | 1 |  | 2 | 0.85 |  | 0.4 | 16x0.3mm | [15 empty] | Off
|  | ExtrUHR 0.4 U30u | [22 empty] | 0.4 | 0.1 | 111 | U30u medium smooth | Extremity | [13 empty] | manual
|  | ExtrUHR 0.4 U40u | [22 empty] | 0.4 | 0.1 | 111 | U40u medium | Extremity | [13 empty] | manual
|  | ExtrUHR 0.4 U70u | [22 empty] | 0.4 | 0.1 | 111 | U70u sharp | Extremity | [13 empty] | manual
|  | ExtrUHR 0.4 U70u | [22 empty] | 0.4 | 0.1 | 111 | U70u sharp | Extremity | [13 empty] | manual
|  | ExtrUHR 0.4 U30u | [22 empty] | 0.4 | 0.1 | 111 | U30u medium smooth | Extremity | [13 empty] | manual
|  | ExtrUHR 0.4 U40u | [22 empty] | 0.4 | 0.1 | 111 | U40u medium | Extremity | [13 empty] | manual
|  | ExtrUHR 0.4 U70u | [22 empty] | 0.4 | 0.1 | 111 | U70u sharp | Extremity | [13 empty] | manual
|  | ExtrUHR 0.4 U70u | [22 empty] | 0.4 | 0.1 | 111 | U70u sharp | Extremity | [13 empty] | manual
CATPHAN_UHR_LENG_Customized_Customized1 (Adult)
CATPHAN_UHR_LENG_Customized_Customized1 (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body |  |  |  |  | ROEMIR001a
| ExtrUHR |  |  | 120 | 180 | 180 |  |  |  | 5 | On | CARE Dose | 17.65 | 22s | 25% |  |  | 1 |  | 2 | 0.85 |  | 2 | 16x0.3mm | [15 empty] | Off
|  | ExtrUHR 2.0 U90u | [22 empty] | 2 | 2 | 90 | U90u ultra sharp | Extremity | [13 empty] | manual
| ExtrUHR |  |  | 120 | 180 | 1278 |  |  |  | 5 | Off | CARE Dose | 92.31 | 22s | 25% |  |  | 1 |  | 2 | 0.85 |  | 0.4 | 16x0.3mm | [15 empty] | Off
|  | ExtrUHR 0.4 U30u | [22 empty] | 0.4 | 0.1 | 111 | U30u medium smooth | Extremity | [13 empty] | manual
|  | ExtrUHR 0.4 U40u | [22 empty] | 0.4 | 0.1 | 111 | U40u medium | Extremity | [13 empty] | manual
|  | ExtrUHR 0.4 U70u | [22 empty] | 0.4 | 0.1 | 111 | U70u sharp | Extremity | [13 empty] | manual
|  | ExtrUHR 0.4 U70u | [22 empty] | 0.4 | 0.1 | 111 | U70u sharp | Extremity | [13 empty] | manual
|  | ExtrUHR 0.4 U30u | [22 empty] | 0.4 | 0.1 | 111 | U30u medium smooth | Extremity | [13 empty] | manual
|  | ExtrUHR 0.4 U40u | [22 empty] | 0.4 | 0.1 | 111 | U40u medium | Extremity | [13 empty] | manual
|  | ExtrUHR 0.4 U70u | [22 empty] | 0.4 | 0.1 | 111 | U70u sharp | Extremity | [13 empty] | manual
|  | ExtrUHR 0.4 U70u | [22 empty] | 0.4 | 0.1 | 111 | U70u sharp | Extremity | [13 empty] | manual
CATPHAN_UHR_LENG_Customized_Customized2 (Adult)
CATPHAN_UHR_LENG_Customized_Customized2 (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body |  |  |  |  | ROEMIR001a
| ExtrUHR |  |  | 120 | 180 | 180 |  |  |  | 5 | On | CARE Dose | 17.65 | 22s | 25% |  |  | 1 |  | 2 | 0.85 |  | 2 | 16x0.3mm | [15 empty] | Off
|  | ExtrUHR 2.0 U90u | [22 empty] | 2 | 2 | 90 | U90u ultra sharp | Extremity | [13 empty] | manual
| ExtrUHR |  |  | 120 | 180 | 1278 |  |  |  | 5 | Off | CARE Dose | 92.31 | 22s | 25% |  |  | 1 |  | 2 | 0.85 |  | 0.4 | 16x0.3mm | [15 empty] | Off
|  | ExtrUHR 0.4 U30u | [22 empty] | 0.4 | 0.1 | 111 | U30u medium smooth | Extremity | [13 empty] | manual
|  | ExtrUHR 0.4 U40u | [22 empty] | 0.4 | 0.1 | 111 | U40u medium | Extremity | [13 empty] | manual
|  | ExtrUHR 0.4 U70u | [22 empty] | 0.4 | 0.1 | 111 | U70u sharp | Extremity | [13 empty] | manual
|  | ExtrUHR 0.4 U70u | [22 empty] | 0.4 | 0.1 | 111 | U70u sharp | Extremity | [13 empty] | manual
|  | ExtrUHR 0.4 U30u | [22 empty] | 0.4 | 0.1 | 111 | U30u medium smooth | Extremity | [13 empty] | manual
|  | ExtrUHR 0.4 U40u | [22 empty] | 0.4 | 0.1 | 111 | U40u medium | Extremity | [13 empty] | manual
|  | ExtrUHR 0.4 U70u | [22 empty] | 0.4 | 0.1 | 111 | U70u sharp | Extremity | [13 empty] | manual
|  | ExtrUHR 0.4 U70u | [22 empty] | 0.4 | 0.1 | 111 | U70u sharp | Extremity | [13 empty] | manual
JMK_IMP_H2O_Customized (Adult)
JMK_IMP_H2O_Customized (Adult) | Topogram | Topogram 0.6 T20s |  | 120 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.2 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 0bc31e7d-0da2-46bb-9861-e7c2746fa66a |  |  |  | ROEMIR001a | TRACKER
| H20 1mm |  | 120 | 120 | 340 | 666 |  |  | Off | 3 | Off | CARE Dose4D | 101.72 |  | 25% |  |  | 1 |  | 2 | 0 | 34 | 1 | 128x0.6mm | [15 empty] | Off
|  | H20 1mm 1.0 H30s | [22 empty] | 1 |  | 34 | H30s medium smooth | 350/40 Soft Tissue |  |  |  | IAU | [9 empty] | manual
| H2O 3mm |  | 120 | 120 | 340 | 796 |  |  | Off | 3 | Off | CARE Dose4D | 123.92 |  | 25% |  |  | 1 |  | 2 | 0 | 34 | 3 | 32x1.2mm | [15 empty] | Off
|  | H2O 3mm 3.0 H30s | [22 empty] | 3 |  | 11 | H30s medium smooth | 350/40 Soft Tissue |  |  |  | IAU | [9 empty] | manual
| H2O 5mm |  | 120 | 120 | 340 | 796 |  |  | Off | 3 | Off | CARE Dose4D | 123.92 |  | 25% |  |  | 1 |  | 2 | 0 | 34 | 5 | 32x1.2mm | [15 empty] | Off
|  | H2O 5mm 5.0 H30s | [22 empty] | 5 |  | 6 | H30s medium smooth | 350/40 Soft Tissue |  |  |  | IAU | [9 empty] | manual
| H2O 7mm |  | 120 | 120 | 340 | 796 |  |  | Off | 3 | Off | CARE Dose4D | 123.92 |  | 25% |  |  | 1 |  | 2 | 0 | 34 | 7 | 32x1.2mm | [15 empty] | Off
|  | H2O 7mm 7.0 H30s | [22 empty] | 7 |  | 4 | H30s medium smooth | 350/40 Soft Tissue |  |  |  | IAU | [9 empty] | manual
| H2O 10mm |  | 120 | 120 | 340 | 796 |  |  | Off | 3 | Off | CARE Dose4D | 123.92 |  | 25% |  |  | 1 |  | 2 | 0 | 34 | 10 | 32x1.2mm | [15 empty] | Off
|  | H2O 10mm 10.0 H30s | [22 empty] | 10 |  | 3 | H30s medium smooth | 350/40 Soft Tissue |  |  |  | IAU | [9 empty] | manual
JMK_IMP_Phantom_Customized (Adult)
JMK_IMP_Phantom_Customized (Adult) | Topogram | Topogram 0.6 T20s |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body |  |  |  |  | ROEMIR001a | TRACKER
| Align Front |  | 120 | 120 | 110 | 400 |  |  | Off | 3 | Off | CARE Dose4D | 27.48 |  | 25% |  |  | 1 |  | 2 | 0 | 2 | 1 | 2x1.0mm | [15 empty] | Off
|  | Align Front 1.0 B30s | [22 empty] | 1 |  | 4 | B30s medium smooth | Abdomen | [13 empty] | manual
| Align Back |  | 120 | 120 | 110 | 400 |  |  | Off | 3 | Off | CARE Dose4D | 27.48 |  | 25% |  |  | 1 |  | 2 | 0 | 2 | 1 | 2x1.0mm | [15 empty] | Off
|  | Align Back 1.0 B30s | [22 empty] | 1 |  | 4 | B30s medium smooth | Abdomen | [13 empty] | manual
| CT Number |  | 120 | 120 | 340 | 400 |  |  | Off | 3 | Off | CARE Dose4D | 61.09 |  | 25% |  |  | 1 |  | 2 | 0 | 0 | 3 | 128x0.6mm | [15 empty] | Off
|  | CT Number 3.0 H30s | [22 empty] | 3 |  | 11 | H30s medium smooth | 350/40 Soft Tissue | [13 empty] | manual
| 10mm Images |  | 120 | 120 | 340 | 400 |  |  | Off | 3 | Off | CARE Dose4D | 61.09 |  | 25% |  |  | 1 |  | 2 | 0 | 0 | 10 | 128x0.6mm | [15 empty] | Off
|  | 10mm Images 10.0 H30s | [22 empty] | 10 |  | 5 | H30s medium smooth | 350/40 Soft Tissue | [13 empty] | manual
| Sl Thick 1mm |  | 120 | 120 | 340 | 400 |  |  | Off | 3 | Off | CARE Dose4D | 61.09 |  | 25% |  |  | 1 |  | 2 | 0 | 0 | 1 | 128x0.6mm | [15 empty] | Off
|  | Sl Thick 1.0 H30s | [22 empty] | 1 |  | 34 | H30s medium smooth | 350/40 Soft Tissue | [13 empty] | manual
| Sl Thick 3mm |  | 120 | 120 | 340 | 400 |  |  | Off | 3 | Off | CARE Dose4D | 61.09 |  | 25% |  |  | 1 |  | 2 | 0 | 0 | 3 | 128x0.6mm | [15 empty] | Off
|  | Sl Thick 3.0 H30s | [22 empty] | 3 |  | 11 | H30s medium smooth | 350/40 Soft Tissue | [13 empty] | manual
| Sl Thick 5mm |  | 120 | 120 | 340 | 400 |  |  | Off | 3 | Off | CARE Dose4D | 61.09 |  | 25% |  |  | 1 |  | 2 | 0 | 0 | 5 | 128x0.6mm | [15 empty] | Off
|  | Sl Thick 5.0 H30s | [22 empty] | 5 |  | 6 | H30s medium smooth | 350/40 Soft Tissue | [13 empty] | manual
| Sl Thick 7mm |  | 120 | 120 | 340 | 400 |  |  | Off | 3 | Off | CARE Dose4D | 61.09 |  | 25% |  |  | 1 |  | 2 | 0 | 0 | 7 | 128x0.6mm | [15 empty] | Off
|  | Sl Thick 7.0 H30s | [22 empty] | 7 |  | 4 | H30s medium smooth | 350/40 Soft Tissue | [13 empty] | manual
| LCR 120kV |  | 120 | 120 | 220 | 1400 |  |  | Off | 3 | Off | CARE Dose4D | 86.3 |  | 25% |  |  | 1 |  | 2 | 0 | 10 | 10 | 1x10.0mm | [15 empty] | Off
|  | LCR 120kV 10.0 B30s | [22 empty] | 10 |  | 3 | B30s medium smooth | ACR_100_100 | [13 empty] | manual
| LCR 140kV |  | 120 | 140 | 220 | 1200 |  |  | Off | 3 | Off | CARE Dose4D | 107.71 |  | 25% |  |  | 1 |  | 2 | 0 | 10 | 10 | 1x10.0mm | [15 empty] | Off
|  | LCR 140kV 10.0 B30s | [22 empty] | 10 |  | 3 | B30s medium smooth | ACR_100_100 | [13 empty] | manual
| LCR 120 Abd |  | 120 | 120 | 220 | 240 |  |  | Off | 3 | Off | CARE Dose4D | 14.79 |  | 25% |  |  | 1 |  | 2 | 0 | 10 | 10 | 1x10.0mm | [15 empty] | Off
|  | LCR 120 Abd 10.0 B30s | [22 empty] | 10 |  | 1 | B30s medium smooth | ACR_100_100 | [13 empty] | manual
| LCR 3mm |  | 120 | 120 | 340 | 1332 |  |  | Off | 3 | Off | CARE Dose4D | 203.43 |  | 25% |  |  | 1 |  | 2 | 0 | 0 | 3 | 128x0.6mm | [15 empty] | Off
|  | LCR 3mm H30s | [22 empty] | 3 |  | 11 | H30s medium smooth | 350/40 Soft Tissue | [13 empty] | manual
| Uniform 10s |  | 120 | 120 | 220 | 1400 |  |  | Off | 3 | Off | CARE Dose4D | 86.3 |  | 25% |  |  | 1 |  | 2 | 0 | 10 | 10 | 1x10.0mm | [15 empty] | Off
|  | Uniform 10s 10.0 B30s | [22 empty] | 10 |  | 3 | B30s medium smooth | Abdomen | [13 empty] | manual
| Uniform 3mm |  | 120 | 120 | 340 | 400 |  |  | Off | 3 | Off | CARE Dose4D | 61.09 |  | 25% |  |  | 1 |  | 2 | 0 | 34 | 3 | 128x0.6mm | [15 empty] | Off
|  | Uniform 3mm H30s | [22 empty] | 3 |  | 11 | H30s medium smooth | 350/40 Soft Tissue | [13 empty] | manual
| HCR Sharp |  | 120 | 120 | 110 | 400 |  |  | Off | 3 | Off | CARE Dose4D | 27.48 |  | 25% |  |  | 1 |  | 2 | 0 | 2 | 1 | 2x1.0mm | [15 empty] | Off
|  | HCR Sharp 1.0 B70s | [22 empty] | 1 |  | 10 | B70s very sharp | Abdomen | [13 empty] | manual
|  | HCR Sharp zoom 1.0 B70s | [22 empty] | 1 |  | 10 | B70s very sharp | Abdomen | [13 empty] | manual
| HCR Standard |  | 120 | 120 | 110 | 400 |  |  | Off | 3 | Off | CARE Dose4D | 27.48 |  | 25% |  |  | 1 |  | 2 | 0 | 2 | 2 | 2x1.0mm | [15 empty] | Off
|  | HCR Standard 2.0 B30s | [22 empty] | 2 |  | 5 | B30s medium smooth | Abdomen | [13 empty] | manual
| Whole 3mm |  | 120 | 120 | 340 | 1332 |  |  | Off | 3 | Off | CARE Dose4D | 226.43 |  | 25% |  |  | 1 |  | 2 | 0 | 34 | 3 | 128x0.6mm | [15 empty] | Off
|  | Whole 3mm H30s | [22 empty] | 3 |  | 57 | H30s medium smooth | 350/40 Soft Tissue | [13 empty] | manual
QCT_QA_DE_MRB_Customized (Adult)
QCT_QA_DE_MRB_Customized (Adult) | Mindways |  |  | 100 | 155 |  | 155 | 120 |  | 3 | Off | CARE Dose4D | 12.07 | 30s | 25% |  |  | 0.5 |  | 2 | 0.8 |  | 2 | 32x0.6mm | [5 empty] | 16461837-488a-46f1-8e79-5818aa951afb | [9 empty] | Off
|  | Mindways 2.0 D30f:-D | [22 empty] | 2 | 2 | 10 | D30f medium smooth | 400/40 Abd | [13 empty] | manual
| ESP Phantom |  |  | 100 | 155 |  | 155 | 120 |  | 3 | Off | CARE Dose4D | 12.07 | 30s | 25% |  |  | 0.5 |  | 2 | 0.8 |  | 2 | 32x0.6mm | [5 empty] | 16461837-488a-46f1-8e79-5818aa951afb | [9 empty] | Off
|  | ESP Phantom 2.0 D30f:-D | [22 empty] | 2 | 2 | 56 | D30f medium smooth | 400/40 Abd |  |  |  | A3Z1 | A3Z2 | 3DQCT | [7 empty] | manual
RoutineQC_Flash_ACRtests_etc_Customized (Adult)
RoutineQC_Flash_ACRtests_etc_Customized (Adult) | Topo 180 | Topo 180 0.6 T20s |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body |  |  |  |  | ROEMIR001a | TRACKER
| Table Rep 1 |  | 120 | 120 | 110 | 240 |  |  | Off | 3 | Off | CARE Dose4D | 16.49 |  | 25% |  |  | 1 |  | 2 | 0 | 0 | 1 | 2x1.0mm | [15 empty] | Off
|  | Table Rep 1 1.0 B40s | [22 empty] | 1 |  | 2 | B40s medium | 400/40 Abd | [13 empty] | manual
| Table Rep 121 |  | 120 | 120 | 110 | 240 |  |  | Off | 3 | Off | CARE Dose4D | 16.49 |  | 25% |  |  | 1 |  | 2 | 0 | 0 | 1 | 2x1.0mm | [15 empty] | Off
|  | Table Rep 121 1.0 B40s | [22 empty] | 1 |  | 2 | B40s medium | 400/40 Abd | [13 empty] | manual
| Table Rep 1 |  | 120 | 120 | 110 | 240 |  |  | Off | 3 | Off | CARE Dose4D | 16.49 |  | 25% |  |  | 1 |  | 2 | 0 | 0 | 1 | 2x1.0mm | [15 empty] | Off
|  | Table Rep 1 1.0 B40s | [22 empty] | 1 |  | 2 | B40s medium | 400/40 Abd | [13 empty] | manual
| Abd 128x.6 |  | 120 | 120 | 210 | 240 |  |  | Off | 3 | Off | CARE Dose4D | 16.19 | 30s | 25% |  |  | 0.5 |  | 2 | 1 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | Abd 128x.6 5.0 B40f | [22 empty] | 5 | 40 | 4 | B40f medium | Abdomen | [13 empty] | manual
|  | Abd 128x.6 1.0 B40f | [22 empty] | 1 | 1 | 1 | B40f medium | Abdomen | [13 empty] | manual
| DS_CorCTA |  | 120 | 100 | 320 |  | 350 |  | Off | 3 | Off | CARE Dose4D | 50.54 |  | 25% |  |  | 0.285 |  | 2 | 0.23 |  | 3 | 128x0.6mm | [15 empty] | Normal | 60% | 70%
|  | DS_CorCTA 3.0 B26f BestDiast | [22 empty] | 3 | 40 | 4 | B26f medium smooth ASA | 400/40 Abd | [13 empty] | BestDiast | 64%
| DS_CorCTA |  | 120 | 120 | 320 |  | 250 |  | Off | 3 | Off | CARE Dose4D | 60.23 |  | 25% |  |  | 0.285 |  | 2 | 0.23 |  | 3 | 128x0.6mm | [15 empty] | Normal | 60% | 70%
|  | DS_CorCTA 3.0 B26f BestDiast | [22 empty] | 3 | 3 | 1 | B26f medium smooth ASA | 400/40 Abd | [13 empty] | BestDiast | 67%
| Fl_CorCTA |  | 120 | 120 | 370 |  | 370 |  | Off | 3 | Off | CARE Dose4D | 6.02 |  | 25% |  |  | 0.285 |  | 2 | 3.4 |  | 3 | 128x0.6mm
|  | Fl_CorCTA 3.0 B26f 60% | [22 empty] | 3 | 40 | 4 | B26f medium smooth ASA | 400/40 Abd | [14 empty] | 60%
| DE_Ab100/140 |  |  | 100 | 202 |  | 200 | 155 |  | 3 | Off | CARE Dose4D | 15.54 | 30s | 25% |  |  | 0.5 |  | 2 | 0.6 |  | 5 | 32x0.6mm | [15 empty] | Off
|  | DE_Ab100/140 5.0 D30f | [22 empty] | 5 | 40 | 4 | D30f medium smooth | 400/40 Abd | [13 empty] | manual
| DE_Abd80/140 |  |  | 80 | 341 |  | 425 | 164 |  | 3 | Off | CARE Dose4D | 16.04 | 30s | 25% |  |  | 0.5 |  | 2 | 0.6 |  | 5 | 32x0.6mm | [15 empty] | Off
|  | DE_Abd80/140 5.0 D30f | [22 empty] | 5 | 40 | 4 | D30f medium smooth | 400/40 Abd | [13 empty] | manual
| HeadSeq 1x5 |  | 120 | 120 | 320 | 500 |  |  | Off | 3 | Off | CARE Dose4D | 8.69 |  | 25% |  |  | 1 |  | 2 | 0 | 40 | 5 | 1x5.0mm | [15 empty] | Off
|  | HeadSeq 1x5 5.0 H45s | [22 empty] | 5 |  | 4 | H45s medium | 400/40 Abd | [13 empty] | manual
| AbdSeq 15.3 |  | 120 | 120 | 230 | 240 |  |  | Off | 3 | Off | CARE Dose4D | 16.19 |  | 25% |  |  | 0.5 |  | 2 | 0 | 0 | 3 | 128x0.6mm | [15 empty] | Off
|  | AbdSeq 15.3 3.0 B40f | [22 empty] | 3 |  | 11 | B40f medium | 400/40 Abd | [13 empty] | manual
| KneeUHR 0.3 |  |  | 120 | 140 | 240 |  |  |  | 3 | Off | CARE Dose | 23.08 | 30s | 25% |  |  | 1 |  | 2 | 1 |  | 0.4 | 16x0.3mm | [15 empty] | Off
|  | KneeUHR 0.3 0.4 U40u | [22 empty] | 0.4 | 0.4 | 1 | U40u medium | 400/40 Abd | [13 empty] | manual
| KneeUHR 0.6 |  |  | 120 | 140 | 240 |  |  |  | 3 | Off | CARE Dose | 23.08 | 30s | 25% |  |  | 1 |  | 2 | 1 |  | 0.6 | 16x0.6mm | [15 empty] | Off
|  | KneeUHR 0.6 0.6 U40u | [22 empty] | 0.6 | 0.6 | 1 | U40u medium | 400/40 Abd | [13 empty] | manual
| Slice Pres |  |  | 120 | 140 | 240 |  |  |  | 3 | Off | CARE Dose | 23.08 | 30s | 25% |  |  | 1 |  | 2 | 1 |  | 0.6 | 16x0.6mm | [15 empty] | Off
|  | Slice Pres 0.6 U40u | [22 empty] | 0.6 | 0.6 | 15 | U40u medium | 400/40 Abd | [13 empty] | manual
| Head Uni20x.6 |  | 120 | 120 | 320 | 350 |  |  | Off | 3 | Off | CARE Dose4D | 64.46 | 30s | 25% |  |  | 1 |  | 2 | 0.8 |  | 5 | 20x0.6mm | [15 empty] | Off
|  | Head Uni20x.6 5.0 H41s | [22 empty] | 5 | 5 | 1 | H41s medium + | 400/40 Abd | [13 empty] | manual
| DE_HeadAngio |  |  | 100 | 132 |  | 330 | 330 |  | 11 | Off | CARE Dose4D | 56.73 | 30s | 25% |  |  | 0.5 |  | 2 | 0.7 |  | 5 | 32x0.6mm | [15 empty] | Off
|  | DE_HeadAngio 5.0 H40f | [22 empty] | 5 | 5 | 1 | H40f medium | 400/40 Abd | [13 empty] | manual
|  | DE_HeadAngio 5.0 D20f | [22 empty] | 5 | 5 | 1 | D20f smooth | 400/40 Abd | [13 empty] | manual
| LgUni 128x.6 |  | 120 | 120 | 210 | 200 |  |  | Off | 3 | Off | CARE Dose4D | 13.49 | 30s | 25% |  |  | 0.5 |  | 2 | 1 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | LgUni 128x.6 5.0 B40f | [22 empty] | 5 | 5 | 1 | B40f medium | Abdomen | [13 empty] | manual
| Lg Uni Flash |  | 120 | 120 | 130 |  | 150 |  | Off | 3 | Off | CARE Dose4D | 8.29 | 30s | 25% |  |  | 0.285 |  | 2 | 3 |  | 5 | 128x0.6mm
|  | Lg Uni Flash 5.0 B40f | [22 empty] | 5 | 305 | 1 | B40f medium | 400/40 Abd
| LgUni DE100 |  |  | 100 | 202 |  | 200 | 155 |  | 3 | Off | CARE Dose4D | 15.54 | 30s | 25% |  |  | 0.5 |  | 2 | 0.6 |  | 5 | 32x0.6mm | [15 empty] | Off
|  | LgUni DE100 5.0 D30f | [22 empty] | 5 | 5 | 1 | D30f medium smooth | 400/40 Abd | [13 empty] | manual
| LgUni DE80 |  |  | 80 | 404 |  | 360 | 139 |  | 3 | Off | CARE Dose4D | 13.59 | 30s | 25% |  |  | 0.5 |  | 2 | 0.6 |  | 5 | 32x0.6mm | [15 empty] | Off
|  | LgUni DE80 5.0 D30f | [22 empty] | 5 | 5 | 1 | D30f medium smooth | 400/40 Abd | [13 empty] | manual
| Cross Axial |  | 120 | 120 | 320 | 240 |  |  | Off | 3 | Off | CARE Dose4D | 33.37 |  | 25% |  |  | 1 |  | 2 | 0 | 0 | 5 | 1x5.0mm | [15 empty] | Off
|  | Cross Axial 5.0 H40s | [22 empty] | 5 |  | 1 | H40s medium | 400/40 Abd | [13 empty] | manual
|  | Cross Axial 5.0 H40s | [22 empty] | 5 |  | 1 | H40s medium | 400/40 Abd | [13 empty] | manual
| Cross Topo | Cross Topo 0.6 T20s |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body |  |  |  |  | ROEMIR001a | TRACKER
| Increm Topo | Increm Topo 0.6 T20s |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body |  |  |  |  | ROEMIR001a | TRACKER
| Incrementation |  | 120 | 120 | 110 | 240 |  |  | Off | 3 | Off | CARE Dose4D | 0.66 |  | 25% |  |  | 1 |  | 2 | 0 | 50 | 1 | 2x1.0mm | [15 empty] | Off
|  | Incrementation 1.0 B40s | [22 empty] | 1 |  | 14 | B40s medium | 400/40 Abd | [13 empty] | manual
RoutineQC_Flash_ACR_Align (Adult)
RoutineQC_Flash_ACR_Align (Adult) | ACR Align 3 |  | 120 | 120 | 110 | 240 |  |  | Off | 3 | Off | CARE Dose4D | 16.49 |  | 25% |  |  | 1 |  | 2 | 0 | 2 | 1 | 2x1.0mm | [15 empty] | Off
|  | ACR Align 3 1.0 B40s | [22 empty] | 1 |  | 6 | B40s medium | 400/40 Abd | [13 empty] | manual
| ACR Align 2.5 |  | 120 | 120 | 110 | 240 |  |  | Off | 3 | Off | CARE Dose4D | 16.49 |  | 25% |  |  | 1 |  | 2 | 0 | 2 | 1 | 2x1.0mm | [15 empty] | Off
|  | ACR Align 2.5 1.0 B40s | [22 empty] | 1 |  | 6 | B40s medium | 400/40 Abd | [13 empty] | manual
RoutineQC_Flash_ACR_Align_Customized (Adult)
RoutineQC_Flash_ACR_Align_Customized (Adult) | Topo 180 | Topo 180 0.6 T20s |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body |  |  |  |  | ROEMIR001a | TRACKER
| ACR Align 3 |  | 120 | 120 | 110 | 240 |  |  | Off | 3 | Off | CARE Dose4D | 16.49 |  | 25% |  |  | 1 |  | 2 | 0 | 2 | 1 | 2x1.0mm | [15 empty] | Off
|  | ACR Align 3 1.0 B40s | [22 empty] | 1 |  | 6 | B40s medium | 400/40 Abd | [13 empty] | manual
| ACR Align 2.5 |  | 120 | 120 | 110 | 240 |  |  | Off | 3 | Off | CARE Dose4D | 16.49 |  | 25% |  |  | 1 |  | 2 | 0 | 2 | 1 | 2x1.0mm | [15 empty] | Off
|  | ACR Align 2.5 1.0 B40s | [22 empty] | 1 |  | 6 | B40s medium | 400/40 Abd | [13 empty] | manual
RoutineQC_Flash_AT_Customized (Adult)
RoutineQC_Flash_AT_Customized (Adult) | Topo 180 | Topo 180 0.6 T20s |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body |  |  |  |  | ROEMIR001a | TRACKER
| Abd 128x.6 |  | 120 | 120 | 210 | 240 |  |  | Off | 3 | Off | CARE Dose4D | 16.19 | 30s | 25% |  |  | 0.5 |  | 2 | 1 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | Abd 128x.6 5.0 B40f | [22 empty] | 5 | 40 | 4 | B40f medium | Abdomen | [13 empty] | manual
|  | Abd 128x.6 1.0 B40f | [22 empty] | 1 | 40 | 4 | B40f medium | Abdomen | [13 empty] | manual
| xAbd 128x.6 |  | 120 | 120 | 210 | 240 |  |  | Off | 3 | Off | CARE Dose4D | 16.19 | 30s | 25% |  |  | 0.5 |  | 2 | 1 |  | 0.6 | 128x0.6mm | [15 empty] | Off
|  | xAbd 128x.6 0.6 B40f | [22 empty] | 0.6 | 40 | 4 | B40f medium | Abdomen | [13 empty] | manual
|  | xAbd 128x.6 1.5 B40f | [22 empty] | 1.5 | 40 | 4 | B40f medium | Abdomen | [13 empty] | manual
|  | xAbd 128x.6 2.0 B40f | [22 empty] | 2 | 40 | 4 | B40f medium | Abdomen | [13 empty] | manual
|  | xAbd 128x.6 3.0 B40f | [22 empty] | 3 | 40 | 4 | B40f medium | Abdomen | [13 empty] | manual
| yAbd 128x.6 |  | 120 | 120 | 210 | 240 |  |  | Off | 3 | Off | CARE Dose4D | 16.19 | 30s | 25% |  |  | 0.5 |  | 2 | 0.5 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | yAbd 128x.6 5.0 B40f | [22 empty] | 5 | 40 | 4 | B40f medium | Abdomen | [13 empty] | manual
| zAbd 128x.6 |  | 120 | 120 | 210 | 240 |  |  | Off | 3 | Off | CARE Dose4D | 16.19 | 30s | 25% |  |  | 0.5 |  | 2 | 1.5 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | zAbd 128x.6 5.0 B40f | [22 empty] | 5 | 40 | 4 | B40f medium | Abdomen | [13 empty] | manual
| xAbd 32x1.2 |  | 120 | 120 | 210 | 240 |  |  | Off | 3 | Off | CARE Dose4D | 16.56 | 30s | 25% |  |  | 0.5 |  | 2 | 1 |  | 5 | 32x1.2mm | [15 empty] | Off
|  | xAbd 32x1.2 5.0 B40f | [22 empty] | 5 | 40 | 4 | B40f medium | Abdomen | [13 empty] | manual
| DS_CorCTA |  | 120 | 100 | 320 |  | 350 |  | Off | 3 | Off | CARE Dose4D | 50.54 |  | 25% |  |  | 0.285 |  | 2 | 0.23 |  | 3 | 128x0.6mm | [15 empty] | Normal | 60% | 70%
|  | DS_CorCTA 3.0 B26f BestDiast | [22 empty] | 3 | 40 | 4 | B26f medium smooth ASA | 400/40 Abd | [13 empty] | BestDiast | 65%
| DS_CorCTA |  | 120 | 120 | 320 |  | 250 |  | Off | 3 | Off | CARE Dose4D | 60.23 |  | 25% |  |  | 0.285 |  | 2 | 0.23 |  | 3 | 128x0.6mm | [15 empty] | Normal | 60% | 70%
|  | DS_CorCTA 3.0 B26f BestDiast | [22 empty] | 3 | 40 | 4 | B26f medium smooth ASA | 400/40 Abd | [13 empty] | BestDiast | 62%
| Fl_CorCTA |  | 120 | 120 | 370 |  | 370 |  | Off | 3 | Off | CARE Dose4D | 6.02 |  | 25% |  |  | 0.285 |  | 2 | 3.4 |  | 3 | 128x0.6mm
|  | Fl_CorCTA 3.0 B26f 60% | [22 empty] | 3 | 40 | 4 | B26f medium smooth ASA | 400/40 Abd | [14 empty] | 60%
| DE_Abd140/80 |  |  | 140 | 39 |  | 60 | 330 |  | 3 | Off | CARE Dose4D | 16.18 | 30s | 25% |  |  | 0.5 |  | 2 | 0.6 |  | 5 | 32x0.6mm | [15 empty] | Off
|  | DE_Abd140/80 5.0 D30f | [22 empty] | 5 | 40 | 4 | D30f medium smooth | 400/40 Abd | [13 empty] | manual
| DE_Ab100/140 |  |  | 100 | 202 |  | 200 | 155 |  | 3 | Off | CARE Dose4D | 15.54 | 30s | 25% |  |  | 0.5 |  | 2 | 0.6 |  | 5 | 32x0.6mm | [15 empty] | Off
|  | DE_Ab100/140 5.0 D30f | [22 empty] | 5 | 40 | 4 | D30f medium smooth | 400/40 Abd | [13 empty] | manual
| DE_Abd80/140 |  |  | 80 | 341 |  | 425 | 164 |  | 3 | Off | CARE Dose4D | 16.04 | 30s | 25% |  |  | 0.5 |  | 2 | 0.6 |  | 5 | 32x0.6mm | [15 empty] | Off
|  | DE_Abd80/140 5.0 D30f | [22 empty] | 5 | 40 | 4 | D30f medium smooth | 400/40 Abd | [13 empty] | manual
| ThorHRSeq |  | 120 | 120 | 110 | 240 |  |  | Off | 3 | Off | CARE Dose4D | 16.49 |  | 25% |  |  | 1 |  | 2 | 0 | 0 | 1 | 2x1.0mm | [15 empty] | Off
|  | ThorHRSeq 1.0 B40s | [22 empty] | 1 |  | 2 | B40s medium | 400/40 Abd | [13 empty] | manual
| HeadSeq 1x5 |  | 120 | 120 | 320 | 500 |  |  | Off | 3 | Off | CARE Dose4D | 8.69 |  | 25% |  |  | 1 |  | 2 | 0 | 40 | 5 | 1x5.0mm | [15 empty] | Off
|  | HeadSeq 1x5 5.0 H45s | [22 empty] | 5 |  | 4 | H45s medium | 400/40 Abd | [13 empty] | manual
| AbdSeq 15.3 |  | 120 | 120 | 230 | 240 |  |  | Off | 3 | Off | CARE Dose4D | 16.19 |  | 25% |  |  | 0.5 |  | 2 | 0 | 0 | 3 | 128x0.6mm | [15 empty] | Off
|  | AbdSeq 15.3 3.0 B40f | [22 empty] | 3 |  | 11 | B40f medium | 400/40 Abd | [13 empty] | manual
| xSpineSeq |  | 120 | 120 | 220 | 240 |  |  | Off | 3 | Off | CARE Dose4D | 14.79 |  | 25% |  |  | 1 |  | 2 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | xSpineSeq 10.0 B40s | [22 empty] | 10 |  | 1 | B40s medium | 400/40 Abd | [13 empty] | manual
| xAbdSq32 15.3 |  | 120 | 120 | 230 | 240 |  |  | Off | 3 | Off | CARE Dose4D | 16.56 |  | 25% |  |  | 0.5 |  | 2 | 0 | 0 | 3 | 32x1.2mm | [15 empty] | Off
|  | xAbdSq32 15.3 3.0 B40f | [22 empty] | 3 |  | 11 | B40f medium | 400/40 Abd | [13 empty] | manual
|  | xAbdSq32 15.3 5.0 B40f | [22 empty] | 5 |  | 6 | B40f medium | 400/40 Abd | [13 empty] | manual
| KneeUHR 0.3 |  |  | 120 | 140 | 240 |  |  |  | 3 | Off | CARE Dose | 23.08 | 30s | 25% |  |  | 1 |  | 2 | 1 |  | 0.4 | 16x0.3mm | [15 empty] | Off
|  | KneeUHR 0.3 0.4 U40u | [22 empty] | 0.4 | 0.4 | 1 | U40u medium | 400/40 Abd | [13 empty] | manual
| KneeUHR 0.6 |  |  | 120 | 140 | 240 |  |  |  | 3 | Off | CARE Dose | 23.08 | 30s | 25% |  |  | 1 |  | 2 | 1 |  | 0.6 | 16x0.6mm | [15 empty] | Off
|  | KneeUHR 0.6 0.6 U40u | [22 empty] | 0.6 | 0.6 | 1 | U40u medium | 400/40 Abd | [13 empty] | manual
| Slice Pres |  |  | 120 | 140 | 240 |  |  |  | 3 | Off | CARE Dose | 23.08 | 30s | 25% |  |  | 1 |  | 2 | 1 |  | 0.6 | 16x0.6mm | [15 empty] | Off
|  | Slice Pres 0.6 U40u | [22 empty] | 0.6 | 0.5 | 7 | U40u medium | 400/40 Abd | [13 empty] | manual
| Table Rep 0 |  | 120 | 120 | 110 | 240 |  |  | Off | 3 | Off | CARE Dose4D | 16.49 |  | 25% |  |  | 1 |  | 2 | 0 | 0 | 1 | 2x1.0mm | [15 empty] | Off
|  | Table Rep 0 1.0 B40s | [22 empty] | 1 |  | 2 | B40s medium | 400/40 Abd | [13 empty] | manual
| Table Rep 120 |  | 120 | 120 | 110 | 240 |  |  | Off | 3 | Off | CARE Dose4D | 16.49 |  | 25% |  |  | 1 |  | 2 | 0 | 0 | 1 | 2x1.0mm | [15 empty] | Off
|  | Table Rep 120 1.0 B40s | [22 empty] | 1 |  | 2 | B40s medium | 400/40 Abd | [13 empty] | manual
| Table Rep 0 |  | 120 | 120 | 110 | 240 |  |  | Off | 3 | Off | CARE Dose4D | 16.49 |  | 25% |  |  | 1 |  | 2 | 0 | 0 | 1 | 2x1.0mm | [15 empty] | Off
|  | Table Rep 0 1.0 B40s | [22 empty] | 1 |  | 2 | B40s medium | 400/40 Abd | [13 empty] | manual
| Head Uni20x.6 |  | 120 | 120 | 320 | 350 |  |  | Off | 3 | Off | CARE Dose4D | 64.46 | 30s | 25% |  |  | 1 |  | 2 | 0.8 |  | 5 | 20x0.6mm | [15 empty] | Off
|  | Head Uni20x.6 5.0 H41s | [22 empty] | 5 | 5 | 1 | H41s medium + | 400/40 Abd | [13 empty] | manual
| DE_HeadAngio |  |  | 100 | 132 |  | 330 | 330 |  | 11 | Off | CARE Dose4D | 56.73 | 30s | 25% |  |  | 0.5 |  | 2 | 0.7 |  | 5 | 32x0.6mm | [15 empty] | Off
|  | DE_HeadAngio 5.0 H40f | [22 empty] | 5 | 5 | 1 | H40f medium | 400/40 Abd | [13 empty] | manual
|  | DE_HeadAngio 5.0 D20f | [22 empty] | 5 | 5 | 1 | D20f smooth | 400/40 Abd | [13 empty] | manual
| xSinus |  | 120 | 120 | 80 | 200 |  |  | Off | 11 | Off | CARE Dose4D | 28.17 | 30s | 25% |  |  | 1 |  | 2 | 1 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | xSinus 5.0 H41s | [22 empty] | 5 | 5 | 1 | H41s medium + | 400/40 Abd | [13 empty] | manual
| xHeadSq 63.8 |  | 120 | 120 | 340 | 240 |  |  | Off | 3 | Off | CARE Dose4D | 37.36 |  | 25% |  |  | 1 |  | 2 | 0 | 0 | 3 | 32x1.2mm | [15 empty] | Off
|  | xHeadSq 63.8 3.0 H41s | [22 empty] | 3 |  | 11 | H41s medium + | 400/40 Abd | [13 empty] | manual
| LgUni 128x.6 |  | 120 | 120 | 210 | 200 |  |  | Off | 3 | Off | CARE Dose4D | 13.49 | 30s | 25% |  |  | 0.5 |  | 2 | 1 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | LgUni 128x.6 5.0 B40f | [22 empty] | 5 | 5 | 1 | B40f medium | Abdomen | [13 empty] | manual
| LgUni DE100 |  |  | 100 | 202 |  | 200 | 155 |  | 3 | Off | CARE Dose4D | 15.54 | 30s | 25% |  |  | 0.5 |  | 2 | 0.6 |  | 5 | 32x0.6mm | [15 empty] | Off
|  | LgUni DE100 5.0 D30f | [22 empty] | 5 | 5 | 1 | D30f medium smooth | 400/40 Abd | [13 empty] | manual
RoutineQC_Flash_Gaf_Film_Customized (Adult)
RoutineQC_Flash_Gaf_Film_Customized (Adult) | Ext Laser |  | 120 | 120 | 110 | 300 |  |  | Off | 3 | Off | CARE Dose4D | 20.61 |  | 25% |  |  | 1 |  | 2 | 0 | 0 | 1 | 2x1.0mm | [15 empty] | Off
|  | Ext Laser 1.0 B80s | [22 empty] | 1 |  | 2 | B80s ultra sharp | Lung | [13 empty] | manual
| Int Laser 2x1 |  | 120 | 120 | 110 | 300 |  |  | Off | 3 | Off | CARE Dose4D | 20.61 |  | 25% |  |  | 1 |  | 2 | 0 | 0 | 1 | 2x1.0mm | [15 empty] | Off
|  | Int Laser 2x1 1.0 B80s | [22 empty] | 1 |  | 2 | B80s ultra sharp | Lung | [13 empty] | manual
| HeadSeq 1x5 |  | 120 | 120 | 320 | 300 |  |  | Off | 3 | Off | CARE Dose4D | 41.71 |  | 25% |  |  | 1 |  | 2 | 0 | 0 | 5 | 1x5.0mm | [15 empty] | Off
|  | HeadSeq 1x5 5.0 H31s | [22 empty] | 5 |  | 1 | H31s medium smooth + | Base Orbita | [13 empty] | manual
| AbdSeq 128x.6 |  | 120 | 120 | 230 | 300 |  |  | Off | 3 | Off | CARE Dose4D | 20.24 |  | 25% |  |  | 1 |  | 2 | 0 | 0 | 5 | 128x0.6mm | [15 empty] | Off
|  | AbdSeq 128x.6 5.0 B30s | [22 empty] | 5 |  | 6 | B30s medium smooth | Abdomen | [13 empty] | manual
RoutineQC_Flash_IQ_Dose (Adult)
RoutineQC_Flash_IQ_Dose (Adult) | Topo 180 | Topo 180 0.6 T20f |  | 120 |  | 529 |  |  |  | 3 | Off | CARE Dose | 0.4 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Gaf 2x1 -105 |  | 120 | 120 | 110 | 300 |  |  | Off | 3 | Off | CARE Dose4D | 20.61 |  | 25% |  |  | 1 |  | 2 | 0 | 0 | 1 | 2x1.0mm | [15 empty] | Off
|  | Gaf 2x1 -105 1.0 B40s | [22 empty] | 1 |  | 2 | B40s medium | ACR_400/40 | [13 empty] | manual
| Gaf 128x.6-130 |  | 120 | 120 | 210 | 300 |  |  | Off | 7 | Off | CARE Dose4D | 20.24 |  | 25% |  |  | 1 |  | 2 | 0 | 0 | 5 | 128x0.6mm | [15 empty] | Off
|  | Gaf 128x.6-130 5.0 B30s | [22 empty] | 5 |  | 6 | B30s medium smooth | ACR_400/40 | [13 empty] | manual
| Slice Pres BB |  | 120 | 120 | 210 | 200 |  |  | Off | 7 | Off | CARE Dose4D | 13.52 | 20s | 25% |  |  | 0.5 |  | 2 | 0.6 |  | 0.5 | 128x0.6mm | [15 empty] | Off
|  | Slice Pres BB 0.5 I44f 3 | [22 empty] | 0.5 | 0.5 | 7 | I44f medium | ACR_400/40 | [13 empty] | manual
| x_y Laser -20 |  | 120 | 120 | 210 | 200 |  |  | Off | 7 | Off | CARE Dose4D | 13.52 | 20s | 25% |  |  | 0.5 |  | 2 | 0.6 |  | 10 | 128x0.6mm | [15 empty] | Off
|  | x_y Laser -20 10.0 B40f | [22 empty] | 10 | 10 | 1 | B40f medium | ACR_400/40 | [13 empty] | manual
| Incrment/BB 0 |  | 120 | 120 | 110 | 200 |  |  | Off | 3 | Off | CARE Dose4D | 0.69 |  | 25% |  |  | 0.5 |  | 2 | 0 | 40 | 1 | 2x1.0mm | [15 empty] | Off
|  | Incrment/BB 0 1.0 B40s | [22 empty] | 1 |  | 8 | B40s medium | ACR_400/40 | [13 empty] | manual
| Repositioning 0 |  | 120 | 120 | 110 | 200 |  |  | Off | 3 | Off | CARE Dose4D | 13.74 |  | 25% |  |  | 0.5 |  | 2 | 0 | 0 | 1 | 2x1.0mm | [15 empty] | Off
|  | Repositioning 0 1.0 B40s | [22 empty] | 1 |  | 2 | B40s medium | ACR_400/40 | [13 empty] | manual
| Adult Abd 70 0 |  | 120 | 70 | 200 | 200 |  |  | Off | 8 | Off | CARE Dose4D | 2.49 | 20s | 25% | 50 |  | 0.5 |  | 2 | 0.8 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | Adult Abd 70 0 5.0 I30f 2 | [22 empty] | 5 | 5 | 1 | I30f medium smooth | 400/40 Abd | [13 empty] | manual
| Adult Abd 80 0 |  | 120 | 80 | 200 | 200 |  |  | Off | 8 | Off | CARE Dose4D | 3.91 | 20s | 25% | 50 |  | 0.5 |  | 2 | 0.8 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | Adult Abd 80 0 5.0 I30f 2 | [22 empty] | 5 | 5 | 1 | I30f medium smooth | 400/40 Abd | [13 empty] | manual
| Adult Abd 100 0 |  | 120 | 100 | 200 | 200 |  |  | Off | 8 | Off | CARE Dose4D | 8.23 | 20s | 25% | 50 |  | 0.5 |  | 2 | 0.8 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | Adult Abd 100 0 5.0 I30f 2 | [22 empty] | 5 | 5 | 1 | I30f medium smooth | 400/40 Abd | [13 empty] | manual
| Adult Abd 120 0 |  | 120 | 120 | 200 | 200 |  |  | Off | 8 | Off | CARE Dose4D | 13.5 | 20s | 25% | 50 |  | 0.5 |  | 2 | 0.8 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | Adult Abd 120 0 5.0 I30f 2 | [22 empty] | 5 | 5 | 25 | I30f medium smooth | 400/40 Abd | [13 empty] | manual
|  | Adult Abd 120 0 1.0 I30f 2 | [22 empty] | 1 | 1 | 1 | I30f medium smooth | 400/40 Abd | [13 empty] | manual
| Adult Abd 140 0 |  | 120 | 140 | 200 | 200 |  |  | Off | 8 | Off | CARE Dose4D | 19.65 | 20s | 25% | 50 |  | 0.5 |  | 2 | 0.8 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | Adult Abd 140 0 5.0 I30f 2 | [22 empty] | 5 | 5 | 1 | I30f medium smooth | 400/40 Abd | [13 empty] | manual
| Ad DEAbd100 0 |  |  | 100 | 240 |  | 240 | 185 |  | 3 | Off | CARE Dose4D | 18.61 | 20s | 25% |  |  | 0.5 |  | 2 | 0.6 |  | 3 | 32x0.6mm | [15 empty] | Off
|  | Ad DEAbd100 0 3.0 I30f 3 | [22 empty] | 3 | 40 | 4 | I30f medium smooth | 400/40 Abd | [13 empty] | manual
| Adult Head 0 |  | 120 | 120 | 250 | 250 |  |  | Off | 3 | Off | CARE Dose4D | 38.25 | 15s | 25% |  |  | 1 |  | 2 | 0.6 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | Adult Head 0 5.0 J40s 2 | [22 empty] | 5 | 5 | 25 | J40s medium | ACR_400/40 | [13 empty] | manual
|  | Adult Head 0 5.0 J70h 2 | [22 empty] | 5 | 5 | 1 | J70h very sharp | ACR_100/1055 | [13 empty] | manual
| Peds Head 0 |  | 120 | 120 | 220 | 111 |  |  | Off | 3 | Off | CARE Dose4D | 17.08 | 12s | 25% |  |  | 1 |  | 2 | 0.6 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | Peds Head 0 5.0 J30s 2 | [22 empty] | 5 | 5 | 25 | J30s medium smooth | ACR_400/40 | [13 empty] | manual
| Peds Abd 0 |  | 120 | 120 | 330 | 44 |  |  | Off | 7 | Off | CARE Dose4D | 2.45 | 20s | 25% | 10 |  | 0.285 |  | 2 | 1.4 |  | 3 | 128x0.6mm | [15 empty] | Off
|  | Peds Abd 0 3.0 B40f | [22 empty] | 3 | 3 | 41 | B40f medium | ACR_400/40 | [13 empty] | manual
| Lg Uni 120 -277 |  | 120 | 120 | 200 | 200 |  |  | Off | 8 | Off | CARE Dose4D | 13.5 | 20s | 25% | 50 |  | 0.5 |  | 2 | 0.8 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | Lg Uni 120 -277 5.0 I30f 2 | [22 empty] | 5 | 5 | 1 | I30f medium smooth | ACR_100/100 | [13 empty] | manual
| Lg Uni DE-277 |  |  | 100 | 240 |  | 240 | 185 |  | 3 | Off | CARE Dose4D | 18.61 | 20s | 25% |  |  | 0.5 |  | 2 | 0.6 |  | 3 | 32x0.6mm | [15 empty] | Off
|  | Lg Uni DE -277 3.0 I30f 3 | [22 empty] | 3 | 3 | 1 | I30f medium smooth | ACR_100/100 | [13 empty] | manual
| meter prep -265 |  | 120 | 120 | 210 | 160 |  |  | Off | 7 | Off | CARE Dose4D | 10.79 |  | 25% |  |  | 0.5 |  | 2 | 0 | 0 | 5 | 128x0.6mm | [15 empty] | Off
|  | meter prep -265 5.0 B30f | [22 empty] | 5 |  | 6 | B30f medium smooth | ACR_400/40 | [13 empty] | manual
| Adult Abd -265 |  | 120 | 120 | 210 | 160 |  |  | Off | 7 | Off | CARE Dose4D | 10.79 |  | 25% |  |  | 0.5 |  | 4 | 0 | 0 | 5 | 128x0.6mm | [15 empty] | Off
|  | Adult Abd -265 5.0 B30f | [22 empty] | 5 |  | 6 | B30f medium smooth | ACR_400/40 | [13 empty] | manual
| Adult Abd -265 |  | 120 | 120 | 210 | 160 |  |  | Off | 7 | Off | CARE Dose4D | 10.79 |  | 25% |  |  | 0.5 |  | 4 | 0 | 0 | 5 | 128x0.6mm | [15 empty] | Off
|  | Adult Abd -265 5.0 B30f | [22 empty] | 5 |  | 6 | B30f medium smooth | ACR_400/40 | [13 empty] | manual
| Adult Abd -265 |  | 120 | 120 | 210 | 160 |  |  | Off | 7 | Off | CARE Dose4D | 10.79 |  | 25% |  |  | 0.5 |  | 4 | 0 | 0 | 5 | 128x0.6mm | [15 empty] | Off
|  | Adult Abd -265 5.0 B30f | [22 empty] | 5 |  | 6 | B30f medium smooth | ACR_400/40 | [13 empty] | manual
| Ad DEAbd -273 |  |  | 100 | 171 |  | 144 | 111 |  | 3 | Off | CARE Dose4D | 11.15 |  | 25% |  |  | 0.5 |  | 4 | 0 | 0 | 5 | 32x0.6mm | [15 empty] | Off
|  | Ad DEAbd -273 5.0 B40f | [22 empty] | 5 |  | 3 | B40f medium | Abdomen | [13 empty] | manual
| meter prep -265 |  | 120 | 120 | 210 | 160 |  |  | Off | 7 | Off | CARE Dose4D | 10.79 |  | 25% |  |  | 0.5 |  | 2 | 0 | 0 | 5 | 128x0.6mm | [15 empty] | Off
|  | meter prep -265 5.0 B30f | [22 empty] | 5 |  | 6 | B30f medium smooth | ACR_400/40 | [13 empty] | manual
| Adult Abd -265 |  | 120 | 120 | 210 | 160 |  |  | Off | 7 | Off | CARE Dose4D | 10.79 |  | 25% |  |  | 0.5 |  | 4 | 0 | 0 | 5 | 128x0.6mm | [15 empty] | Off
|  | Adult Abd -265 5.0 B30f | [22 empty] | 5 |  | 6 | B30f medium smooth | ACR_400/40 | [13 empty] | manual
| Adult Abd -265 |  | 120 | 120 | 210 | 160 |  |  | Off | 7 | Off | CARE Dose4D | 10.79 |  | 25% |  |  | 0.5 |  | 4 | 0 | 0 | 5 | 128x0.6mm | [15 empty] | Off
|  | Adult Abd -265 5.0 B30f | [22 empty] | 5 |  | 6 | B30f medium smooth | ACR_400/40 | [13 empty] | manual
| Adult Abd -265 |  | 120 | 120 | 210 | 160 |  |  | Off | 7 | Off | CARE Dose4D | 10.79 |  | 25% |  |  | 0.5 |  | 4 | 0 | 0 | 5 | 128x0.6mm | [15 empty] | Off
|  | Adult Abd -265 5.0 B30f | [22 empty] | 5 |  | 6 | B30f medium smooth | ACR_400/40 | [13 empty] | manual
| Ad DEAbd -273 |  |  | 100 | 171 |  | 144 | 111 |  | 3 | Off | CARE Dose4D | 11.15 |  | 25% |  |  | 0.5 |  | 4 | 0 | 0 | 5 | 32x0.6mm | [15 empty] | Off
|  | Ad DEAbd -273 5.0 B40f | [22 empty] | 5 |  | 3 | B40f medium | Abdomen | [13 empty] | manual
| meter prep -265 |  |  | 120 | 150 | 62 |  |  |  | 11 | Off | CARE Dose | 4.19 |  | 25% |  |  | 0.285 |  | 2 | 0 | 0 | 10 | 128x0.6mm | [15 empty] | Off
|  | meter prep -265 10.0 B40f | [22 empty] | 10 |  | 3 | B40f medium | 400/40 Abd | [13 empty] | manual
| Peds Abd -265 |  |  | 120 | 150 | 62 |  |  |  | 11 | Off | CARE Dose | 4.19 |  | 25% |  |  | 0.285 |  | 4 | 0 | 0 | 10 | 128x0.6mm | [15 empty] | Off
|  | Peds Abd -265 10.0 B40f | [22 empty] | 10 |  | 3 | B40f medium | 400/40 Abd | [13 empty] | manual
| Peds Abd -265 |  |  | 120 | 150 | 62 |  |  |  | 11 | Off | CARE Dose | 4.19 |  | 25% |  |  | 0.285 |  | 4 | 0 | 0 | 10 | 128x0.6mm | [15 empty] | Off
|  | Peds Abd -265 10.0 B40f | [22 empty] | 10 |  | 3 | B40f medium | 400/40 Abd | [13 empty] | manual
| Peds Abd -265 |  |  | 120 | 150 | 62 |  |  |  | 11 | Off | CARE Dose | 4.19 |  | 25% |  |  | 0.285 |  | 4 | 0 | 0 | 10 | 128x0.6mm | [15 empty] | Off
|  | Peds Abd -265 10.0 B40f | [22 empty] | 10 |  | 3 | B40f medium | 400/40 Abd | [13 empty] | manual
| meter prep -265 |  |  | 120 | 150 | 62 |  |  |  | 11 | Off | CARE Dose | 4.19 |  | 25% |  |  | 0.285 |  | 2 | 0 | 0 | 10 | 128x0.6mm | [15 empty] | Off
|  | meter prep -265 10.0 B40f | [22 empty] | 10 |  | 3 | B40f medium | 400/40 Abd | [13 empty] | manual
| Peds Abd -265 |  |  | 120 | 150 | 62 |  |  |  | 11 | Off | CARE Dose | 4.19 |  | 25% |  |  | 0.285 |  | 4 | 0 | 0 | 10 | 128x0.6mm | [15 empty] | Off
|  | Peds Abd -265 10.0 B40f | [22 empty] | 10 |  | 3 | B40f medium | 400/40 Abd | [13 empty] | manual
| Peds Abd -265 |  |  | 120 | 150 | 62 |  |  |  | 11 | Off | CARE Dose | 4.19 |  | 25% |  |  | 0.285 |  | 4 | 0 | 0 | 10 | 128x0.6mm | [15 empty] | Off
|  | Peds Abd -265 10.0 B40f | [22 empty] | 10 |  | 3 | B40f medium | 400/40 Abd | [13 empty] | manual
| Peds Abd -265 |  |  | 120 | 150 | 62 |  |  |  | 11 | Off | CARE Dose | 4.19 |  | 25% |  |  | 0.285 |  | 4 | 0 | 0 | 10 | 128x0.6mm | [15 empty] | Off
|  | Peds Abd -265 10.0 B40f | [22 empty] | 10 |  | 3 | B40f medium | 400/40 Abd | [13 empty] | manual
| meter prep 14.8 |  | 120 | 120 | 340 | 150 |  |  | Off | 3 | Off | CARE Dose4D | 22.91 |  | 25% |  |  | 1 |  | 2 | 0 | 0 | 5 | 128x0.6mm | [15 empty] | Off
|  | meter prep 14.8 5.0 H31s | [22 empty] | 5 |  | 6 | H31s medium smooth + | ACR_400/40 | [13 empty] | manual
| Adult Head 14.8 |  | 120 | 120 | 340 | 150 |  |  | Off | 3 | Off | CARE Dose4D | 22.91 |  | 25% |  |  | 1 |  | 4 | 0 | 0 | 5 | 128x0.6mm | [15 empty] | Off
|  | Adult Head 14.8 5.0 H31s | [22 empty] | 5 |  | 6 | H31s medium smooth + | ACR_400/40 | [13 empty] | manual
| Adult Head 14.8 |  | 120 | 120 | 340 | 150 |  |  | Off | 3 | Off | CARE Dose4D | 22.91 |  | 25% |  |  | 1 |  | 4 | 0 | 0 | 5 | 128x0.6mm | [15 empty] | Off
|  | Adult Head 14.8 5.0 H31s | [22 empty] | 5 |  | 6 | H31s medium smooth + | ACR_400/40 | [13 empty] | manual
| Adult Head 14.8 |  | 120 | 120 | 340 | 150 |  |  | Off | 3 | Off | CARE Dose4D | 22.91 |  | 25% |  |  | 1 |  | 4 | 0 | 0 | 5 | 128x0.6mm | [15 empty] | Off
|  | Adult Head 14.8 5.0 H31s | [22 empty] | 5 |  | 6 | H31s medium smooth + | ACR_400/40 | [13 empty] | manual
| Peds Head 14.8 |  | 120 | 120 | 340 | 67 |  |  | Off | 3 | Off | CARE Dose4D | 10.23 |  | 25% |  |  | 1 |  | 4 | 0 | 0 | 5 | 128x0.6mm | [15 empty] | Off
|  | Peds Head 14.8 5.0 H31s | [22 empty] | 5 |  | 6 | H31s medium smooth + | ACR_400/40 | [13 empty] | manual
| Peds Head 14.8 |  | 120 | 120 | 340 | 67 |  |  | Off | 3 | Off | CARE Dose4D | 10.23 |  | 25% |  |  | 1 |  | 4 | 0 | 0 | 5 | 128x0.6mm | [15 empty] | Off
|  | Peds Head 14.8 5.0 H31s | [22 empty] | 5 |  | 6 | H31s medium smooth + | ACR_400/40 | [13 empty] | manual
| Peds Head 14.8 |  | 120 | 120 | 340 | 67 |  |  | Off | 3 | Off | CARE Dose4D | 10.23 |  | 25% |  |  | 1 |  | 4 | 0 | 0 | 5 | 128x0.6mm | [15 empty] | Off
|  | Peds Head 14.8 5.0 H31s | [22 empty] | 5 |  | 6 | H31s medium smooth + | ACR_400/40 | [13 empty] | manual
| meter prep 14.8 |  | 120 | 120 | 340 | 150 |  |  | Off | 3 | Off | CARE Dose4D | 22.91 |  | 25% |  |  | 1 |  | 2 | 0 | 0 | 5 | 128x0.6mm | [15 empty] | Off
|  | meter prep 14.8 5.0 H31s | [22 empty] | 5 |  | 6 | H31s medium smooth + | ACR_400/40 | [13 empty] | manual
| Adult Head 14.8 |  | 120 | 120 | 340 | 150 |  |  | Off | 3 | Off | CARE Dose4D | 22.91 |  | 25% |  |  | 1 |  | 4 | 0 | 0 | 5 | 128x0.6mm | [15 empty] | Off
|  | Adult Head 14.8 5.0 H31s | [22 empty] | 5 |  | 6 | H31s medium smooth + | ACR_400/40 | [13 empty] | manual
| Adult Head 14.8 |  | 120 | 120 | 340 | 150 |  |  | Off | 3 | Off | CARE Dose4D | 22.91 |  | 25% |  |  | 1 |  | 4 | 0 | 0 | 5 | 128x0.6mm | [15 empty] | Off
|  | Adult Head 14.8 5.0 H31s | [22 empty] | 5 |  | 6 | H31s medium smooth + | ACR_400/40 | [13 empty] | manual
| Adult Head 14.8 |  | 120 | 120 | 340 | 150 |  |  | Off | 3 | Off | CARE Dose4D | 22.91 |  | 25% |  |  | 1 |  | 4 | 0 | 0 | 5 | 128x0.6mm | [15 empty] | Off
|  | Adult Head 14.8 5.0 H31s | [22 empty] | 5 |  | 6 | H31s medium smooth + | ACR_400/40 | [13 empty] | manual
| Peds Head 14.8 |  | 120 | 120 | 340 | 67 |  |  | Off | 3 | Off | CARE Dose4D | 10.23 |  | 25% |  |  | 1 |  | 4 | 0 | 0 | 5 | 128x0.6mm | [15 empty] | Off
|  | Peds Head 14.8 5.0 H31s | [22 empty] | 5 |  | 6 | H31s medium smooth + | ACR_400/40 | [13 empty] | manual
| Peds Head 14.8 |  | 120 | 120 | 340 | 67 |  |  | Off | 3 | Off | CARE Dose4D | 10.23 |  | 25% |  |  | 1 |  | 4 | 0 | 0 | 5 | 128x0.6mm | [15 empty] | Off
|  | Peds Head 14.8 5.0 H31s | [22 empty] | 5 |  | 6 | H31s medium smooth + | ACR_400/40 | [13 empty] | manual
| Peds Head 14.8 |  | 120 | 120 | 340 | 67 |  |  | Off | 3 | Off | CARE Dose4D | 10.23 |  | 25% |  |  | 1 |  | 4 | 0 | 0 | 5 | 128x0.6mm | [15 empty] | Off
|  | Peds Head 14.8 5.0 H31s | [22 empty] | 5 |  | 6 | H31s medium smooth + | ACR_400/40 | [13 empty] | manual
RoutineQC_Flash_SpotCheck_Customized (Adult)
RoutineQC_Flash_SpotCheck_Customized (Adult) | Topo 180 | Topo 180 0.6 T20s |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body |  |  |  |  | ROEMIR001a | TRACKER
| Abd 128x.6 |  | 120 | 120 | 210 | 240 |  |  | Off | 3 | Off | CARE Dose4D | 16.19 | 30s | 25% |  |  | 0.5 |  | 2 | 1 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | Abd 128x.6 5.0 B40f | [22 empty] | 5 | 40 | 4 | B40f medium | Abdomen | [13 empty] | manual
|  | Abd 128x.6 1.0 B40f | [22 empty] | 1 | 40 | 4 | B40f medium | Abdomen | [13 empty] | manual
| DS_CorCTA |  | 120 | 120 | 320 |  | 250 |  | Off | 3 | Off | CARE Dose4D | 60.23 |  | 25% |  |  | 0.285 |  | 2 | 0.23 |  | 3 | 128x0.6mm | [15 empty] | Normal | 60% | 70%
|  | DS_CorCTA 3.0 B26f BestDiast | [22 empty] | 3 | 40 | 4 | B26f medium smooth ASA | 400/40 Abd | [13 empty] | BestDiast | 66%
| Fl_CorCTA |  | 120 | 120 | 370 |  | 370 |  | Off | 3 | Off | CARE Dose4D | 6.02 |  | 25% |  |  | 0.285 |  | 2 | 3.4 |  | 3 | 128x0.6mm
|  | Fl_CorCTA 3.0 B26f 60% | [22 empty] | 3 | 40 | 4 | B26f medium smooth ASA | 400/40 Abd | [14 empty] | 60%
| DE_Abdomen |  |  | 100 | 202 |  | 200 | 155 |  | 3 | Off | CARE Dose4D | 15.54 | 30s | 25% |  |  | 0.5 |  | 2 | 0.6 |  | 5 | 32x0.6mm | [15 empty] | Off
|  | DE_Abdomen 5.0 D30f | [22 empty] | 5 | 40 | 4 | D30f medium smooth | 400/40 Abd | [13 empty] | manual
| Head 20x.6 |  | 120 | 120 | 320 | 425 |  |  | Off | 3 | Off | CARE Dose4D | 78.27 | 30s | 25% |  |  | 1 |  | 2 | 0.8 |  | 5 | 20x0.6mm | [15 empty] | Off
|  | Head 20x.6 5.0 H41s | [22 empty] | 5 | 40 | 4 | H41s medium + | 400/40 Abd | [13 empty] | manual
| LgUni 128x.6 |  | 120 | 120 | 210 | 200 |  |  | Off | 3 | Off | CARE Dose4D | 13.49 | 30s | 25% |  |  | 0.5 |  | 2 | 1 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | LgUni 128x.6 5.0 B40f | [22 empty] | 5 | 5 | 1 | B40f medium | Abdomen | [13 empty] | manual
| LgUni DE100 |  |  | 100 | 202 |  | 200 | 155 |  | 3 | Off | CARE Dose4D | 15.54 | 30s | 25% |  |  | 0.5 |  | 2 | 0.6 |  | 5 | 32x0.6mm | [15 empty] | Off
|  | LgUni DE100 5.0 D30f | [22 empty] | 5 | 5 | 1 | D30f medium smooth | 400/40 Abd | [13 empty] | manual
WIP_Flash_CAP_Customized (Adult)
WIP_Flash_CAP_Customized (Adult) | Topogram | Topogram 0.6 T20s |  | 120 |  | 275 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body |  |  |  | GEPACS | ROEMIR001a | TRACKER
| Fl_CAP |  | 120 | 120 | 210 |  | 130 |  | Off | 3 | On | CARE Dose4D | 7.88 | 30s | 25% | 50 |  | 0.285 |  | 65 | 3.2 |  | 5 | 128x0.6mm
|  | Fl_CAP 5.0 B40f | [22 empty] | 5 | 5 | 99 | B40f medium | 400/40 Abd |  | Nonionic
|  | Fl_COR AP | [22 empty] | 3 | 2.5 | 133 | B40f medium | 400/40 Abd |  | Nonionic
|  | Fl_HR 1.5 B46f | [22 empty] | 1.5 | 1.5 | 330 | B46f HeartView sharp ASA | 1500/600 Lung |  | Nonionic
|  | Fl_Chest 20.0 MIP | [22 empty] | 20 | 10 | 67 | B50f medium sharp | 1500/600 Lung |  | Nonionic
|  | Fl_COR AP (OPT) | [22 empty] | 3 | 2.5 | 153 | B40f medium | 400/40 Abd |  | Nonionic
YZ_120mAs_Customized (Adult)
YZ_120mAs_Customized (Adult) | Topogram | Topogram 0.6 T20s |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd |  |  |  |  | ROEMIR001a | TRACKER
| 120mAs |  | 120 | 120 | 120 | 480 |  |  | Off | 8 | On | CARE Dose4D | 8.1 | 30s | 25% | 50 |  | 0.5 |  | 2 | 0.75 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | 120mAs 5.0 B40f | [22 empty] | 5 | 5 | 15 | B40f medium | 400/40 Abd | [13 empty] | manual
|  | B40f | [22 empty] | 5 | 5 | 15 | B40f medium | 400/40 Abd | [13 empty] | manual
|  | I40f | [22 empty] | 5 | 5 | 15 | I40f medium | 400/40 Abd | [13 empty] | manual
YZ_180mAs_Customized (Adult)
YZ_180mAs_Customized (Adult) | Topogram | Topogram 0.6 T20s |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd |  |  |  |  | ROEMIR001a | TRACKER
| 180mAs |  | 120 | 120 | 180 | 480 |  |  | Off | 8 | On | CARE Dose4D | 12.14 | 30s | 25% | 50 |  | 0.5 |  | 2 | 0.75 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | 180mAs 5.0 B40f | [22 empty] | 5 | 5 | 15 | B40f medium | 400/40 Abd | [13 empty] | manual
|  | B40f | [22 empty] | 5 | 5 | 15 | B40f medium | 400/40 Abd | [13 empty] | manual
|  | I40f | [22 empty] | 5 | 5 | 15 | I40f medium | 400/40 Abd | [13 empty] | manual
YZ_240mAs_Customized (Adult)
YZ_240mAs_Customized (Adult) | Topogram | Topogram 0.6 T20s |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd |  |  |  |  | ROEMIR001a | TRACKER
| 240mAs |  | 120 | 120 | 240 | 480 |  |  | Off | 8 | On | CARE Dose4D | 16.19 | 30s | 25% | 50 |  | 0.5 |  | 2 | 0.75 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | 240mAs 5.0 B40f | [22 empty] | 5 | 5 | 15 | B40f medium | 400/40 Abd | [13 empty] | manual
|  | B40f | [22 empty] | 5 | 5 | 15 | B40f medium | 400/40 Abd | [13 empty] | manual
|  | I40f | [22 empty] | 5 | 5 | 15 | I40f medium | 400/40 Abd | [13 empty] | manual
YZ_360mAs_Customized (Adult)
YZ_360mAs_Customized (Adult) | Topogram | Topogram 0.6 T20s |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd |  |  |  |  | ROEMIR001a | TRACKER
| 360mA |  | 120 | 120 | 360 | 480 |  |  | Off | 8 | On | CARE Dose4D | 24.29 | 30s | 25% | 50 |  | 0.5 |  | 2 | 0.75 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | 360mA 5.0 B40f | [22 empty] | 5 | 5 | 15 | B40f medium | 400/40 Abd | [13 empty] | manual
|  | B40f | [22 empty] | 5 | 5 | 15 | B40f medium | 400/40 Abd | [13 empty] | manual
|  | I40f | [22 empty] | 5 | 5 | 15 | I40f medium | 400/40 Abd | [13 empty] | manual
YZ_480mAs_Customized (Adult)
YZ_480mAs_Customized (Adult) | Topogram | Topogram 0.6 T20s |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd |  |  |  |  | ROEMIR001a | TRACKER
| 480mAs |  | 120 | 120 | 480 | 480 |  |  | Off | 8 | On | CARE Dose4D | 32.38 | 30s | 25% | 50 |  | 0.5 |  | 2 | 0.75 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | 480mAs 5.0 B40f | [22 empty] | 5 | 5 | 15 | B40f medium | 400/40 Abd | [13 empty] | manual
|  | B40f | [22 empty] | 5 | 5 | 15 | B40f medium | 400/40 Abd | [13 empty] | manual
|  | I40f | [22 empty] | 5 | 5 | 15 | I40f medium | 400/40 Abd | [13 empty] | manual
|  | B40f thin | [22 empty] | 1 | 1 | 31 | B40f medium | 400/40 Abd | [13 empty] | manual
YZ_test_Customized (Adult)
YZ_test_Customized (Adult) | Topogram | Topogram 0.6 T20s |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd |  |  |  |  | ROEMIR001a | TRACKER
| 120mAs |  | 120 | 120 | 120 | 480 |  |  | Off | 8 | On | CARE Dose4D | 8.1 | 30s | 25% | 50 |  | 0.5 |  | 2 | 0.75 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | I40f | [22 empty] | 5 | 5 | 15 | I40f medium | 400/40 Abd | [13 empty] | manual
|  | 3D oblique | [22 empty] | 5 | 5 | 15 | B40f medium | 400/40 Abd | [13 empty] | manual
|  | ax 3D | [22 empty] | 5 | 5 | 15 | B40f medium | 400/40 Abd | [13 empty] | manual
|  | con 3D | [22 empty] | 5 | 5 | 81 | B40f medium | 400/40 Abd | [13 empty] | manual
|  | spine oblique 3D | [22 empty] | 5 | 5 | 81 | B40f medium | 400/40 Abd | [13 empty] | manual
|  | spine disc 3D | [22 empty] | 5 | 5 | 81 | B40f medium | 400/40 Abd | [13 empty] | manual
|  | spine vertebra 3D | [22 empty] | 5 | 5 | 81 | B40f medium | 400/40 Abd | [13 empty] | manual
|  | spine multi 3D | [22 empty] | 5 | 5 | 0 | B40f medium | 400/40 Abd | [13 empty] | manual
YZ_UHRtest_Customized (Adult)
YZ_UHRtest_Customized (Adult) | Topogram PA | Topogram PA 0.6 T80s |  | 120 |  | 318 |  |  |  | 3 | Off | CARE Dose | 0.24 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Body |  |  |  |  | ROEMIR001a | TRACKER
| Topogram PA | Topogram PA 0.6 T80s |  | 120 |  | 318 |  |  |  | 3 | Off | CARE Dose | 0.24 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Body |  |  |  |  | ROEMIR001a | TRACKER
| Pelvis |  | 140 | 140 | 250 | 250 |  |  | Off | 2 | Off | CARE Dose4D | 24.55 | 30s | 25% | 50 |  | 1 |  | 2 | 1 |  | 2 | 128x0.6mm | [15 empty] | Off
|  | Pelvis 2.0 B70s | [22 empty] | 2 | 1 | 11 | B70s very sharp | 2000/500 MSK | [13 empty] | manual
|  | Pelvis 0.6 B70s | [22 empty] | 0.6 | 0.6 | 17 | B70s very sharp | 2000/500 MSK | [13 empty] | manual
|  | Pelvis 0.6 B80s | [22 empty] | 0.6 | 0.6 | 17 | B80s ultra sharp | 2000/500 MSK | [13 empty] | manual
|  | Pelvis 1.0 B80s | [22 empty] | 1 | 1 | 11 | B80s ultra sharp | 2000/500 MSK | [13 empty] | manual
|  | Pelvis 1.0 B70s | [22 empty] | 1 | 1 | 11 | B70s very sharp | 2000/500 MSK | [13 empty] | manual
| wrist UHR |  |  | 140 | 180 | 166 |  |  |  | 3 | Off | CARE Dose | 24.57 | 30s | 25% | 50 |  | 1 |  | 2 | 1 |  | 2 | 16x0.6mm | [15 empty] | Off
|  | wrist UHR 2.0 V80u 1 | [22 empty] | 2 | 1 | 11 | V80u very sharp | 2000/500 MSK | [13 empty] | manual
|  | wrist UHR 0.6 V80u Saf 1 | [22 empty] | 0.6 | 0.6 | 17 | V80u very sharp | 2000/500 MSK | [13 empty] | manual
|  | wrist UHR 0.6 V80u Saf 2 | [22 empty] | 0.6 | 0.6 | 17 | V80u very sharp | 2000/500 MSK | [13 empty] | manual
|  | wrist UHR 1.0 V80u Saf 1 | [22 empty] | 1 | 1 | 11 | V80u very sharp | 2000/500 MSK | [13 empty] | manual
| wrist UHR |  |  | 140 | 180 | 166 |  |  |  | 3 | Off | CARE Dose | 24.57 | 30s | 25% | 50 |  | 1 |  | 2 | 1 |  | 0.6 | 16x0.6mm | [15 empty] | Off
|  | wrist UHR 0.6 V80u 1 | [22 empty] | 0.6 | 0.6 | 17 | V80u very sharp | 2000/500 MSK | [13 empty] | manual
| Pelvis |  | 140 | 140 | 250 | 250 |  |  | Off | 2 | Off | CARE Dose4D | 24.55 | 30s | 25% | 50 |  | 1 |  | 2 | 1 |  | 0.6 | 128x0.6mm | [15 empty] | Off
|  | Pelvis 0.6 B70s | [22 empty] | 0.6 | 0.6 | 17 | B70s very sharp | 2000/500 MSK | [13 empty] | manual
| Pelvis 45 |  | 140 | 140 | 250 | 458 |  |  | Off | 2 | Off | CARE Dose4D | 44.97 | 30s | 25% | 50 |  | 1 |  | 2 | 1 |  | 0.6 | 128x0.6mm | [15 empty] | Off
|  | Pelvis 45 0.6 B80s | [22 empty] | 0.6 | 0.6 | 17 | B80s ultra sharp | 2000/500 MSK | [13 empty] | manual
|  | Pelvis 45 1.0 B80s | [22 empty] | 1 | 1 | 11 | B80s ultra sharp | 2000/500 MSK | [13 empty] | manual
|  | Pelvis 45 1.0 B70s | [22 empty] | 1 | 1 | 11 | B70s very sharp | 2000/500 MSK | [13 empty] | manual
| wrist UHR 45 |  |  | 140 | 180 | 304 |  |  |  | 3 | Off | CARE Dose | 45 | 30s | 25% | 50 |  | 1 |  | 2 | 1 |  | 0.6 | 16x0.6mm | [15 empty] | Off
|  | wrist UHR 45 0.6 V80u 1 | [22 empty] | 0.6 | 0.6 | 17 | V80u very sharp | 2000/500 MSK | [13 empty] | manual
|  | wrist UHR 45 1.0 V80u 1 | [22 empty] | 1 | 1 | 11 | V80u very sharp | 2000/500 MSK | [13 empty] | manual
| Pelvis 30 |  | 140 | 140 | 250 | 306 |  |  | Off | 2 | Off | CARE Dose4D | 30.05 | 30s | 25% | 50 |  | 1 |  | 2 | 1 |  | 0.6 | 128x0.6mm | [15 empty] | Off
|  | Pelvis 30 0.6 B80s | [22 empty] | 0.6 | 0.6 | 17 | B80s ultra sharp | 2000/500 MSK | [13 empty] | manual
|  | Pelvis 30 1.0 B80s | [22 empty] | 1 | 1 | 11 | B80s ultra sharp | 2000/500 MSK | [13 empty] | manual
|  | Pelvis 30 1.0 B70s | [22 empty] | 1 | 1 | 11 | B70s very sharp | 2000/500 MSK | [13 empty] | manual
| wrist UHR 30 |  |  | 140 | 180 | 203 |  |  |  | 3 | Off | CARE Dose | 30.05 | 30s | 25% | 50 |  | 1 |  | 2 | 1 |  | 0.6 | 16x0.6mm | [15 empty] | Off
|  | wrist UHR 30 0.6 V80u 1 | [22 empty] | 0.6 | 0.6 | 17 | V80u very sharp | 2000/500 MSK | [13 empty] | manual
|  | wrist UHR 30 1.0 V80u 1 | [22 empty] | 1 | 1 | 11 | V80u very sharp | 2000/500 MSK | [13 empty] | manual
| Pelvis 20 |  | 140 | 140 | 250 | 204 |  |  | Off | 2 | Off | CARE Dose4D | 20.03 | 30s | 25% | 50 |  | 1 |  | 2 | 1 |  | 0.6 | 128x0.6mm | [15 empty] | Off
|  | Pelvis 20 0.6 B80s | [22 empty] | 0.6 | 0.6 | 17 | B80s ultra sharp | 2000/500 MSK | [13 empty] | manual
|  | Pelvis 20 1.0 B80s | [22 empty] | 1 | 1 | 11 | B80s ultra sharp | 2000/500 MSK | [13 empty] | manual
|  | Pelvis 20 1.0 B70s | [22 empty] | 1 | 1 | 11 | B70s very sharp | 2000/500 MSK | [13 empty] | manual
| wrist UHR 20 |  |  | 140 | 180 | 135 |  |  |  | 3 | Off | CARE Dose | 19.98 | 30s | 25% | 50 |  | 1 |  | 2 | 1 |  | 0.6 | 16x0.6mm | [15 empty] | Off
|  | wrist UHR 20 0.6 V80u 1 | [22 empty] | 0.6 | 0.6 | 17 | V80u very sharp | 2000/500 MSK | [13 empty] | manual
|  | wrist UHR 20 1.0 V80u 1 | [22 empty] | 1 | 1 | 11 | V80u very sharp | 2000/500 MSK | [13 empty] | manual
| Pelvis 40 |  | 140 | 140 | 250 | 407 |  |  | Off | 2 | Off | CARE Dose4D | 39.96 | 30s | 25% | 50 |  | 1 |  | 2 | 1 |  | 0.6 | 128x0.6mm | [15 empty] | Off
|  | Pelvis 40 0.6 B80s | [22 empty] | 0.6 | 0.6 | 17 | B80s ultra sharp | 2000/500 MSK | [13 empty] | manual
|  | Pelvis 40 1.0 B80s | [22 empty] | 1 | 1 | 11 | B80s ultra sharp | 2000/500 MSK | [13 empty] | manual
|  | Pelvis 40 1.0 B70s | [22 empty] | 1 | 1 | 11 | B70s very sharp | 2000/500 MSK | [13 empty] | manual
| wrist UHR 40 |  |  | 140 | 180 | 270 |  |  |  | 3 | Off | CARE Dose | 39.96 | 30s | 25% | 50 |  | 1 |  | 2 | 1 |  | 0.6 | 16x0.6mm | [15 empty] | Off
|  | wrist UHR 40 0.6 V80u 1 | [22 empty] | 0.6 | 0.6 | 17 | V80u very sharp | 2000/500 MSK | [13 empty] | manual
|  | wrist UHR 40 1.0 V80u 1 | [22 empty] | 1 | 1 | 11 | V80u very sharp | 2000/500 MSK | [13 empty] | manual
| Pelvis 15 |  | 140 | 140 | 250 | 153 |  |  | Off | 2 | Off | CARE Dose4D | 15.02 | 30s | 25% | 50 |  | 1 |  | 2 | 1 |  | 0.6 | 128x0.6mm | [15 empty] | Off
|  | Pelvis 15 0.6 B80s | [22 empty] | 0.6 | 0.6 | 17 | B80s ultra sharp | 2000/500 MSK | [13 empty] | manual
|  | Pelvis 15 1.0 B80s | [22 empty] | 1 | 1 | 11 | B80s ultra sharp | 2000/500 MSK | [13 empty] | manual
|  | Pelvis 15 1.0 B70s | [22 empty] | 1 | 1 | 11 | B70s very sharp | 2000/500 MSK | [13 empty] | manual
| wrist UHR 15 |  |  | 140 | 180 | 101 |  |  |  | 3 | Off | CARE Dose | 14.95 | 30s | 25% | 50 |  | 1 |  | 2 | 1 |  | 0.6 | 16x0.6mm | [15 empty] | Off
|  | wrist UHR 15 0.6 V80u 1 | [22 empty] | 0.6 | 0.6 | 17 | V80u very sharp | 2000/500 MSK | [13 empty] | manual
|  | wrist UHR 15 1.0 V80u 1 | [22 empty] | 1 | 1 | 11 | V80u very sharp | 2000/500 MSK | [13 empty] | manual
| Pelvis 10 |  | 140 | 140 | 250 | 102 |  |  | Off | 2 | Off | CARE Dose4D | 10.02 | 30s | 25% | 50 |  | 1 |  | 2 | 1 |  | 0.6 | 128x0.6mm | [15 empty] | Off
|  | Pelvis 10 0.6 B80s | [22 empty] | 0.6 | 0.6 | 17 | B80s ultra sharp | 2000/500 MSK | [13 empty] | manual
|  | Pelvis 10 1.0 B80s | [22 empty] | 1 | 1 | 11 | B80s ultra sharp | 2000/500 MSK | [13 empty] | manual
|  | Pelvis 10 1.0 B70s | [22 empty] | 1 | 1 | 11 | B70s very sharp | 2000/500 MSK | [13 empty] | manual
| wrist UHR 10 |  |  | 140 | 180 | 68 |  |  |  | 3 | Off | CARE Dose | 10.07 | 30s | 25% | 50 |  | 1 |  | 2 | 1 |  | 0.6 | 16x0.6mm | [15 empty] | Off
|  | wrist UHR 10 0.6 V80u 1 | [22 empty] | 0.6 | 0.6 | 17 | V80u very sharp | 2000/500 MSK | [13 empty] | manual
|  | wrist UHR 10 1.0 V80u 1 | [22 empty] | 1 | 1 | 11 | V80u very sharp | 2000/500 MSK | [13 empty] | manual
RT
RT_Abdomen (Adult)
RT_Abdomen (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Abdomen |  | 120 | 120 | 250 | 250 |  |  | Semi | 3 | On | CARE Dose4D | 17.28 | 99s | 25% |  |  | 1 |  | 2 | 0.6 |  | 4 | 32x1.2mm | [15 empty] | Off
|  | Abdomen 4.0 B30s | [22 empty] | 4 | 4 | 51 | B30s medium smooth | Abdomen | [13 empty] | manual
RT_Head (Adult)
RT_Head (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.29 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head
| Head |  | 120 | 120 | 410 | 410 |  |  | Off | 3 | On | CARE Dose4D | 64.03 | 99s | 25% |  |  | 1 |  | 2 | 0.55 |  | 3 | 32x1.2mm | [15 empty] | Off
|  | Head 3.0 H31s | [22 empty] | 3 | 3 | 41 | H31s medium smooth + | Cerebrum | [13 empty] | manual
RT_HeadNeckShoulder (Adult)
RT_HeadNeckShoulder (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| HeadNeckShoul |  | 120 | 120 | 200 | 200 |  |  | Semi | 3 | On | CARE Dose4D | 16.69 | 99s | 25% |  |  | 1 |  | 2 | 0.8 |  | 4 | 32x1.2mm | [15 empty] | Off
|  | HeadNeckShoul 4.0 B30s | [22 empty] | 4 | 4 | 38 | B30s medium smooth | Larynx | [13 empty] | manual
RT_Pelvis (Adult)
RT_Pelvis (Adult) | Topogram | Topogram 0.6 T80f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Body
| Pelvis |  | 120 | 120 | 270 | 270 |  |  | Semi | 3 | On | CARE Dose4D | 18.65 | 99s | 25% |  |  | 1 |  | 2 | 0.8 |  | 4 | 32x1.2mm | [15 empty] | Off
|  | Pelvis 4.0 B30s | [22 empty] | 4 | 4 | 51 | B30s medium smooth | Pelvis | [13 empty] | manual
RT_Thorax (Adult)
RT_Thorax (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Thorax |  | 120 | 120 | 140 | 140 |  |  | Semi | 3 | On | CARE Dose4D | 9.68 | 99s | 25% |  |  | 1 |  | 2 | 1.2 |  | 4 | 32x1.2mm | [15 empty] | Off
|  | Thorax 4.0 B30s | [22 empty] | 4 | 4 | 76 | B30s medium smooth | Mediastinum | [13 empty] | manual
DualEnergy
DE_AbdomenSeq (Adult)
DE_AbdomenSeq (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| DE_AbdSeq |  |  | 100 | 171 |  | 171 | 132 |  | 3 | On | CARE Dose4D | 14.96 |  | 25% |  |  | 0.5 |  | 2 | 0 | 17 | 5 | 32x0.6mm | [15 empty] | Off
|  | DE_AbdSeq 5.0 B31f | [22 empty] | 5 |  | 37 | B31f medium smooth + | Abdomen | [13 empty] | manual
|  | DE #PP DE_AbdSeq 1.5 D30f | [22 empty] | 1.5 |  | 186 | D30f medium smooth | Abdomen | [13 empty] | manual
DE_Abdomen_KidneyStones (Adult)
DE_Abdomen_KidneyStones (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| DE_Abdomen |  |  | 100 | 210 |  | 210 | 162 |  | 3 | On | CARE Dose4D | 16.67 | 30s | 25% |  |  | 0.33 |  | 2 | 0.7 |  | 5 | 64x0.6mm | [15 empty] | Off
|  | DE_Abdomen 5.0 B30f | [22 empty] | 5 | 5 | 41 | B30f medium smooth | Abdomen | [13 empty] | manual
|  | DE #PP DE_Abdomen 1.5 D30f | [22 empty] | 1.5 | 1 | 201 | D30f medium smooth | Abdomen | [13 empty] | manual
DE_Abdomen_VNC (Adult)
DE_Abdomen_VNC (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| DE_Abdomen |  |  | 100 | 230 |  | 230 | 178 |  | 3 | On | CARE Dose4D | 17.88 | 30s | 25% |  |  | 0.5 |  | 2 | 0.6 |  | 5 | 32x0.6mm | [15 empty] | Off
|  | DE_Abdomen 5.0 B30f | [22 empty] | 5 | 5 | 41 | B30f medium smooth | Abdomen | [13 empty] | manual
|  | DE #PP DE_Abdomen 1.5 D30f | [22 empty] | 1.5 | 1 | 201 | D30f medium smooth | Abdomen | [13 empty] | manual
DE_BodyAngioFast_BoneRem (Adult)
DE_BodyAngioFast_BoneRem (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| DE_BodyAngio |  |  | 100 | 119 |  | 119 | 92 |  | 3 | On | CARE Dose4D | 8.28 | 6s | 25% |  |  | 0.33 |  | 2 | 0.7 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | DE_BodyAngio 5.0 B30f | [22 empty] | 5 | 5 | 81 | B30f medium smooth | Abdomen | [13 empty] | manual
|  | DE #PP DE_BodyAngio 1.5 D26f | [22 empty] | 1.5 | 1 | 401 | D26f med. smooth ASA | Abdomen | [13 empty] | manual
DE_BodyAngio_BoneRem (Adult)
DE_BodyAngio_BoneRem (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| DE_BodyAngio |  |  | 100 | 119 |  | 119 | 92 |  | 3 | On | CARE Dose4D | 9.29 | 15s | 25% |  |  | 0.5 |  | 2 | 0.9 |  | 5 | 32x0.6mm | [15 empty] | Off
|  | DE_BodyAngio 5.0 B30f | [22 empty] | 5 | 5 | 81 | B30f medium smooth | Abdomen | [13 empty] | manual
|  | DE #PP DE_BodyAngio 1.5 D26f | [22 empty] | 1.5 | 1 | 401 | D26f med. smooth ASA | Abdomen | [13 empty] | manual
DE_BodyAngio_BoneRem_IR (Adult)
DE_BodyAngio_BoneRem_IR (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| DE_BodyAngio |  |  | 100 | 95 |  | 95 | 73 |  | 3 | On | CARE Dose4D | 7.39 | 15s | 25% |  |  | 0.5 |  | 2 | 0.9 |  | 5 | 32x0.6mm | [15 empty] | Off
|  | DE_BodyAngio 5.0 I30f 3 | [22 empty] | 5 | 5 | 81 | I30f medium smooth | Abdomen | [13 empty] | manual
|  | DE #PP DE_BodyAngio 1.5 Q30f 3 | [22 empty] | 1.5 | 1 | 401 | Q30f medium smooth | Abdomen | [13 empty] | manual
DE_Body_BoneMarrow (Adult)
DE_Body_BoneMarrow (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 365 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Topogram | Topogram 0.6 T20f |  | 120 |  | 365 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| DE_Body |  |  | 100 | 230 |  | 230 | 178 |  | 3 | On | CARE Dose4D | 17.88 | 30s | 25% |  |  | 0.5 |  | 2 | 0.6 |  | 5 | 32x0.6mm | [15 empty] | Off
|  | DE_Body 5.0 B30f | [22 empty] | 5 | 5 | 141 | B30f medium smooth | Abdomen | [13 empty] | manual
|  | DE #PP DE_Body 1.5 D34f | [22 empty] | 1.5 | 1 | 701 | D34f med. smooth BHCBone | Abdomen | [13 empty] | manual
DE_Body_OptiCon (Adult)
DE_Body_OptiCon (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| DE_Body |  |  | 140 | 32 |  | 32 | 176 |  | 3 | On | CARE Dose4D | 8.67 | 19s | 25% |  |  | 0.5 |  | 2 | 0.7 |  | 5 | 32x0.6mm | [15 empty] | Off
|  | DE_Body 5.0 B30f | [22 empty] | 5 | 5 | 81 | B30f medium smooth | Abdomen | [13 empty] | manual
|  | DE #PP DE_Body 1.5 D30f | [22 empty] | 1.5 | 1 | 401 | D30f medium smooth | Abdomen | [13 empty] | manual
DE_CarotidAngioFast_BoneRem (Adult)
DE_CarotidAngioFast_BoneRem (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| DE_CarotidAngio |  |  | 100 | 104 |  | 104 | 104 |  | 3 | On | CARE Dose4D | 8.09 | 10s | 25% |  |  | 0.285 |  | 2 | 0.7 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | DE_CarotidAngio 5.0 B30f | [22 empty] | 5 | 5 | 51 | B30f medium smooth | Larynx | [13 empty] | manual
|  | DE #PP DE_CarotidAngio 1.0 D26f | [22 empty] | 1 | 0.7 | 358 | D26f med. smooth ASA | Larynx | [13 empty] | manual
DE_CarotidAngio_BoneRem (Adult)
DE_CarotidAngio_BoneRem (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| DE_CarotidAngio |  |  | 100 | 114 |  | 114 | 114 |  | 3 | On | CARE Dose4D | 8.03 | 10s | 25% |  |  | 0.285 |  | 2 | 0.9 |  | 5 | 64x0.6mm | [15 empty] | Off
|  | DE_CarotidAngio 5.0 B30f | [22 empty] | 5 | 5 | 51 | B30f medium smooth | Larynx | [13 empty] | manual
|  | DE #PP DE_CarotidAngio 1.0 D26f | [22 empty] | 1 | 0.7 | 358 | D26f med. smooth ASA | Larynx | [13 empty] | manual
DE_CoronaryCTA_HeartPBV (Adult)
DE_CoronaryCTA_HeartPBV (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| DE_CorCTA |  |  | 100 | 165 |  | 165 | 140 |  | 3 | On | CARE Dose4D | 64.62 |  | 25% |  |  | 0.285 |  | 2 | 0.17 |  | 3 | 64x0.6mm | [15 empty] | Normal | 60% | 80%
|  | DE_CorCTA 3.0 B30f 70% | [22 empty] | 3 | 3 | 41 | B30f medium smooth | Cardiac | [13 empty] | manual | 70%
|  | QCA DE_CorCTA 0.6 B26f 70% | [22 empty] | 0.6 | 0.3 | 401 | B26f medium smooth ASA | Cardiac | [13 empty] | manual | 70%
|  | DE #PP DE_CorCTA 1.0 D33f 60% | [22 empty] | 1 | 0.7 | 172 | D33f med. smooth BHCIodine | Cardiac | [13 empty] | manual | 60%
DE_CoronaryCTA_Seq (Adult)
DE_CoronaryCTA_Seq (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| DE_CorSeq |  |  | 100 | 165 |  | 165 | 140 |  | 3 | On | CARE Dose4D | 12.16 |  | 25% |  |  | 0.285 |  | 2 | 0 | 17 | 3 | 64x0.6mm | [13 empty] | 60% | 60% | Off
|  | DE_CorSeq 3.0 B30f 60% | [22 empty] | 3 |  | 39 | B30f medium smooth | Cardiac | [13 empty] | manual | 60%
|  | DE #PP DE_CorSeq 1.5 D33f 60% | [22 empty] | 1.5 |  | 118 | D33f med. smooth BHCIodine | Cardiac | [13 empty] | manual | 60%
DE_CTA_Spine_Customized (Adult)
DE_CTA_Spine_Customized (Adult) | Topogram PA | Topogram PA 0.6 T20s |  | 120 |  | 393 |  |  |  | 3 | Off | CARE Dose | 0.2 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body |  |  |  |  | ROEMIR001a | TRACKER
| Topogram LAT | Topogram LAT 0.6 T20s |  | 120 |  | 393 |  |  |  | 3 | Off | CARE Dose | 0.2 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body |  |  |  |  | ROEMIR001a | TRACKER
| PreMonitoring |  | 120 | 120 | 300 | 100 |  |  | Off | 11 | Off | CARE Dose4D | 6.16 |  | 25% |  |  | 0.5 |  | 2 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | PreMonitoring 10.0 B40s | [22 empty] | 10 |  | 1 | B40s medium | 400/40 Spine Std | [13 empty] | manual
| Monitoring |  | 120 | 120 | 300 | 100 |  |  | Off | 11 | Off | CARE Dose4D | 246.56 |  | 25% |  |  | 0.5 |  | 10 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | Monitoring 10.0 B40s | [22 empty] | 10 |  | 40 | B40s medium | 400/40 Spine Std | [13 empty] | manual
| DE_CTA Spine |  |  | 100 | 300 |  | 119 | 92 |  | 11 | On | CARE Dose4D | 23.37 | 30s | 25% |  |  | 0.5 |  | 5 | 0.9 |  | 3 | 32x0.6mm | [15 empty] | Off
|  | DE_CTA Spine 3.0 B40f | [22 empty] | 3 | 3 | 134 | B40f medium | 600/175 Neuro CTA |  |  |  | GO_NEURO | [9 empty] | manual
|  | DE_CTA Spine 0.75 D30f | [22 empty] | 0.75 | 0.7 | 572 | D30f medium smooth | 600/175 Neuro CTA |  |  |  | GO_NEURO | [9 empty] | manual
| DE_CTA Spine |  |  | 100 | 300 |  | 119 | 92 |  | 11 | On | CARE Dose4D | 23.37 | 30s | 25% |  |  | 0.5 |  | 2 | 0.9 |  | 3 | 32x0.6mm | [15 empty] | Off
|  | DE_CTA Spine 3.0 B40f | [22 empty] | 3 | 3 | 134 | B40f medium | 600/175 Neuro CTA |  |  |  | GO_NEURO | [9 empty] | manual
|  | DE_CTA Spine 0.75 D30f | [22 empty] | 0.75 | 0.7 | 572 | D30f medium smooth | 600/175 Neuro CTA |  |  |  | GO_NEURO | [9 empty] | manual
DE_ExtrAngio_BoneRem (Adult)
DE_ExtrAngio_BoneRem (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 365 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Topogram | Topogram 0.6 T20f |  | 120 |  | 365 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| DE_ExtrAngio |  |  | 80 | 230 |  | 230 | 98 |  | 3 | On | CARE Dose4D | 9.24 | 30s | 25% |  |  | 0.33 |  | 2 | 0.7 |  | 5 | 64x0.6mm | [15 empty] | Off
|  | DE_ExtrAngio 5.0 B30f | [22 empty] | 5 | 5 | 201 | B30f medium smooth | Abdomen | [13 empty] | manual
|  | DE #PP DE_ExtrAngio 1.5 D30f | [22 empty] | 1.5 | 1 | 1001 | D30f medium smooth | Abdomen | [13 empty] | manual
DE_Foot_Gout (Adult)
DE_Foot_Gout (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| DE_Foot |  |  | 80 | 250 |  | 250 | 125 |  | 3 | On | CARE Dose | 10.8 | 10s | 25% |  |  | 0.33 |  | 2 | 0.7 |  | 2 | 64x0.6mm | [15 empty] | Off
|  | DE_Foot 2.0 B20f | [22 empty] | 2 | 2 | 76 | B20f smooth | Extremity | [13 empty] | manual
|  | DE #PP DE_Foot 0.75 D34f | [22 empty] | 0.75 | 0.5 | 301 | D34f med. smooth BHCBone | Bone | [13 empty] | manual
DE_Hand_Gout (Adult)
DE_Hand_Gout (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body |  |  |  |  | ROEMIR001a | TRACKER
| Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body |  |  |  |  | ROEMIR001a | TRACKER
| DE_Hand |  |  | 80 | 205 |  | 205 | 103 |  | 3 | On | CARE Dose | 8.89 | 6s | 25% |  |  | 0.33 |  | 2 | 0.7 |  | 2 | 64x0.6mm | [15 empty] | Off
|  | DE_Hand 2.0 B20f | [22 empty] | 2 | 2 | 76 | B20f smooth | Extremity |  |  |  | GEPACS | [9 empty] | manual
|  | DE #PP DE_Hand 0.6 D45f | [22 empty] | 0.6 | 0.3 | 501 | D45f medium sharp | Bone |  |  |  | GEPACS | [9 empty] | manual
DE_HeadAngio_BoneRem (Adult)
DE_HeadAngio_BoneRem (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.29 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head
| Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.29 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head
| DE_HeadAngio |  |  | 80 | 222 |  | 222 | 111 |  | 3 | On | CARE Dose4D | 19.4 | 10s | 25% |  |  | 0.33 |  | 2 | 0.7 |  | 4 | 64x0.6mm | [15 empty] | Off
|  | DE_HeadAngio 4.0 H20f | [22 empty] | 4 | 4 | 31 | H20f smooth | Base Orbita | [13 empty] | manual
|  | DE #PP DE_HeadAngio 1.0 D26f | [22 empty] | 1 | 0.7 | 172 | D26f med. smooth ASA | Larynx | [13 empty] | manual
DE_Head_BrainHem (Adult)
DE_Head_BrainHem (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.29 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head
| Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.29 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head
| DE_Head |  |  | 80 | 310 |  | 310 | 155 |  | 3 | On | CARE Dose4D | 28.96 | 16s | 25% |  |  | 0.5 |  | 2 | 0.7 |  | 4 | 40x0.6mm | [15 empty] | Off
|  | DE_Head 4.0 H20f | [22 empty] | 4 | 4 | 31 | H20f smooth | Base Orbita | [13 empty] | manual
|  | DE #PP DE_Head 1.0 D34f | [22 empty] | 1 | 0.7 | 172 | D34f med. smooth BHCBone | Base Orbita | [13 empty] | manual
DE_Knee_BoneMarrow (Adult)
DE_Knee_BoneMarrow (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| DE_Knee |  |  | 80 | 240 |  | 240 | 120 |  | 3 | On | CARE Dose | 10.18 | 8s | 25% |  |  | 0.5 |  | 2 | 0.7 |  | 3 | 32x0.6mm | [15 empty] | Off
|  | DE_Knee 3.0 B30f | [22 empty] | 3 | 3 | 51 | B30f medium smooth | Extremity | [13 empty] | manual
|  | DE #PP DE_Knee 0.75 D30f | [22 empty] | 0.75 | 0.5 | 301 | D30f medium smooth | Bone | [13 empty] | manual
DE_Knee_Tendon (Adult)
DE_Knee_Tendon (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| DE_Knee |  |  | 140 | 45 |  | 45 | 191 |  | 3 | On | CARE Dose | 11.57 | 23s | 25% |  |  | 1 |  | 2 | 0.7 |  | 2 | 40x0.6mm | [15 empty] | Off
|  | DE_Knee 2.0 B20s | [22 empty] | 2 | 2 | 76 | B20s smooth | Extremity | [13 empty] | manual
|  | DE #PP DE_Knee 0.75 D34s | [22 empty] | 0.75 | 0.5 | 301 | D34s med. smooth BHCBone | Bone | [13 empty] | manual
DE_Spine_Metal_Vol (Adult)
DE_Spine_Metal_Vol (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| DE_Spine |  |  | 100 | 250 |  | 250 | 483 |  | 3 | Off | CARE Dose4D | 31.12 | 40s | 25% |  |  | 0.5 |  | 2 | 0.5 |  | 3 | 32x0.6mm | [15 empty] | Off
|  | DE_Spine 3.0 B30f | [22 empty] | 3 | 3 | 54 | B30f medium smooth | Spine | [13 empty] | manual
|  | DE_Spine 1.5 D30f | [22 empty] | 1.5 | 1 | 161 | D30f medium smooth | Spine | [13 empty] | manual
|  | DE_Spine 1.5 B20f | [22 empty] | 1.5 | 1 | 161 | B20f smooth | Spine | [13 empty] | manual
|  | DE_Spine 2.0 MPR spine multi | [22 empty] | 2 | 2 | 0 | B20f smooth | Spine | [13 empty] | manual
|  | DE_Spine 2.0 MPR spine multi | [22 empty] | 2 | 2 | 0 | B60f sharp | Spine | [13 empty] | manual
DE_Thorax_Analysis (Adult)
DE_Thorax_Analysis (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| DE_Thorax |  |  | 100 | 89 |  | 89 | 76 |  | 3 | On | CARE Dose4D | 7.4 | 10s | 25% |  |  | 0.285 |  | 2 | 0.55 |  | 5 | 64x0.6mm | [15 empty] | Off
|  | DE_Thorax 5.0 B31f | [22 empty] | 5 | 5 | 61 | B31f medium smooth + | Mediastinum | [13 empty] | manual
|  | DE_Thorax 5.0 B80f | [22 empty] | 5 | 5 | 61 | B80f ultra sharp | Lung | [13 empty] | manual
|  | DE #PP DE_Thorax 2.0 D33f | [22 empty] | 2 | 1.4 | 215 | D33f med. smooth BHCIodine | Mediastinum | [13 empty] | manual
DE_Thorax_Analysis_IR (Adult)
DE_Thorax_Analysis_IR (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| DE_Thorax |  |  | 100 | 71 |  | 71 | 60 |  | 3 | On | CARE Dose4D | 5.89 | 10s | 25% |  |  | 0.285 |  | 2 | 0.55 |  | 5 | 64x0.6mm | [15 empty] | Off
|  | DE_Thorax 5.0 I31f 3 | [22 empty] | 5 | 5 | 61 | I31f medium smooth | Mediastinum | [13 empty] | manual
|  | DE_Thorax 5.0 I70f 3 | [22 empty] | 5 | 5 | 61 | I70f very sharp ASA | Lung | [13 empty] | manual
|  | DE #PP DE_Thorax 2.0 Q30f 3 | [22 empty] | 2 | 1.4 | 215 | Q30f medium smooth | Mediastinum | [13 empty] | manual
DE_Thorax_Analysis_XCARE (Adult)
DE_Thorax_Analysis_XCARE (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| DE_Thorax |  |  | 100 | 89 |  | 89 | 76 |  | 3 | On | CARE Dose4D | 7.4 | 10s | 25% |  |  | 0.285 |  | 2 | 0.55 |  | 5 | 64x0.6mm | [15 empty] | Off
|  | DE_Thorax 5.0 B31f | [22 empty] | 5 | 5 | 61 | B31f medium smooth + | Mediastinum |  |  | X-CARE | [10 empty] | manual
|  | DE_Thorax 5.0 B80f | [22 empty] | 5 | 5 | 61 | B80f ultra sharp | Lung |  |  | X-CARE | [10 empty] | manual
|  | DE #PP DE_Thorax 2.0 D30f | [22 empty] | 2 | 1.4 | 215 | D30f medium smooth | Mediastinum |  |  | X-CARE | [10 empty] | manual
DE_Thorax_LungNodules (Adult)
DE_Thorax_LungNodules (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| DE_Thorax |  |  | 80 | 205 |  | 205 | 87 |  | 3 | On | CARE Dose4D | 8.25 | 10s | 25% |  |  | 0.285 |  | 2 | 0.55 |  | 5 | 64x0.6mm | [15 empty] | Off
|  | DE_Thorax 5.0 B31f | [22 empty] | 5 | 5 | 61 | B31f medium smooth + | Mediastinum | [13 empty] | manual
|  | DE_Thorax 5.0 B80f | [22 empty] | 5 | 5 | 61 | B80f ultra sharp | Lung | [13 empty] | manual
|  | DE #PP DE_Thorax 1.5 D30f | [22 empty] | 1.5 | 1 | 301 | D30f medium smooth | Mediastinum | [13 empty] | manual
DE_Thorax_Xenon (Adult)
DE_Thorax_Xenon (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| DE_Thorax |  |  | 100 | 89 |  | 89 | 76 |  | 3 | On | CARE Dose4D | 7.4 | 10s | 25% |  |  | 0.285 |  | 2 | 0.55 |  | 5 | 64x0.6mm | [15 empty] | Off
|  | DE_Thorax 5.0 B31f | [22 empty] | 5 | 5 | 61 | B31f medium smooth + | Mediastinum | [13 empty] | manual
|  | DE_Thorax 5.0 B80f | [22 empty] | 5 | 5 | 61 | B80f ultra sharp | Lung | [13 empty] | manual
|  | DE #PP DE_Thorax 2.0 D34f | [22 empty] | 2 | 1.4 | 215 | D34f med. smooth BHCBone | Lung | [13 empty] | manual
DE_Wrist_Tendon (Adult)
DE_Wrist_Tendon (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| DE_Wrist |  |  | 140 | 37 |  | 37 | 157 |  | 3 | On | CARE Dose | 8.85 | 5s | 25% |  |  | 0.33 |  | 2 | 0.7 |  | 2 | 64x0.6mm | [15 empty] | Off
|  | DE_Wrist 2.0 B20f | [22 empty] | 2 | 2 | 31 | B20f smooth | Extremity | [13 empty] | manual
|  | DE #PP DE_Wrist 0.75 D34f | [22 empty] | 0.75 | 0.5 | 121 | D34f med. smooth BHCBone | Bone | [13 empty] | manual
RDE_LiverIron_Customized (Adult)
RDE_LiverIron_Customized (Adult) | Topogram | Topogram 0.6 T20s |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 67092aa9-b6e0-47d9-8c91-ba48e48b197a |  |  |  | ROEMIR001a | TRACKER
| Topogram | Topogram 0.6 T20s |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 67092aa9-b6e0-47d9-8c91-ba48e48b197a |  |  |  | ROEMIR001a | TRACKER
| DE_80/140kv |  |  | 140 | 106 |  | 71 | 391 |  | 3 | Off | CARE Dose4D | 19.22 | 30s | 25% | 50 |  | 0.5 |  | 5 | 0.7 |  | 2 | 32x0.6mm | [5 empty] | 67092aa9-b6e0-47d9-8c91-ba48e48b197a | [9 empty] | Off
|  | DE_80/140kv 2.0 B40f | [22 empty] | 2 | 1.5 | 101 | B40f medium | 400/40 Abd | [13 empty] | manual
|  | DE_80/140kv 5.0 D40f | [22 empty] | 5 | 1 | 151 | D40f medium | 400/40 Abd | [13 empty] | manual
| DE_100/140sn |  |  | 100 | 240 |  | 310 | 240 |  | 3 | Off | CARE Dose4D | 24.11 | 30s | 25% | 50 |  | 0.5 |  | 11 | 0.7 |  | 2 | 32x0.6mm | [5 empty] | 67092aa9-b6e0-47d9-8c91-ba48e48b197a | [9 empty] | Off
|  | DE_100/140sn 2.0 B40f | [22 empty] | 2 | 1.5 | 101 | B40f medium | 400/40 Abd | [13 empty] | manual
|  | DE_100/140sn 5.0 D40f | [22 empty] | 5 | 1 | 151 | D40f medium | 400/40 Abd | [13 empty] | manual
| DE_80/140sn |  |  | 80 | 404 |  | 461 | 178 |  | 3 | Off | CARE Dose4D | 17.44 | 30s | 25% | 50 |  | 0.5 |  | 11 | 0.7 |  | 2 | 32x0.6mm | [5 empty] | 67092aa9-b6e0-47d9-8c91-ba48e48b197a | [9 empty] | Off
|  | DE_80/140sn 2.0 B40f | [22 empty] | 2 | 1.5 | 101 | B40f medium | 400/40 Abd | [13 empty] | manual
|  | DE_80/140sn 5.0 B40f | [22 empty] | 5 | 1 | 151 | B40f medium | 400/40 Abd | [13 empty] | manual
Siemens default scan protocols
Head
HeadRoutine (Child)
HeadRoutine (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.16 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head
| Head |  | 100 | 100 | 582 | 291 |  |  | On | 3 | On | CARE Dose4D | 56.22 | 11s | 25% |  |  | 1 |  | 2 | 0.8 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | Head 5.0 H31s | [22 empty] | 5 | 5 | 25 | H31s medium smooth + | Baby Brain | [13 empty] | manual
HeadRoutineSeq (Child)
HeadRoutineSeq (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.16 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head
| HeadSeq |  | 100 | 100 | 582 | 291 |  |  | On | 3 | On | CARE Dose4D | 62.48 |  | 25% |  |  | 1 |  | 2 | 0 | 34 | 5 | 128x0.6mm | [15 empty] | Off
|  | HeadSeq 5.0 H31s | [22 empty] | 5 |  | 27 | H31s medium smooth + | Baby Brain | [13 empty] | manual
Head_IR (Child)
Head_IR (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.16 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head
| Head |  | 100 | 100 | 407 | 204 |  |  | On | 3 | On | CARE Dose4D | 39.33 | 11s | 25% |  |  | 1 |  | 2 | 0.8 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | Head 5.0 J30s 3 | [22 empty] | 5 | 5 | 25 | J30s medium smooth | Baby Brain | [13 empty] | manual
Head_XCARE (Child)
Head_XCARE (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.16 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head
| Head |  | 100 | 100 | 582 | 291 |  |  | On | 3 | On | CARE Dose4D | 56.36 | 14s | 25% |  |  | 1 |  | 2 | 0.6 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | Head 5.0 H31s | [22 empty] | 5 | 5 | 25 | H31s medium smooth + | Baby Brain |  |  | X-CARE | [10 empty] | manual
Head_XCARE_IR (Child)
Head_XCARE_IR (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.16 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head
| Head |  | 100 | 100 | 407 | 204 |  |  | On | 3 | On | CARE Dose4D | 39.45 | 14s | 25% |  |  | 1 |  | 2 | 0.6 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | Head 5.0 J30s 3 | [22 empty] | 5 | 5 | 25 | J30s medium smooth | Baby Brain |  |  | X-CARE | [10 empty] | manual
InnerEar (Child)
InnerEar (Child) | Topogram | Topogram 0.6 T80f |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.16 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Head
| InnerEar |  | 120 | 120 | 180 | 90 |  |  | On | 5 | On | CARE Dose4D | 22.46 | 6s | 25% |  |  | 1 |  | 2 | 0.8 |  | 0.6 | 128x0.6mm | [15 empty] | Off
|  | InnerEar 0.6 H60s | [22 empty] | 0.6 | 0.6 | 67 | H60s sharp FR | Inner Ear | [13 empty] | manual
InnerEarUHR (Child)
InnerEarUHR (Child) | Topogram | Topogram 0.6 T80f |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.16 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Head
| InnerEarUHR |  | 120 | 120 | 180 | 90 |  |  | On | 5 | On | CARE Dose4D | 39.82 | 24s | 25% |  |  | 1 |  | 2 | 0.85 |  | 0.6 | 16x0.6mm | [15 empty] | Off
|  | InnerEarUHR 0.6 U75u | [22 empty] | 0.6 | 0.6 | 67 | U75u very sharp ASA | Inner Ear | [13 empty] | manual
InnerEarUHRSeq (Child)
InnerEarUHRSeq (Child) | Topogram | Topogram 0.6 T80f |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.16 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Head
| InnEarUHRSeq |  | 120 | 120 | 180 | 90 |  |  | On | 5 | On | CARE Dose4D | 41.65 |  | 25% |  |  | 1 |  | 2 | 0 | 4 | 0.6 | 16x0.6mm | [15 empty] | Off
|  | InnEarUHRSeq 0.6 U75u | [22 empty] | 0.6 |  | 68 | U75u very sharp ASA | Inner Ear | [13 empty] | manual
InnerEarZUHR (Child)
InnerEarZUHR (Child) | Topogram | Topogram 0.6 T80f |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.16 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Head
| InnerEarUHR |  | 120 | 120 | 230 | 115 |  |  | On | 5 | On | CARE Dose4D | 51.02 | 24s | 25% |  |  | 1 |  | 2 | 0.85 |  | 0.4 | 16x0.3mm | [15 empty] | Off
|  | InnerEarUHR 0.4 U75u | [22 empty] | 0.4 | 0.4 | 101 | U75u very sharp ASA | Inner Ear | [13 empty] | manual
InnerEarZUHR_IR (Child)
InnerEarZUHR_IR (Child) | Topogram | Topogram 0.6 T80f |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.16 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Head
| InnerEarUHR |  | 120 | 120 | 161 | 81 |  |  | On | 5 | On | CARE Dose4D | 35.66 | 24s | 25% |  |  | 1 |  | 2 | 0.85 |  | 0.4 | 16x0.3mm | [15 empty] | Off
|  | InnerEarUHR 0.4 V80u 3 | [22 empty] | 0.4 | 0.4 | 101 | V80u very sharp | Inner Ear | [13 empty] | manual
NeonateHead (Child)
NeonateHead (Child) | Topogram | Topogram 0.6 T20f |  | 80 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.08 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head
| NeonateHead |  | 100 | 100 | 582 | 150 |  |  | On | 3 | On | CARE Dose4D | 56.22 | 11s | 25% |  |  | 1 |  | 2 | 0.8 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | NeonateHead 5.0 C30s | [22 empty] | 5 | 5 | 25 | C30s med. smooth FR | Baby Brain | [13 empty] | manual
Orbit (Child)
Orbit (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.16 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head
| Orbit |  | 120 | 120 | 125 | 63 |  |  | On | 5 | On | CARE Dose4D | 15.6 | 10s | 25% |  |  | 1 |  | 2 | 0.8 |  | 3 | 128x0.6mm | [15 empty] | Off
|  | Orbit 3.0 H60s | [22 empty] | 3 | 3 | 14 | H60s sharp FR | Base Orbita | [13 empty] | manual
|  | Orbit 3.0 H30s | [22 empty] | 3 | 3 | 14 | H30s medium smooth | Base Orbita | [13 empty] | manual
|  | Orbit 0.6 H60s | [22 empty] | 0.6 | 0.4 | 101 | H60s sharp FR | Base Orbita | [13 empty] | manual
Sinus (Child)
Sinus (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.16 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head
| Sinus |  | 120 | 120 | 50 | 25 |  |  | On | 5 | On | CARE Dose4D | 6.24 | 10s | 25% |  |  | 1 |  | 2 | 0.8 |  | 3 | 128x0.6mm | [15 empty] | Off
|  | Sinus 3.0 H60s | [22 empty] | 3 | 3 | 21 | H60s sharp FR | Sinuses | [13 empty] | manual
|  | Sinus 3.0 H30s | [22 empty] | 3 | 3 | 21 | H30s medium smooth | Sinuses | [13 empty] | manual
|  | Sinus 1.0 H60s | [22 empty] | 1 | 0.7 | 86 | H60s sharp FR | Sinuses | [13 empty] | manual
|  | Sinus 1.0 H30s | [22 empty] | 1 | 0.7 | 86 | H30s medium smooth | Sinuses | [13 empty] | manual
Neck
NeckRoutine (Child)
NeckRoutine (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.07 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Neck |  | 100 | 100 | 233 | 117 |  |  | On | 7 | On | CARE Dose4D | 7.95 | 9s | 25% |  |  | 1 |  | 2 | 0.8 |  | 5 | 32x1.2mm | [15 empty] | Off
|  | Neck 5.0 B30s | [22 empty] | 5 | 5 | 35 | B30s medium smooth | Baby Neck | [13 empty] | manual
|  | Neck 1.5 B60s | [22 empty] | 1.5 | 1 | 171 | B60s sharp | Baby Neck | [13 empty] | manual
Neck_IR (Child)
Neck_IR (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.07 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Neck |  | 100 | 100 | 163 | 82 |  |  | On | 7 | On | CARE Dose4D | 5.57 | 9s | 25% |  |  | 1 |  | 2 | 0.8 |  | 5 | 32x1.2mm | [15 empty] | Off
|  | Neck 5.0 I30s 3 | [22 empty] | 5 | 5 | 35 | I30s medium smooth | Baby Neck | [13 empty] | manual
|  | Neck 1.5 I50s 3 | [22 empty] | 1.5 | 1 | 171 | I50s medium sharp ASA | Baby Neck | [13 empty] | manual
Shoulder
ShoulderRoutine (Child)
ShoulderRoutine (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.07 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Shoulder |  | 100 | 100 | 254 | 127 |  |  | On | 5 | On | CARE Dose4D | 8.44 | 7s | 25% |  |  | 1 |  | 2 | 0.8 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | Shoulder 5.0 B60s | [22 empty] | 5 | 5 | 27 | B60s sharp | Bone | [13 empty] | manual
|  | Shoulder 0.75 B31s | [22 empty] | 0.75 | 0.5 | 261 | B31s medium smooth + | Shoulder | [13 empty] | manual
|  | Shoulder 0.75 B60s | [22 empty] | 0.75 | 0.5 | 261 | B60s sharp | Bone | [13 empty] | manual
Thorax
Flash_Thorax (Child)
Flash_Thorax (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.07 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Fl_Thorax |  | 120 | 120 | 112 |  | 56 |  | On | 7 | On | CARE Dose4D | 6.2 | 5s | 25% |  |  | 0.285 |  | 2 | 3 |  | 3 | 128x0.6mm
|  | Fl_Thorax 3.0 B30f | [22 empty] | 3 | 3 | 51 | B30f medium smooth | Mediastinum
|  | Fl_Thorax 3.0 B60f | [22 empty] | 3 | 3 | 51 | B60f sharp | Baby Lung
|  | Fl_Thorax 1.5 B30f | [22 empty] | 1.5 | 1 | 151 | B30f medium smooth | Mediastinum
|  | Fl_Thorax 1.5 B60f | [22 empty] | 1.5 | 1 | 151 | B60f sharp | Baby Lung
Flash_Thorax_IR (Child)
Flash_Thorax_IR (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.07 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Fl_Thorax |  | 120 | 120 | 68 |  | 34 |  | On | 7 | On | CARE Dose4D | 3.76 | 5s | 25% |  |  | 0.285 |  | 2 | 3 |  | 3 | 128x0.6mm
|  | Fl_Thorax 3.0 I30f 3 | [22 empty] | 3 | 3 | 51 | I30f medium smooth | Mediastinum
|  | Fl_Thorax 3.0 I50f 3 | [22 empty] | 3 | 3 | 51 | I50f medium sharp ASA | Baby Lung
|  | Fl_Thorax 1.5 I30f 3 | [22 empty] | 1.5 | 1 | 151 | I30f medium smooth | Mediastinum
|  | Fl_Thorax 1.5 I50f 3 | [22 empty] | 1.5 | 1 | 151 | I50f medium sharp ASA | Baby Lung
ThoraxRoutine (Child)
ThoraxRoutine (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.07 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Thorax |  | 100 | 100 | 155 | 78 |  |  | On | 7 | On | CARE Dose4D | 5.28 | 5s | 25% |  |  | 0.5 |  | 2 | 1.4 |  | 3 | 32x1.2mm | [15 empty] | Off
|  | Thorax 3.0 B30f | [22 empty] | 3 | 3 | 51 | B30f medium smooth | Mediastinum | [13 empty] | manual
|  | Thorax 3.0 B60f | [22 empty] | 3 | 3 | 51 | B60f sharp | Baby Lung | [13 empty] | manual
|  | Thorax 1.5 B30f | [22 empty] | 1.5 | 1 | 151 | B30f medium smooth | Mediastinum | [13 empty] | manual
|  | Thorax 1.5 B60f | [22 empty] | 1.5 | 1 | 151 | B60f sharp | Baby Lung | [13 empty] | manual
ThoraxSeqHR (Child)
ThoraxSeqHR (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.07 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| ThorHRSeq |  | 120 | 120 | 110 | 55 |  |  | On | 3 | On | CARE Dose4D | 1.24 |  | 25% |  |  | 0.5 |  | 2 | 0 | 10 | 1 | 2x1.0mm | [15 empty] | Off
|  | ThorHRSeq 1.0 B70s | [22 empty] | 1 |  | 32 | B70s very sharp | Baby Lung | [13 empty] | manual
Thorax_IR (Child)
Thorax_IR (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.07 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Thorax |  | 100 | 100 | 93 | 47 |  |  | On | 7 | On | CARE Dose4D | 3.17 | 5s | 25% |  |  | 0.5 |  | 2 | 1.4 |  | 3 | 32x1.2mm | [15 empty] | Off
|  | Thorax 3.0 I30f 3 | [22 empty] | 3 | 3 | 51 | I30f medium smooth | Mediastinum | [13 empty] | manual
|  | Thorax 3.0 I50f 3 | [22 empty] | 3 | 3 | 51 | I50f medium sharp ASA | Baby Lung | [13 empty] | manual
|  | Thorax 1.5 I30f 3 | [22 empty] | 1.5 | 1 | 151 | I30f medium smooth | Mediastinum | [13 empty] | manual
|  | Thorax 1.5 I50f 3 | [22 empty] | 1.5 | 1 | 151 | I50f medium sharp ASA | Baby Lung | [13 empty] | manual
Thorax_XCARE (Child)
Thorax_XCARE (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.07 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Thorax |  | 100 | 100 | 155 | 78 |  |  | On | 7 | On | CARE Dose4D | 5.29 | 5s | 25% |  |  | 0.285 |  | 2 | 0.6 |  | 3 | 32x1.2mm | [15 empty] | Off
|  | Thorax 3.0 B30f | [22 empty] | 3 | 3 | 51 | B30f medium smooth | Mediastinum |  |  | X-CARE | [10 empty] | manual
|  | Thorax 3.0 B60f | [22 empty] | 3 | 3 | 51 | B60f sharp | Baby Lung |  |  | X-CARE | [10 empty] | manual
|  | Thorax 1.5 B30f | [22 empty] | 1.5 | 1 | 151 | B30f medium smooth | Mediastinum |  |  | X-CARE | [10 empty] | manual
|  | Thorax 1.5 B60f | [22 empty] | 1.5 | 1 | 151 | B60f sharp | Baby Lung |  |  | X-CARE | [10 empty] | manual
Thorax_XCARE_IR (Child)
Thorax_XCARE_IR (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.07 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Thorax |  | 100 | 100 | 93 | 47 |  |  | On | 7 | On | CARE Dose4D | 3.17 | 5s | 25% |  |  | 0.285 |  | 2 | 0.6 |  | 3 | 32x1.2mm | [15 empty] | Off
|  | Thorax 3.0 I30f 3 | [22 empty] | 3 | 3 | 51 | I30f medium smooth | Mediastinum |  |  | X-CARE | [10 empty] | manual
|  | Thorax 3.0 I50f 3 | [22 empty] | 3 | 3 | 51 | I50f medium sharp ASA | Baby Lung |  |  | X-CARE | [10 empty] | manual
|  | Thorax 1.5 I30f 3 | [22 empty] | 1.5 | 1 | 151 | I30f medium smooth | Mediastinum |  |  | X-CARE | [10 empty] | manual
|  | Thorax 1.5 I50f 3 | [22 empty] | 1.5 | 1 | 151 | I50f medium sharp ASA | Baby Lung |  |  | X-CARE | [10 empty] | manual
Abdomen
AbdomenRoutine (Child)
AbdomenRoutine (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.07 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Abdomen |  | 100 | 100 | 296 | 148 |  |  | On | 7 | On | CARE Dose4D | 7.9 | 30s | 25% |  |  | 0.5 |  | 2 | 1.4 |  | 5 | 32x1.2mm | [15 empty] | Off
|  | Abdomen 5.0 B30f | [22 empty] | 5 | 5 | 41 | B30f medium smooth | Baby Abdomen | [13 empty] | manual
Abdomen_IR (Child)
Abdomen_IR (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.07 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Abdomen |  | 100 | 100 | 207 | 104 |  |  | On | 7 | On | CARE Dose4D | 7.05 | 30s | 25% |  |  | 0.5 |  | 2 | 1.4 |  | 5 | 32x1.2mm | [15 empty] | Off
|  | Abdomen 5.0 I30f 3 | [22 empty] | 5 | 5 | 41 | I30f medium smooth | Baby Abdomen | [13 empty] | manual
Flash_Abdomen (Child)
Flash_Abdomen (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.07 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Fl_Abdomen |  | 120 | 120 | 212 |  | 106 |  | On | 7 | On | CARE Dose4D | 8.4 | 30s | 25% |  |  | 0.285 |  | 2 | 3 |  | 5 | 128x0.6mm
|  | Fl_Abdomen 5.0 B30f | [22 empty] | 5 | 5 | 41 | B30f medium smooth | Baby Abdomen
Flash_Abdomen_IR (Child)
Flash_Abdomen_IR (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.07 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Fl_Abdomen |  | 120 | 120 | 148 |  | 74 |  | On | 7 | On | CARE Dose4D | 8.18 | 30s | 25% |  |  | 0.285 |  | 2 | 3 |  | 5 | 128x0.6mm
|  | Fl_Abdomen 5.0 I30f 3 | [22 empty] | 5 | 5 | 41 | I30f medium smooth | Baby Abdomen
Spine
SpineRoutine (Child)
SpineRoutine (Child) | Topogram | Topogram 0.6 T80f |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.07 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Body
| Spine |  | 120 | 120 | 330 | 165 |  |  | On | 3 | On | CARE Dose4D | 18.26 | 40s | 25% |  |  | 1 |  | 2 | 0.8 |  | 2 | 128x0.6mm | [15 empty] | Off
|  | Spine 2.0 B30s | [22 empty] | 2 | 2 | 81 | B30s medium smooth | Baby Spine | [13 empty] | manual
|  | Spine 0.6 B60s | [22 empty] | 0.6 | 0.4 | 401 | B60s sharp | Baby Spine | [13 empty] | manual
SpineRoutine_Vol (Child)
SpineRoutine_Vol (Child) | Topogram | Topogram 0.6 T80f |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.07 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Body
| Spine |  | 120 | 120 | 330 | 165 |  |  | On | 3 | On | CARE Dose4D | 18.26 | 40s | 25% |  |  | 1 |  | 2 | 0.8 |  | 2 | 128x0.6mm | [15 empty] | Off
|  | Spine 2.0 B30s | [22 empty] | 2 | 2 | 81 | B30s medium smooth | Baby Spine | [13 empty] | manual
|  | Spine 0.6 MPR spine multi | [22 empty] | 0.6 | 0.4 | 0 | B30s medium smooth | Baby Spine | [13 empty] | manual
|  | Spine 0.6 MPR spine multi | [22 empty] | 0.6 | 0.4 | 0 | B60s sharp | Baby Spine | [13 empty] | manual
Spine_IR (Child)
Spine_IR (Child) | Topogram | Topogram 0.6 T80f |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.07 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Body
| Spine |  | 120 | 120 | 231 | 116 |  |  | On | 3 | On | CARE Dose4D | 12.79 | 40s | 25% |  |  | 1 |  | 2 | 0.8 |  | 2 | 128x0.6mm | [15 empty] | Off
|  | Spine 2.0 I30s 3 | [22 empty] | 2 | 2 | 81 | I30s medium smooth | Baby Spine | [13 empty] | manual
|  | Spine 0.6 I50s 3 | [22 empty] | 0.6 | 0.4 | 401 | I50s medium sharp ASA | Baby Spine | [13 empty] | manual
Spine_IR_Vol (Child)
Spine_IR_Vol (Child) | Topogram | Topogram 0.6 T80f |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.07 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Body
| Spine |  | 120 | 120 | 231 | 116 |  |  | On | 3 | On | CARE Dose4D | 12.79 | 40s | 25% |  |  | 1 |  | 2 | 0.8 |  | 2 | 128x0.6mm | [15 empty] | Off
|  | Spine 2.0 I30s 3 | [22 empty] | 2 | 2 | 81 | I30s medium smooth | Baby Spine | [13 empty] | manual
|  | Spine 0.6 MPR spine multi | [22 empty] | 0.6 | 0.4 | 0 | I30s medium smooth | Baby Spine | [13 empty] | manual
|  | Spine 0.6 MPR spine multi | [22 empty] | 0.6 | 0.4 | 0 | I50s medium sharp ASA | Baby Spine | [13 empty] | manual
Pelvis
PelvisRoutine (Child)
PelvisRoutine (Child) | Topogram | Topogram 0.6 T80f |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.07 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Body
| Pelvis |  | 100 | 100 | 296 | 148 |  |  | On | 7 | On | CARE Dose4D | 10.08 | 30s | 25% |  |  | 0.5 |  | 2 | 0.8 |  | 5 | 32x1.2mm | [15 empty] | Off
|  | Pelvis 5.0 B31f | [22 empty] | 5 | 5 | 31 | B31f medium smooth + | Pelvis | [13 empty] | manual
Pelvis_IR (Child)
Pelvis_IR (Child) | Topogram | Topogram 0.6 T80f |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.07 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Body
| Pelvis |  | 100 | 100 | 207 | 104 |  |  | On | 7 | On | CARE Dose4D | 7.06 | 30s | 25% |  |  | 0.5 |  | 2 | 0.8 |  | 5 | 32x1.2mm | [15 empty] | Off
|  | Pelvis 5.0 I31f 3 | [22 empty] | 5 | 5 | 31 | I31f medium smooth | Pelvis | [13 empty] | manual
Specials
Flash_NeonateBody (Child)
Flash_NeonateBody (Child) | Topogram | Topogram 0.6 T20f |  | 80 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.03 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Fl_Neonate |  | 80 | 80 | 276 |  | 94 |  | On | 7 | On | CARE Dose4D | 1.94 | 5s | 25% |  |  | 0.285 |  | 2 | 3 |  | 5 | 128x0.6mm
|  | Fl_Neonate 5.0 B30f | [22 empty] | 5 | 5 | 31 | B30f medium smooth | Baby Abdomen
Flash_NeonateBody_70kV (Child)
Flash_NeonateBody_70kV (Child) | Topogram | Topogram 0.6 T20f |  | 70 |  | 164 |  |  |  | 3 | Off | CARE Dose | 0.03 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Fl_Neonate |  | 70 | 70 | 508 |  | 172 |  | On | 7 | On | CARE Dose4D | 1.79 | 5s | 25% |  |  | 0.285 |  | 2 | 1.55 |  | 5 | 128x0.6mm
|  | Fl_Neonate 5.0 B30f | [22 empty] | 5 | 5 | 31 | B30f medium smooth | Baby Abdomen
Flash_NeonateBody_IR (Child)
Flash_NeonateBody_IR (Child) | Topogram | Topogram 0.6 T20f |  | 80 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.03 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Fl_Neonate |  | 80 | 80 | 166 |  | 56 |  | On | 7 | On | CARE Dose4D | 1.94 | 5s | 25% |  |  | 0.285 |  | 2 | 3 |  | 5 | 128x0.6mm
|  | Fl_Neonate 5.0 I30f 3 | [22 empty] | 5 | 5 | 31 | I30f medium smooth | Baby Abdomen
NeonateBody (Child)
NeonateBody (Child) | Topogram | Topogram 0.6 T20f |  | 80 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.03 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| NeonateBody |  | 80 | 80 | 276 | 94 |  |  | On | 7 | On | CARE Dose4D | 3.76 | 5s | 25% |  |  | 0.5 |  | 2 | 1.4 |  | 5 | 32x1.2mm | [15 empty] | Off
|  | NeonateBody 5.0 B30f | [22 empty] | 5 | 5 | 31 | B30f medium smooth | Baby Abdomen | [13 empty] | manual
NeonateBody_70kV (Child)
NeonateBody_70kV (Child) | Topogram | Topogram 0.6 T20f |  | 70 |  | 164 |  |  |  | 3 | Off | CARE Dose | 0.03 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| NeonateBody |  | 70 | 70 | 507 | 172 |  |  | On | 7 | On | CARE Dose4D | 1.78 | 5s | 25% |  |  | 0.5 |  | 2 | 1.4 |  | 5 | 32x1.2mm | [15 empty] | Off
|  | NeonateBody 5.0 B30f | [22 empty] | 5 | 5 | 31 | B30f medium smooth | Baby Abdomen | [13 empty] | manual
NeonateBody_IR (Child)
NeonateBody_IR (Child) | Topogram | Topogram 0.6 T20f |  | 80 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.03 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| NeonateBody |  | 80 | 80 | 166 | 56 |  |  | On | 7 | On | CARE Dose4D | 2.69 | 5s | 25% |  |  | 0.5 |  | 2 | 1.4 |  | 5 | 32x1.2mm | [15 empty] | Off
|  | NeonateBody 5.0 I30f 3 | [22 empty] | 5 | 5 | 31 | I30f medium smooth | Baby Abdomen | [13 empty] | manual
UpperExtremities
Extremity (Child)
Extremity (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.07 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Extremity |  |  | 100 | 100 | 35 |  |  |  | 3 | On | CARE Dose | 1.16 | 6s | 25% |  |  | 1 |  | 2 | 0.8 |  | 3 | 128x0.6mm | [15 empty] | Off
|  | Extremity 3.0 B30s | [22 empty] | 3 | 3 | 34 | B30s medium smooth | Baby Extremity | [13 empty] | manual
|  | Extremity 1.0 B30s | [22 empty] | 1 | 0.7 | 143 | B30s medium smooth | Baby Extremity | [13 empty] | manual
ExtrRoutineZUHR (Child)
ExtrRoutineZUHR (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.07 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| ExtrUHR |  |  | 120 | 180 | 90 |  |  |  | 3 | On | CARE Dose | 8.88 | 20s | 25% |  |  | 1 |  | 2 | 0.85 |  | 0.4 | 16x0.3mm | [15 empty] | Off
|  | ExtrUHR 0.4 U80u | [22 empty] | 0.4 | 0.4 | 151 | U80u very sharp | Baby Extremity | [13 empty] | manual
ExtrUHR (Child)
ExtrUHR (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.07 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| ExtrUHR |  |  | 120 | 120 | 42 |  |  |  | 3 | On | CARE Dose | 4.15 | 20s | 25% |  |  | 1 |  | 2 | 0.85 |  | 0.6 | 16x0.6mm | [15 empty] | Off
|  | ExtrUHR 0.6 U80u | [22 empty] | 0.6 | 0.6 | 101 | U80u very sharp | Baby Extremity | [13 empty] | manual
ExtrZUHR_IR (Child)
ExtrZUHR_IR (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.07 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| ExtrUHR |  |  | 120 | 126 | 65 |  |  |  | 3 | On | CARE Dose | 6.46 | 20s | 25% |  |  | 1 |  | 2 | 0.85 |  | 0.4 | 16x0.3mm | [15 empty] | Off
|  | ExtrUHR 0.4 V80u 3 | [22 empty] | 0.4 | 0.4 | 151 | V80u very sharp | Baby Extremity | [13 empty] | manual
LowerExtremities
Extremity (Child)
Extremity (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.07 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Extremity |  |  | 100 | 100 | 35 |  |  |  | 3 | On | CARE Dose | 1.16 | 6s | 25% |  |  | 1 |  | 2 | 0.8 |  | 3 | 128x0.6mm | [15 empty] | Off
|  | Extremity 3.0 B30s | [22 empty] | 3 | 3 | 34 | B30s medium smooth | Baby Extremity | [13 empty] | manual
|  | Extremity 1.0 B30s | [22 empty] | 1 | 0.7 | 143 | B30s medium smooth | Baby Extremity | [13 empty] | manual
ExtrRoutineZUHR (Child)
ExtrRoutineZUHR (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.07 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| ExtrUHR |  |  | 120 | 180 | 90 |  |  |  | 3 | On | CARE Dose | 8.88 | 20s | 25% |  |  | 1 |  | 2 | 0.85 |  | 0.4 | 16x0.3mm | [15 empty] | Off
|  | ExtrUHR 0.4 U80u | [22 empty] | 0.4 | 0.4 | 151 | U80u very sharp | Baby Extremity | [13 empty] | manual
ExtrUHR (Child)
ExtrUHR (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.07 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| ExtrUHR |  |  | 120 | 120 | 42 |  |  |  | 3 | On | CARE Dose | 4.15 | 20s | 25% |  |  | 1 |  | 2 | 0.85 |  | 0.6 | 16x0.6mm | [15 empty] | Off
|  | ExtrUHR 0.6 U80u | [22 empty] | 0.6 | 0.6 | 101 | U80u very sharp | Baby Extremity | [13 empty] | manual
ExtrZUHR_IR (Child)
ExtrZUHR_IR (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.07 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| ExtrUHR |  |  | 120 | 126 | 65 |  |  |  | 3 | On | CARE Dose | 6.46 | 20s | 25% |  |  | 1 |  | 2 | 0.85 |  | 0.4 | 16x0.3mm | [15 empty] | Off
|  | ExtrUHR 0.4 V80u 3 | [22 empty] | 0.4 | 0.4 | 151 | V80u very sharp | Baby Extremity | [13 empty] | manual
Cardiac
Flash_Cardio (Child)
Flash_Cardio (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.07 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Fl_CorCTA |  | 80 | 80 | 594 |  | 298 |  | On | 11 | On | CARE Dose4D | 1.71 |  | 25% |  |  | 0.285 |  | 2 | 3.4 |  | 3 | 128x0.6mm
|  | Fl_CorCTA 3.0 B30f 30% | [22 empty] | 3 | 3 | 24 | B30f medium smooth | Cardiac | [14 empty] | 30%
|  | Fl_CorCTA 0.6 B26f 30% | [22 empty] | 0.6 | 0.3 | 234 | B26f medium smooth ASA | Cardiac | [14 empty] | 30%
Flash_Cardio_IR (Child)
Flash_Cardio_IR (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.07 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Fl_CorCTA |  | 80 | 80 | 416 |  | 298 |  | On | 11 | On | CARE Dose4D | 1.71 |  | 25% |  |  | 0.285 |  | 2 | 3.4 |  | 3 | 128x0.6mm
|  | Fl_CorCTA 3.0 I30f 3 30% | [22 empty] | 3 | 3 | 24 | I30f medium smooth | Cardiac | [14 empty] | 30%
|  | Fl_CorCTA 0.6 I26f 3 30% | [22 empty] | 0.6 | 0.3 | 234 | I26f medium smooth ASA | Cardiac | [14 empty] | 30%
Vascular
BodyAngio (Child)
BodyAngio (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.07 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| BodyAngio |  | 80 | 80 | 223 | 112 |  |  | On | 11 | On | CARE Dose4D | 3.51 | 5s | 25% |  |  | 0.5 |  | 2 | 1.4 |  | 3 | 128x0.6mm | [15 empty] | Off
|  | BodyAngio 3.0 B30f | [22 empty] | 3 | 3 | 67 | B30f medium smooth | Baby Abdomen | [13 empty] | manual
|  | BodyAngio 0.6 B26f | [22 empty] | 0.6 | 0.4 | 501 | B26f medium smooth ASA | CT Angio | [13 empty] | manual
BodyAngio_IR (Child)
BodyAngio_IR (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.07 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| BodyAngio |  | 80 | 80 | 156 | 78 |  |  | On | 11 | On | CARE Dose4D | 2.46 | 5s | 25% |  |  | 0.5 |  | 2 | 1.4 |  | 3 | 128x0.6mm | [15 empty] | Off
|  | BodyAngio 3.0 I30f 3 | [22 empty] | 3 | 3 | 67 | I30f medium smooth | Baby Abdomen | [13 empty] | manual
|  | BodyAngio 0.6 I26f 3 | [22 empty] | 0.6 | 0.4 | 501 | I26f medium smooth ASA | CT Angio | [13 empty] | manual
CarotidAngio (Child)
CarotidAngio (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.07 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| CarotidAngio |  | 80 | 80 | 223 | 112 |  |  | On | 11 | On | CARE Dose4D | 3.51 | 5s | 25% |  |  | 0.5 |  | 2 | 1.4 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | CarotidAngio 5.0 B30f | [22 empty] | 5 | 5 | 35 | B30f medium smooth | Baby Neck | [13 empty] | manual
|  | CarotidAngio 0.6 B26f | [22 empty] | 0.6 | 0.4 | 426 | B26f medium smooth ASA | CT Angio | [13 empty] | manual
CarotidAngio_IR (Child)
CarotidAngio_IR (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.07 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| CarotidAngio |  | 80 | 80 | 156 | 78 |  |  | On | 11 | On | CARE Dose4D | 2.46 | 5s | 25% |  |  | 0.5 |  | 2 | 1.4 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | CarotidAngio 5.0 I30f 3 | [22 empty] | 5 | 5 | 35 | I30f medium smooth | Baby Neck | [13 empty] | manual
|  | CarotidAngio 0.6 I26f 3 | [22 empty] | 0.6 | 0.4 | 426 | I26f medium smooth ASA | CT Angio | [13 empty] | manual
Flash_BodyAngio (Child)
Flash_BodyAngio (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.07 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Fl_BodyAngio |  | 80 | 80 | 224 |  | 112 |  | On | 11 | On | CARE Dose4D | 1.94 | 5s | 25% |  |  | 0.285 |  | 2 | 3 |  | 3 | 128x0.6mm
|  | Fl_BodyAngio 3.0 B30f | [22 empty] | 3 | 3 | 67 | B30f medium smooth | Baby Abdomen
|  | Fl_BodyAngio 0.6 B26f | [22 empty] | 0.6 | 0.4 | 501 | B26f medium smooth ASA | CT Angio
Flash_BodyAngio_IR (Child)
Flash_BodyAngio_IR (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.07 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Fl_BodyAngio |  | 80 | 80 | 158 |  | 80 |  | On | 11 | On | CARE Dose4D | 1.94 | 5s | 25% |  |  | 0.285 |  | 2 | 3 |  | 3 | 128x0.6mm
|  | Fl_BodyAngio 3.0 I30f 3 | [22 empty] | 3 | 3 | 67 | I30f medium smooth | Baby Abdomen
|  | Fl_BodyAngio 0.6 I26f 3 | [22 empty] | 0.6 | 0.4 | 501 | I26f medium smooth ASA | CT Angio
HeadAngio (Child)
HeadAngio (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.16 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head
| HeadAngio |  | 80 | 80 | 221 | 111 |  |  | On | 11 | On | CARE Dose4D | 8.43 | 5s | 25% |  |  | 0.5 |  | 2 | 1.4 |  | 2 | 128x0.6mm | [15 empty] | Off
|  | HeadAngio 2.0 H20f | [22 empty] | 2 | 2 | 31 | H20f smooth | Baby Brain | [13 empty] | manual
|  | HeadAngio 0.6 H10f | [22 empty] | 0.6 | 0.4 | 151 | H10f very smooth | CT Angio | [13 empty] | manual
HeadAngio_IR (Child)
HeadAngio_IR (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.16 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head
| HeadAngio |  | 80 | 80 | 155 | 78 |  |  | On | 11 | On | CARE Dose4D | 5.91 | 5s | 25% |  |  | 0.5 |  | 2 | 1.4 |  | 2 | 128x0.6mm | [15 empty] | Off
|  | HeadAngio 2.0 J30f 3 | [22 empty] | 2 | 2 | 31 | J30f medium smooth | Baby Brain | [13 empty] | manual
|  | HeadAngio 0.6 J30f 3 | [22 empty] | 0.6 | 0.4 | 151 | J30f medium smooth | CT Angio | [13 empty] | manual
HeadAngio_XCARE (Child)
HeadAngio_XCARE (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.16 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head
| HeadAngio |  | 80 | 80 | 221 | 111 |  |  | On | 11 | On | CARE Dose4D | 8.45 | 5s | 25% |  |  | 0.285 |  | 2 | 0.6 |  | 2 | 128x0.6mm | [15 empty] | Off
|  | HeadAngio 2.0 H20f | [22 empty] | 2 | 2 | 31 | H20f smooth | Baby Brain |  |  | X-CARE | [10 empty] | manual
|  | HeadAngio 0.6 H10f | [22 empty] | 0.6 | 0.4 | 151 | H10f very smooth | CT Angio |  |  | X-CARE | [10 empty] | manual
HeadAngio_XCARE_IR (Child)
HeadAngio_XCARE_IR (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.16 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head
| HeadAngio |  | 80 | 80 | 155 | 78 |  |  | On | 11 | On | CARE Dose4D | 5.93 | 5s | 25% |  |  | 0.285 |  | 2 | 0.6 |  | 2 | 128x0.6mm | [15 empty] | Off
|  | HeadAngio 2.0 J30f 3 | [22 empty] | 2 | 2 | 31 | J30f medium smooth | Baby Brain |  |  | X-CARE | [10 empty] | manual
|  | HeadAngio 0.6 J30f 3 | [22 empty] | 0.6 | 0.4 | 151 | J30f medium smooth | CT Angio |  |  | X-CARE | [10 empty] | manual
RT
RT_Abdomen (Child)
RT_Abdomen (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.07 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Abdomen |  | 120 | 120 | 250 | 125 |  |  | Semi | 3 | On | CARE Dose4D | 14.21 | 99s | 25% |  |  | 0.5 |  | 2 | 0.6 |  | 4 | 32x1.2mm | [15 empty] | Off
|  | Abdomen 4.0 B30f | [22 empty] | 4 | 4 | 51 | B30f medium smooth | Baby Abdomen | [13 empty] | manual
RT_Head (Child)
RT_Head (Child) | Topogram | Topogram 0.6 T20f |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.16 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head
| Head |  | 120 | 120 | 410 | 205 |  |  | Off | 3 | On | CARE Dose4D | 63.87 | 99s | 25% |  |  | 1 |  | 2 | 0.8 |  | 3 | 32x1.2mm | [15 empty] | Off
|  | Head 3.0 H30s | [22 empty] | 3 | 3 | 41 | H30s medium smooth | Baby Brain | [13 empty] | manual
Head
Dental (Adult)
Dental (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.29 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head
| Dental |  | 120 | 120 | 90 | 90 |  |  | Semi | 5 | On | CARE Dose4D | 12.69 | 7s | 25% |  |  | 1 |  | 2 | 0.8 |  | 0.75 | 128x0.6mm | [15 empty] | Off
|  | Dental 0.75 H60s | [22 empty] | 0.75 | 0.5 | 101 | H60s sharp FR | Dental | [13 empty] | manual
HeadIntervention (Adult)
HeadIntervention (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.29 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head
| i-Spiral |  | 120 | 120 | 420 | 420 |  |  | On | 3 | On | CARE Dose4D | 60.49 | 10s | 25% |  |  | 0.5 |  | 2 | 0.55 |  | 5 | 32x1.2mm | [15 empty] | Off
|  | i-Spiral 5.0 H32f | [22 empty] | 5 | 3.5 | 35 | H32f med. smooth FR | Base Orbita | [13 empty] | manual
HeadNeuro (Adult)
HeadNeuro (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.29 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head
| Head |  | 120 | 120 | 390 | 390 |  |  | On | 3 | On | CARE Dose4D | 73.37 | 40s | 25% |  |  | 1 |  | 2 | 0.55 |  | 5 | 40x0.6mm | [15 empty] | Off
|  | Head 5.0 H31s | [22 empty] | 5 | 5 | 25 | H31s medium smooth + | Base Orbita | [13 empty] | manual
HeadNeuroSeq (Adult)
HeadNeuroSeq (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.29 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head
| HeadSeq |  | 120 | 120 | 390 | 390 |  |  | On | 3 | On | CARE Dose4D | 71.55 |  | 25% |  |  | 1 |  | 2 | 0 | 14 | 4.8 | 12x1.2mm | [15 empty] | Off
|  | HeadSeq 4.8 H31s | [22 empty] | 4.8 |  | 27 | H31s medium smooth + | Base Orbita | [13 empty] | manual
HeadNeuro_XCARE (Adult)
HeadNeuro_XCARE (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.29 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head
| Head |  | 120 | 120 | 390 | 390 |  |  | On | 3 | On | CARE Dose4D | 73.37 | 40s | 25% |  |  | 1 |  | 2 | 0.55 |  | 5 | 40x0.6mm | [15 empty] | Off
|  | Head 5.0 H31s | [22 empty] | 5 | 5 | 25 | H31s medium smooth + | Base Orbita |  |  | X-CARE | [10 empty] | manual
HeadRoutine (Adult)
HeadRoutine (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.29 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head
| Head |  | 120 | 120 | 350 | 350 |  |  | On | 3 | On | CARE Dose4D | 53.64 | 15s | 25% |  |  | 1 |  | 2 | 0.55 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | Head 5.0 H31s | [22 empty] | 5 | 5 | 25 | H31s medium smooth + | Base Orbita | [13 empty] | manual
HeadRoutineSeq (Adult)
HeadRoutineSeq (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.29 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head
| HeadSeq |  | 120 | 120 | 390 | 390 |  |  | On | 3 | On | CARE Dose4D | 67.58 |  | 25% |  |  | 1 |  | 2 | 0 | 34 | 5 | 32x1.2mm | [15 empty] | Off
|  | HeadSeq 5.0 H31s | [22 empty] | 5 |  | 27 | H31s medium smooth + | Base Orbita | [13 empty] | manual
Head_IR (Adult)
Head_IR (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.29 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head
| Head |  | 120 | 120 | 273 | 273 |  |  | On | 3 | On | CARE Dose4D | 41.97 | 15s | 25% |  |  | 1 |  | 2 | 0.55 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | Head 5.0 J30s 3 | [22 empty] | 5 | 5 | 25 | J30s medium smooth | Base Orbita | [13 empty] | manual
Head_XCARE (Adult)
Head_XCARE (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.29 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head
| Head |  | 120 | 120 | 390 | 390 |  |  | On | 3 | On | CARE Dose4D | 59.76 | 15s | 25% |  |  | 1 |  | 2 | 0.55 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | Head 5.0 H31s | [22 empty] | 5 | 5 | 25 | H31s medium smooth + | Base Orbita |  |  | X-CARE | [10 empty] | manual
Head_XCARE_IR (Adult)
Head_XCARE_IR (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.29 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head
| Head |  | 120 | 120 | 273 | 273 |  |  | On | 3 | On | CARE Dose4D | 41.97 | 15s | 25% |  |  | 1 |  | 2 | 0.55 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | Head 5.0 J30s 3 | [22 empty] | 5 | 5 | 25 | J30s medium smooth | Base Orbita |  |  | X-CARE | [10 empty] | manual
InnerEar (Adult)
InnerEar (Adult) | Topogram | Topogram 0.6 T80f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.29 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Head
| InnerEar |  | 120 | 120 | 180 | 180 |  |  | Semi | 5 | On | CARE Dose4D | 27.51 | 6s | 25% |  |  | 1 |  | 2 | 0.8 |  | 2 | 128x0.6mm | [15 empty] | Off
|  | InnerEar 2.0 H60s | [22 empty] | 2 | 2 | 21 | H60s sharp FR | Inner Ear | [13 empty] | manual
|  | R InnerEar 0.6 H60s | [22 empty] | 0.6 | 0.4 | 101 | H60s sharp FR | Inner Ear | [13 empty] | manual
|  | L InnerEar 0.6 H60s | [22 empty] | 0.6 | 0.4 | 101 | H60s sharp FR | Inner Ear | [13 empty] | manual
InnerEarUHR (Adult)
InnerEarUHR (Adult) | Topogram | Topogram 0.6 T80f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.29 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Head
| InnerEarUHR |  | 120 | 120 | 180 | 180 |  |  | Semi | 5 | On | CARE Dose4D | 39.82 | 24s | 25% |  |  | 1 |  | 2 | 0.85 |  | 2 | 16x0.6mm | [15 empty] | Off
|  | InnerEarUHR 2.0 U75u | [22 empty] | 2 | 2 | 21 | U75u very sharp ASA | Inner Ear | [13 empty] | manual
|  | R InnerEarUHR 0.6 U75u | [22 empty] | 0.6 | 0.4 | 101 | U75u very sharp ASA | Inner Ear | [13 empty] | manual
|  | L InnerEarUHR 0.6 U75u | [22 empty] | 0.6 | 0.4 | 101 | U75u very sharp ASA | Inner Ear | [13 empty] | manual
InnerEarUHRSeq (Adult)
InnerEarUHRSeq (Adult) | Topogram | Topogram 0.6 T80f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.29 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Head
| InnEarUHRSeq |  | 120 | 120 | 180 | 180 |  |  | Semi | 5 | On | CARE Dose4D | 41.65 |  | 25% |  |  | 1 |  | 2 | 0 | 4 | 2 | 16x0.6mm | [15 empty] | Off
|  | InnEarUHRSeq 2.0 U75u | [22 empty] | 2 |  | 20 | U75u very sharp ASA | Inner Ear | [13 empty] | manual
|  | InnEarUHRSeq 0.6 U75u | [22 empty] | 0.6 |  | 101 | U75u very sharp ASA | Inner Ear | [13 empty] | manual
InnerEarZUHR (Adult)
InnerEarZUHR (Adult) | Topogram | Topogram 0.6 T80f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.29 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Head
| InnerEarUHR |  | 120 | 120 | 230 | 230 |  |  | Semi | 5 | On | CARE Dose4D | 51.02 | 24s | 25% |  |  | 1 |  | 2 | 0.85 |  | 2 | 16x0.3mm | [15 empty] | Off
|  | InnerEarUHR 2.0 U75u | [22 empty] | 2 | 2 | 21 | U75u very sharp ASA | Inner Ear | [13 empty] | manual
|  | R InnerEarUHR 0.4 U75u | [22 empty] | 0.4 | 0.2 | 201 | U75u very sharp ASA | Inner Ear | [13 empty] | manual
|  | L InnerEarUHR 0.4 U75u | [22 empty] | 0.4 | 0.2 | 201 | U75u very sharp ASA | Inner Ear | [13 empty] | manual
InnerEarZUHR_IR (Adult)
InnerEarZUHR_IR (Adult) | Topogram | Topogram 0.6 T80f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.29 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Head
| InnerEarUHR |  | 120 | 120 | 161 | 161 |  |  | Semi | 5 | On | CARE Dose4D | 35.66 | 24s | 25% |  |  | 1 |  | 2 | 0.85 |  | 2 | 16x0.3mm | [15 empty] | Off
|  | InnerEarUHR 2.0 V80u 3 | [22 empty] | 2 | 2 | 21 | V80u very sharp | Inner Ear | [13 empty] | manual
|  | R InnerEarUHR 0.4 V80u 3 | [22 empty] | 0.4 | 0.2 | 201 | V80u very sharp | Inner Ear | [13 empty] | manual
|  | L InnerEarUHR 0.4 V80u 3 | [22 empty] | 0.4 | 0.2 | 201 | V80u very sharp | Inner Ear | [13 empty] | manual
Orbit (Adult)
Orbit (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.29 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head
| Orbit |  | 120 | 120 | 125 | 125 |  |  | Semi | 5 | On | CARE Dose4D | 17.62 | 10s | 25% |  |  | 1 |  | 2 | 0.8 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | Orbit 5.0 H60s | [22 empty] | 5 | 5 | 11 | H60s sharp FR | Base Orbita | [13 empty] | manual
|  | Orbit 5.0 H30s | [22 empty] | 5 | 5 | 11 | H30s medium smooth | Base Orbita | [13 empty] | manual
|  | Orbit 0.6 H60s | [22 empty] | 0.6 | 0.4 | 126 | H60s sharp FR | Base Orbita | [13 empty] | manual
Sinus (Adult)
Sinus (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.29 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head
| Sinus |  | 120 | 120 | 50 | 50 |  |  | On | 5 | On | CARE Dose4D | 7.05 | 10s | 25% |  |  | 1 |  | 2 | 0.8 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | Sinus 5.0 H60s | [22 empty] | 5 | 5 | 17 | H60s sharp FR | Sinuses | [13 empty] | manual
|  | Sinus 5.0 H30s | [22 empty] | 5 | 5 | 17 | H30s medium smooth | Sinuses | [13 empty] | manual
|  | Sinus 1.0 H60s | [22 empty] | 1 | 0.7 | 115 | H60s sharp FR | Sinuses | [13 empty] | manual
|  | Sinus 1.0 H30s | [22 empty] | 1 | 0.7 | 115 | H30s medium smooth | Sinuses | [13 empty] | manual
Sinus_Vol (Adult)
Sinus_Vol (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.29 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head
| Sinus |  | 120 | 120 | 50 | 50 |  |  | On | 5 | On | CARE Dose4D | 7.05 | 10s | 25% |  |  | 1 |  | 2 | 0.8 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | Sinus 5.0 H60s | [22 empty] | 5 | 5 | 17 | H60s sharp FR | Sinuses | [13 empty] | manual
|  | Sinus 3.0 MPR cor | [22 empty] | 3 | 3 | 37 | H30s medium smooth | Sinuses | [13 empty] | manual
|  | Sinus 3.0 MPR cor | [22 empty] | 3 | 3 | 37 | H60s sharp FR | Sinuses | [13 empty] | manual
Neck
NeckRoutine (Adult)
NeckRoutine (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Neck |  | 120 | 120 | 165 | 165 |  |  | On | 7 | On | CARE Dose4D | 11.14 | 9s | 25% |  |  | 1 |  | 2 | 0.8 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | Neck 5.0 B31s | [22 empty] | 5 | 5 | 37 | B31s medium smooth + | Larynx | [13 empty] | manual
|  | Neck 1.0 B20s | [22 empty] | 1 | 0.7 | 258 | B20s smooth | Larynx | [13 empty] | manual
NeckThorAbd (Adult)
NeckThorAbd (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 365 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Neck |  | 120 | 120 | 165 | 165 |  |  | On | 7 | On | CARE Dose4D | 11.14 | 7s | 25% |  |  | 1 |  | 2 | 0.8 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | Neck 5.0 B31s | [22 empty] | 5 | 5 | 29 | B31s medium smooth + | Larynx | [13 empty] | manual
|  | Neck 1.0 B20s | [22 empty] | 1 | 0.7 | 201 | B20s smooth | Larynx | [13 empty] | manual
| ThorAbd |  | 120 | 120 | 210 | 210 |  |  | On | 7 | On | CARE Dose4D | 14.18 | 40s | 25% |  |  | 0.5 |  | 2 | 0.8 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | Tho ThorAbd 5.0 B31f | [22 empty] | 5 | 5 | 41 | B31f medium smooth + | Mediastinum | [13 empty] | manual
|  | Lun ThorAbd 5.0 B80f | [22 empty] | 5 | 5 | 41 | B80f ultra sharp | Lung | [13 empty] | manual
|  | Abd ThorAbd 5.0 B30f | [22 empty] | 5 | 5 | 81 | B30f medium smooth | Abdomen | [13 empty] | manual
|  | MPR ThorAbd 1.0 B30f | [22 empty] | 1 | 0.7 | 858 | B30f medium smooth | Mediastinum | [13 empty] | manual
NeckThorax (Adult)
NeckThorax (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Neck |  | 120 | 120 | 165 | 165 |  |  | On | 7 | On | CARE Dose4D | 11.14 | 7s | 25% |  |  | 1 |  | 2 | 0.8 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | Neck 5.0 B31s | [22 empty] | 5 | 5 | 29 | B31s medium smooth + | Larynx | [13 empty] | manual
|  | Neck 1.0 B20s | [22 empty] | 1 | 0.7 | 201 | B20s smooth | Larynx | [13 empty] | manual
| Thorax |  | 120 | 120 | 110 | 110 |  |  | On | 7 | On | CARE Dose4D | 7.43 | 7s | 25% |  |  | 0.5 |  | 2 | 0.8 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | Thorax 5.0 B31f | [22 empty] | 5 | 5 | 61 | B31f medium smooth + | Mediastinum | [13 empty] | manual
|  | Thorax 5.0 B70f | [22 empty] | 5 | 5 | 61 | B70f very sharp | Lung | [13 empty] | manual
|  | Thorax 1.0 B80f | [22 empty] | 1 | 0.7 | 429 | B80f ultra sharp | Lung | [13 empty] | manual
Neck_IR (Adult)
Neck_IR (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Neck |  | 120 | 120 | 116 | 116 |  |  | On | 7 | On | CARE Dose4D | 7.85 | 9s | 25% |  |  | 1 |  | 2 | 0.8 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | Neck 5.0 I31s 3 | [22 empty] | 5 | 5 | 37 | I31s medium smooth | Larynx | [13 empty] | manual
|  | Neck 1.0 I26s 3 | [22 empty] | 1 | 0.7 | 258 | I26s medium smooth ASA | Larynx | [13 empty] | manual
Shoulder
ShoulderRoutine (Adult)
ShoulderRoutine (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Shoulder |  | 120 | 120 | 165 | 165 |  |  | On | 5 | On | CARE Dose4D | 11.14 | 8s | 25% |  |  | 1 |  | 2 | 0.8 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | Shoulder 5.0 B60s | [22 empty] | 5 | 5 | 31 | B60s sharp | Bone | [13 empty] | manual
|  | Shoulder 0.6 B31s | [22 empty] | 0.6 | 0.4 | 376 | B31s medium smooth + | Shoulder | [13 empty] | manual
|  | Shoulder 0.6 B60s | [22 empty] | 0.6 | 0.4 | 376 | B60s sharp | Bone | [13 empty] | manual
Thorax
DSXXL_Thorax (Adult)
DSXXL_Thorax (Adult) | Topogram | Topogram 0.6 T20f |  | 140 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.2 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
Flash_Thorax (Adult)
Flash_Thorax (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Fl_Thorax |  | 120 | 120 | 110 |  | 110 |  | On | 7 | On | CARE Dose4D | 6.08 | 5s | 25% |  |  | 0.285 |  | 2 | 3 |  | 5 | 128x0.6mm
|  | Fl_Thorax 5.0 B31f | [22 empty] | 5 | 5 | 61 | B31f medium smooth + | Mediastinum
|  | Fl_Thorax 5.0 B80f | [22 empty] | 5 | 5 | 61 | B80f ultra sharp | Lung
|  | Fl_Thorax 1.0 B31f | [22 empty] | 1 | 0.7 | 429 | B31f medium smooth + | Mediastinum
|  | Fl_Thorax 1.0 B70f | [22 empty] | 1 | 0.7 | 429 | B70f very sharp | Lung
Flash_Thorax_IR (Adult)
Flash_Thorax_IR (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Fl_Thorax |  | 120 | 120 | 68 |  | 68 |  | On | 7 | On | CARE Dose4D | 3.76 | 5s | 25% |  |  | 0.285 |  | 2 | 3 |  | 5 | 128x0.6mm
|  | Fl_Thorax 5.0 I31f 3 | [22 empty] | 5 | 5 | 61 | I31f medium smooth | Mediastinum
|  | Fl_Thorax 5.0 I70f 3 | [22 empty] | 5 | 5 | 61 | I70f very sharp ASA | Lung
|  | Fl_Thorax 1.0 I31f 3 | [22 empty] | 1 | 0.7 | 429 | I31f medium smooth | Mediastinum
|  | Fl_Thorax 1.0 I70f 3 | [22 empty] | 1 | 0.7 | 429 | I70f very sharp ASA | Lung
LungCARE (Adult)
LungCARE (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| LungCARE |  | 120 | 120 | 20 | 20 |  |  | Semi | 1 | On | CARE Dose4D | 1.35 | 10s | 25% |  |  | 0.5 |  | 2 | 1.2 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | LungCARE 5.0 B31f | [22 empty] | 5 | 5 | 61 | B31f medium smooth + | Mediastinum | [13 empty] | manual
|  | LungCARE 5.0 B80f | [22 empty] | 5 | 5 | 61 | B80f ultra sharp | Lung | [13 empty] | manual
|  | LungCARE 1.0 B31f | [22 empty] | 1 | 0.7 | 429 | B31f medium smooth + | Mediastinum | [13 empty] | manual
|  | LungCARE 1.0 B60f | [22 empty] | 1 | 0.7 | 429 | B60f sharp | Lung | [13 empty] | manual
LungLowDose (Adult)
LungLowDose (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| LungLowDose |  | 120 | 120 | 20 | 20 |  |  | On | 1 | On | CARE Dose4D | 1.35 | 10s | 25% |  |  | 0.5 |  | 2 | 1.2 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | LungLowDose 5.0 B31f | [22 empty] | 5 | 5 | 61 | B31f medium smooth + | Mediastinum | [13 empty] | manual
|  | LungLowDose 5.0 B80f | [22 empty] | 5 | 5 | 61 | B80f ultra sharp | Lung | [13 empty] | manual
|  | LungLowDose 1.0 B31f | [22 empty] | 1 | 0.7 | 429 | B31f medium smooth + | Mediastinum | [13 empty] | manual
|  | LungLowDose 1.0 B70f | [22 empty] | 1 | 0.7 | 429 | B70f very sharp | Lung | [13 empty] | manual
ThorAbd (Adult)
ThorAbd (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 275 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| ThorAbd |  | 120 | 120 | 210 | 210 |  |  | On | 7 | On | CARE Dose4D | 14.19 | 40s | 25% |  |  | 0.5 |  | 2 | 0.6 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | Tho ThorAbd 5.0 B31f | [22 empty] | 5 | 5 | 41 | B31f medium smooth + | Mediastinum | [13 empty] | manual
|  | Lun ThorAbd 5.0 B80f | [22 empty] | 5 | 5 | 41 | B80f ultra sharp | Lung | [13 empty] | manual
|  | Abd ThorAbd 5.0 B30f | [22 empty] | 5 | 5 | 81 | B30f medium smooth | Abdomen | [13 empty] | manual
|  | MPR ThorAbd 1.0 B30f | [22 empty] | 1 | 0.7 | 858 | B30f medium smooth | Mediastinum | [13 empty] | manual
ThorAbd_XCARE (Adult)
ThorAbd_XCARE (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 275 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| ThorAbd |  | 120 | 120 | 210 | 210 |  |  | On | 7 | On | CARE Dose4D | 14.19 | 40s | 25% |  |  | 0.5 |  | 2 | 0.6 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | Tho ThorAbd 5.0 B31f | [22 empty] | 5 | 5 | 41 | B31f medium smooth + | Mediastinum |  |  | X-CARE | [10 empty] | manual
|  | Lun ThorAbd 5.0 B80f | [22 empty] | 5 | 5 | 41 | B80f ultra sharp | Lung |  |  | X-CARE | [10 empty] | manual
|  | Abd ThorAbd 5.0 B30f | [22 empty] | 5 | 5 | 81 | B30f medium smooth | Abdomen |  |  | X-CARE | [10 empty] | manual
|  | MPR ThorAbd 1.0 B30f | [22 empty] | 1 | 0.7 | 858 | B30f medium smooth | Mediastinum |  |  | X-CARE | [10 empty] | manual
ThoraxECGSeqHR (Adult)
ThoraxECGSeqHR (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| ThorECGHR |  | 120 | 120 | 110 | 110 |  |  | On | 3 | On | CARE Dose4D | 0.9 |  | 25% |  |  | 0.285 |  | 2 | 0 | 10 | 1 | 2x1.0mm | [13 empty] | 70% | 70% | Off
|  | ThorECGHR 1.0 B80f 70% | [22 empty] | 1 |  | 60 | B80f ultra sharp | Lung | [13 empty] | manual | 70%
ThoraxHR (Adult)
ThoraxHR (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| ThorHR |  | 120 | 120 | 110 | 110 |  |  | On | 3 | On | CARE Dose4D | 7.43 | 10s | 25% |  |  | 0.5 |  | 2 | 1.2 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | ThorHR 5.0 B80f | [22 empty] | 5 | 5 | 61 | B80f ultra sharp | Lung | [13 empty] | manual
|  | ThorHR 5.0 B31f | [22 empty] | 5 | 5 | 61 | B31f medium smooth + | Mediastinum | [13 empty] | manual
|  | ThorHR 0.6 B70f | [22 empty] | 0.6 | 0.4 | 751 | B70f very sharp | Lung | [13 empty] | manual
|  | ThorHR 0.6 B31f | [22 empty] | 0.6 | 0.4 | 751 | B31f medium smooth + | Mediastinum | [13 empty] | manual
ThoraxHR_IR (Adult)
ThoraxHR_IR (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| ThorHR |  | 120 | 120 | 66 | 66 |  |  | On | 3 | On | CARE Dose4D | 4.48 | 10s | 25% |  |  | 0.5 |  | 2 | 1.2 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | ThorHR 5.0 I70f 3 | [22 empty] | 5 | 5 | 61 | I70f very sharp ASA | Lung | [13 empty] | manual
|  | ThorHR 5.0 I31f 3 | [22 empty] | 5 | 5 | 61 | I31f medium smooth | Mediastinum | [13 empty] | manual
|  | ThorHR 0.6 I70f 3 | [22 empty] | 0.6 | 0.4 | 751 | I70f very sharp ASA | Lung | [13 empty] | manual
|  | ThorHR 0.6 I31f 3 | [22 empty] | 0.6 | 0.4 | 751 | I31f medium smooth | Mediastinum | [13 empty] | manual
ThoraxIntervention (Adult)
ThoraxIntervention (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| i-Spiral |  | 120 | 120 | 110 | 110 |  |  | On | 3 | On | CARE Dose4D | 7.6 | 10s | 25% |  |  | 0.5 |  | 2 | 1.2 |  | 5 | 32x1.2mm | [15 empty] | Off
|  | i-Spiral 5.0 B31f | [22 empty] | 5 | 3.5 | 86 | B31f medium smooth + | Mediastinum | [13 empty] | manual
ThoraxRoutine (Adult)
ThoraxRoutine (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Thorax |  | 120 | 120 | 110 | 110 |  |  | On | 7 | On | CARE Dose4D | 7.43 | 10s | 25% |  |  | 0.5 |  | 2 | 1.2 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | Thorax 5.0 B31f | [22 empty] | 5 | 5 | 61 | B31f medium smooth + | Mediastinum | [13 empty] | manual
|  | Thorax 5.0 B80f | [22 empty] | 5 | 5 | 61 | B80f ultra sharp | Lung | [13 empty] | manual
|  | Thorax 1.0 B31f | [22 empty] | 1 | 0.7 | 429 | B31f medium smooth + | Mediastinum | [13 empty] | manual
|  | Thorax 1.0 B70f | [22 empty] | 1 | 0.7 | 429 | B70f very sharp | Lung | [13 empty] | manual
ThoraxSeqHR (Adult)
ThoraxSeqHR (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| ThorHRSeq |  | 120 | 120 | 110 | 110 |  |  | On | 3 | On | CARE Dose4D | 1.51 |  | 25% |  |  | 0.5 |  | 2 | 0 | 10 | 1 | 2x1.0mm | [15 empty] | Off
|  | ThorHRSeq 1.0 B80s | [22 empty] | 1 |  | 60 | B80s ultra sharp | Lung | [13 empty] | manual
Thorax_IR (Adult)
Thorax_IR (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Thorax |  | 120 | 120 | 66 | 66 |  |  | On | 7 | On | CARE Dose4D | 4.48 | 10s | 25% |  |  | 0.5 |  | 2 | 1.2 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | Thorax 5.0 I31f 3 | [22 empty] | 5 | 5 | 61 | I31f medium smooth | Mediastinum | [13 empty] | manual
|  | Thorax 5.0 I70f 3 | [22 empty] | 5 | 5 | 61 | I70f very sharp ASA | Lung | [13 empty] | manual
|  | Thorax 1.0 I31f 3 | [22 empty] | 1 | 0.7 | 429 | I31f medium smooth | Mediastinum | [13 empty] | manual
|  | Thorax 1.0 I70f 3 | [22 empty] | 1 | 0.7 | 429 | I70f very sharp ASA | Lung | [13 empty] | manual
Thorax_XCARE (Adult)
Thorax_XCARE (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Thorax |  | 120 | 120 | 110 | 110 |  |  | On | 7 | On | CARE Dose4D | 7.45 | 10s | 25% |  |  | 0.285 |  | 2 | 0.6 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | Thorax 5.0 B31f | [22 empty] | 5 | 5 | 61 | B31f medium smooth + | Mediastinum |  |  | X-CARE | [10 empty] | manual
|  | Thorax 5.0 B80f | [22 empty] | 5 | 5 | 61 | B80f ultra sharp | Lung |  |  | X-CARE | [10 empty] | manual
|  | Thorax 1.0 B31f | [22 empty] | 1 | 0.7 | 429 | B31f medium smooth + | Mediastinum |  |  | X-CARE | [10 empty] | manual
|  | Thorax 1.0 B70f | [22 empty] | 1 | 0.7 | 429 | B70f very sharp | Lung |  |  | X-CARE | [10 empty] | manual
Thorax_XCARE_IR (Adult)
Thorax_XCARE_IR (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Thorax |  | 120 | 120 | 66 | 66 |  |  | On | 7 | On | CARE Dose4D | 4.46 | 10s | 25% |  |  | 0.285 |  | 2 | 0.6 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | Thorax 5.0 I31f 3 | [22 empty] | 5 | 5 | 61 | I31f medium smooth | Mediastinum |  |  | X-CARE | [10 empty] | manual
|  | Thorax 5.0 I70f 3 | [22 empty] | 5 | 5 | 61 | I70f very sharp ASA | Lung |  |  | X-CARE | [10 empty] | manual
|  | Thorax 1.0 I31f 3 | [22 empty] | 1 | 0.7 | 429 | I31f medium smooth | Mediastinum |  |  | X-CARE | [10 empty] | manual
|  | Thorax 1.0 I70f 3 | [22 empty] | 1 | 0.7 | 429 | I70f very sharp ASA | Lung |  |  | X-CARE | [10 empty] | manual
Abdomen
AbdMultiPhase (Adult)
AbdMultiPhase (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Non Contrast |  | 120 | 120 | 190 | 190 |  |  | On | 3 | On | CARE Dose4D | 13.14 | 30s | 25% |  |  | 0.5 |  | 2 | 0.6 |  | 5 | 32x1.2mm | [15 empty] | Off
|  | Non Contrast 5.0 B30f | [22 empty] | 5 | 5 | 41 | B30f medium smooth | Abdomen | [13 empty] | manual
| Arterial Phase |  | 120 | 120 | 210 | 210 |  |  | On | 7 | On | CARE Dose4D | 14.19 | 30s | 25% |  |  | 0.5 |  | 2 | 0.6 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | Arterial Phase 5.0 B30f | [22 empty] | 5 | 5 | 41 | B30f medium smooth | Abdomen | [13 empty] | manual
|  | Arterial Phase 1.0 B20f | [22 empty] | 1 | 0.7 | 286 | B20f smooth | Abdomen | [13 empty] | manual
| Venous Phase |  | 120 | 120 | 190 | 190 |  |  | On | 7 | On | CARE Dose4D | 13.14 | 30s | 25% |  |  | 0.5 |  | 25 | 0.6 |  | 5 | 32x1.2mm | [15 empty] | Off
|  | Venous Phase 5.0 B30f | [22 empty] | 5 | 5 | 41 | B30f medium smooth | Abdomen | [13 empty] | manual
AbdomenIntervention (Adult)
AbdomenIntervention (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| i-Spiral |  | 120 | 120 | 210 | 210 |  |  | On | 3 | On | CARE Dose4D | 14.5 | 30s | 25% |  |  | 0.5 |  | 2 | 0.8 |  | 5 | 32x1.2mm | [15 empty] | Off
|  | i-Spiral 5.0 B30f | [22 empty] | 5 | 3.5 | 86 | B30f medium smooth | Abdomen | [13 empty] | manual
AbdomenRoutine (Adult)
AbdomenRoutine (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Abdomen |  | 120 | 120 | 210 | 210 |  |  | On | 7 | On | CARE Dose4D | 14.19 | 30s | 25% |  |  | 0.5 |  | 2 | 0.6 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | Abdomen 5.0 B30f | [22 empty] | 5 | 5 | 41 | B30f medium smooth | Abdomen | [13 empty] | manual
|  | Abdomen 1.0 B20f | [22 empty] | 1 | 0.7 | 286 | B20f smooth | Abdomen | [13 empty] | manual
AbdomenSeq (Adult)
AbdomenSeq (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| AbdSeq |  | 120 | 120 | 210 | 210 |  |  | On | 7 | On | CARE Dose4D | 16.13 |  | 25% |  |  | 0.5 |  | 2 | 0 | 34 | 5 | 32x1.2mm | [15 empty] | Off
|  | AbdSeq 5.0 B30f | [22 empty] | 5 |  | 41 | B30f medium smooth | Abdomen | [13 empty] | manual
Abdomen_IR (Adult)
Abdomen_IR (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Abdomen |  | 120 | 120 | 147 | 147 |  |  | On | 7 | On | CARE Dose4D | 9.97 | 30s | 25% |  |  | 0.5 |  | 2 | 0.6 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | Abdomen 5.0 I30f 3 | [22 empty] | 5 | 5 | 41 | I30f medium smooth | Abdomen | [13 empty] | manual
|  | Abdomen 1.0 I26f 3 | [22 empty] | 1 | 0.7 | 286 | I26f medium smooth ASA | Abdomen | [13 empty] | manual
Colonography (Adult)
Colonography (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Colo_supine |  | 120 | 120 | 55 | 55 |  |  | Semi | 1 | On | CARE Dose4D | 3.72 | 30s | 25% |  |  | 0.5 |  | 2 | 0.9 |  | 1 | 128x0.6mm | [15 empty] | Off
|  | Colo_supine 1.0 B20f | [22 empty] | 1 | 0.7 | 572 | B20f smooth | Abdomen | [13 empty] | manual
| Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Colo_prone |  | 120 | 120 | 30 | 30 |  |  | Semi | 1 | On | CARE Dose4D | 2.03 | 30s | 25% |  |  | 0.5 |  | 2 | 1.4 |  | 1 | 128x0.6mm | [15 empty] | Off
|  | Colo_prone 1.0 B10f | [22 empty] | 1 | 0.7 | 572 | B10f very smooth | Abdomen | [13 empty] | manual
DSXXL_Abdomen (Adult)
DSXXL_Abdomen (Adult) | Topogram | Topogram 0.6 T20f |  | 140 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.2 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
Spine
C_Spine (Adult)
C_Spine (Adult) | Topogram | Topogram 0.6 T80f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Body
| C_Spine |  | 120 | 120 | 275 | 275 |  |  | On | 3 | On | CARE Dose4D | 18.57 | 30s | 25% |  |  | 1 |  | 2 | 0.8 |  | 2 | 128x0.6mm | [15 empty] | Off
|  | C_Spine 2.0 B30s | [22 empty] | 2 | 2 | 81 | B30s medium smooth | Spine | [13 empty] | manual
|  | C_Spine 0.6 B20s | [22 empty] | 0.6 | 0.4 | 401 | B20s smooth | Spine | [13 empty] | manual
|  | C_Spine 0.6 B60s | [22 empty] | 0.6 | 0.4 | 401 | B60s sharp | Spine | [13 empty] | manual
C_Spine_IR (Adult)
C_Spine_IR (Adult) | Topogram | Topogram 0.6 T80f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Body
| C_Spine |  | 120 | 120 | 193 | 193 |  |  | On | 3 | On | CARE Dose4D | 13.08 | 30s | 25% |  |  | 1 |  | 2 | 0.8 |  | 2 | 128x0.6mm | [15 empty] | Off
|  | C_Spine 2.0 I30s 3 | [22 empty] | 2 | 2 | 81 | I30s medium smooth | Spine | [13 empty] | manual
|  | C_Spine 0.6 I30s 3 | [22 empty] | 0.6 | 0.4 | 401 | I30s medium smooth | Spine | [13 empty] | manual
|  | C_Spine 0.6 I50s 3 | [22 empty] | 0.6 | 0.4 | 401 | I50s medium sharp ASA | Spine | [13 empty] | manual
C_Spine_IR_Vol (Adult)
C_Spine_IR_Vol (Adult) | Topogram | Topogram 0.6 T80f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Body
| C_Spine |  | 120 | 120 | 193 | 193 |  |  | On | 3 | On | CARE Dose4D | 13.08 | 30s | 25% |  |  | 1 |  | 2 | 0.8 |  | 2 | 128x0.6mm | [15 empty] | Off
|  | C_Spine 2.0 I30s 3 | [22 empty] | 2 | 2 | 81 | I30s medium smooth | Spine | [13 empty] | manual
|  | C_Spine 0.6 MPR spine multi | [22 empty] | 0.6 | 0.4 | 0 | I26s medium smooth ASA | Spine | [13 empty] | manual
|  | C_Spine 2.0 MPR spine | [22 empty] | 2 | 2 | 17 | I30s medium smooth | Spine | [13 empty] | manual
|  | C_Spine 2.0 MPR spine | [22 empty] | 2 | 2 | 17 | I50s medium sharp ASA | Spine | [13 empty] | manual
C_Spine_Vol (Adult)
C_Spine_Vol (Adult) | Topogram | Topogram 0.6 T80f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Body
| C_Spine |  | 120 | 120 | 275 | 275 |  |  | On | 3 | On | CARE Dose4D | 18.57 | 30s | 25% |  |  | 1 |  | 2 | 0.8 |  | 2 | 128x0.6mm | [15 empty] | Off
|  | C_Spine 2.0 B30s | [22 empty] | 2 | 2 | 81 | B30s medium smooth | Spine | [13 empty] | manual
|  | C_Spine 0.6 MPR spine multi | [22 empty] | 0.6 | 0.4 | 0 | B20s smooth | Spine | [13 empty] | manual
|  | C_Spine 2.0 MPR spine | [22 empty] | 2 | 2 | 17 | B20s smooth | Spine | [13 empty] | manual
|  | C_Spine 2.0 MPR spine | [22 empty] | 2 | 2 | 17 | B60s sharp | Spine | [13 empty] | manual
DSXXL_Spine (Adult)
DSXXL_Spine (Adult) | Topogram | Topogram 0.6 T80f |  | 140 |  | 529 |  |  |  | 3 | Off | CARE Dose | 0.58 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Body
DSXXL_Spine_Vol (Adult)
DSXXL_Spine_Vol (Adult) | Topogram | Topogram 0.6 T80f |  | 140 |  | 529 |  |  |  | 3 | Off | CARE Dose | 0.58 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Body
SpineIntervention (Adult)
SpineIntervention (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| i-Spiral |  | 120 | 120 | 275 | 275 |  |  | On | 3 | On | CARE Dose4D | 18.99 | 40s | 25% |  |  | 0.5 |  | 2 | 0.8 |  | 5 | 32x1.2mm | [15 empty] | Off
|  | i-Spiral 5.0 B30f | [22 empty] | 5 | 3.5 | 46 | B30f medium smooth | Spine | [13 empty] | manual
SpineNeuro (Adult)
SpineNeuro (Adult) | Topogram | Topogram 0.6 T80f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Body
| Spine |  | 120 | 120 | 330 | 330 |  |  | On | 3 | On | CARE Dose4D | 26.85 | 40s | 25% |  |  | 1 |  | 2 | 0.8 |  | 3 | 20x0.6mm | [15 empty] | Off
|  | Spine 3.0 B30s | [22 empty] | 3 | 3 | 54 | B30s medium smooth | Spine | [13 empty] | manual
|  | Spine 0.6 B20s | [22 empty] | 0.6 | 0.4 | 401 | B20s smooth | Spine | [13 empty] | manual
|  | Spine 0.6 B60s | [22 empty] | 0.6 | 0.4 | 401 | B60s sharp | Spine | [13 empty] | manual
SpineNeuro_Vol (Adult)
SpineNeuro_Vol (Adult) | Topogram | Topogram 0.6 T80f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Body
| Spine |  | 120 | 120 | 330 | 330 |  |  | On | 3 | On | CARE Dose4D | 26.85 | 40s | 25% |  |  | 1 |  | 2 | 0.8 |  | 3 | 20x0.6mm | [15 empty] | Off
|  | Spine 3.0 B30s | [22 empty] | 3 | 3 | 54 | B30s medium smooth | Spine | [13 empty] | manual
|  | Spine 0.6 MPR spine multi | [22 empty] | 0.6 | 0.4 | 0 | B20s smooth | Spine | [13 empty] | manual
|  | Spine 2.0 MPR spine | [22 empty] | 2 | 2 | 17 | B20s smooth | Spine | [13 empty] | manual
|  | Spine 2.0 MPR spine | [22 empty] | 2 | 2 | 17 | B60s sharp | Spine | [13 empty] | manual
SpineRoutine (Adult)
SpineRoutine (Adult) | Topogram | Topogram 0.6 T80f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Body
| Spine |  | 120 | 120 | 330 | 330 |  |  | On | 3 | On | CARE Dose4D | 22.28 | 40s | 25% |  |  | 1 |  | 2 | 0.8 |  | 3 | 128x0.6mm | [15 empty] | Off
|  | Spine 3.0 B30s | [22 empty] | 3 | 3 | 54 | B30s medium smooth | Spine | [13 empty] | manual
|  | Spine 0.6 B20s | [22 empty] | 0.6 | 0.4 | 401 | B20s smooth | Spine | [13 empty] | manual
|  | Spine 0.6 B60s | [22 empty] | 0.6 | 0.4 | 401 | B60s sharp | Spine | [13 empty] | manual
SpineRoutine_Vol (Adult)
SpineRoutine_Vol (Adult) | Topogram | Topogram 0.6 T80f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Body
| Spine |  | 120 | 120 | 330 | 330 |  |  | On | 3 | On | CARE Dose4D | 22.28 | 40s | 25% |  |  | 1 |  | 2 | 0.8 |  | 3 | 128x0.6mm | [15 empty] | Off
|  | Spine 3.0 B30s | [22 empty] | 3 | 3 | 54 | B30s medium smooth | Spine | [13 empty] | manual
|  | Spine 0.6 MPR spine multi | [22 empty] | 0.6 | 0.4 | 0 | B20s smooth | Spine | [13 empty] | manual
|  | Spine 2.0 MPR spine | [22 empty] | 2 | 2 | 17 | B20s smooth | Spine | [13 empty] | manual
|  | Spine 2.0 MPR spine | [22 empty] | 2 | 2 | 17 | B60s sharp | Spine | [13 empty] | manual
SpineSeq (Adult)
SpineSeq (Adult) | Topogram | Topogram 0.6 T80f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Body
| L3-4 |  | 120 | 120 | 220 | 220 |  |  | On | 3 | On | CARE Dose4D | 17.4 |  | 25% |  |  | 1 |  | 2 | 0 | 14 | 2.4 | 12x1.2mm | [15 empty] | Off
|  | L3-4 2.4 B30s | [22 empty] | 2.4 |  | 6 | B30s medium smooth | Spine | [13 empty] | manual
| L4-5 |  | 120 | 120 | 220 | 220 |  |  | On | 3 | On | CARE Dose4D | 17.4 |  | 25% |  |  | 1 |  | 2 | 0 | 14 | 2.4 | 12x1.2mm | [15 empty] | Off
|  | L4-5 2.4 B30s | [22 empty] | 2.4 |  | 6 | B30s medium smooth | Spine | [13 empty] | manual
| L5-S1 |  | 120 | 120 | 220 | 220 |  |  | On | 3 | On | CARE Dose4D | 17.4 |  | 25% |  |  | 1 |  | 2 | 0 | 14 | 2.4 | 12x1.2mm | [15 empty] | Off
|  | L5-S1 2.4 B30s | [22 empty] | 2.4 |  | 6 | B30s medium smooth | Spine | [13 empty] | manual
Spine_IR (Adult)
Spine_IR (Adult) | Topogram | Topogram 0.6 T80f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Body
| Spine |  | 120 | 120 | 231 | 231 |  |  | On | 3 | On | CARE Dose4D | 15.61 | 40s | 25% |  |  | 1 |  | 2 | 0.8 |  | 3 | 128x0.6mm | [15 empty] | Off
|  | Spine 3.0 I30s 3 | [22 empty] | 3 | 3 | 54 | I30s medium smooth | Spine | [13 empty] | manual
|  | Spine 0.6 I30s 3 | [22 empty] | 0.6 | 0.4 | 401 | I30s medium smooth | Spine | [13 empty] | manual
|  | Spine 0.6 I50s 3 | [22 empty] | 0.6 | 0.4 | 401 | I50s medium sharp ASA | Spine | [13 empty] | manual
Spine_IR_Vol (Adult)
Spine_IR_Vol (Adult) | Topogram | Topogram 0.6 T80f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Body
| Spine |  | 120 | 120 | 231 | 231 |  |  | On | 3 | On | CARE Dose4D | 15.61 | 40s | 25% |  |  | 1 |  | 2 | 0.8 |  | 3 | 128x0.6mm | [15 empty] | Off
|  | Spine 3.0 I30s 3 | [22 empty] | 3 | 3 | 54 | I30s medium smooth | Spine | [13 empty] | manual
|  | Spine 0.6 MPR spine multi | [22 empty] | 0.6 | 0.4 | 0 | I26s medium smooth ASA | Spine | [13 empty] | manual
|  | Spine 2.0 MPR spine | [22 empty] | 2 | 2 | 17 | I30s medium smooth | Spine | [13 empty] | manual
|  | Spine 2.0 MPR spine | [22 empty] | 2 | 2 | 17 | I50s medium sharp ASA | Spine | [13 empty] | manual
Pelvis
DSXXL_Pelvis (Adult)
DSXXL_Pelvis (Adult) | Topogram | Topogram 0.6 T80f |  | 140 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.2 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Body
Hip (Adult)
Hip (Adult) | Topogram | Topogram 0.6 T80f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Body
| Hip |  | 120 | 120 | 190 | 190 |  |  | On | 5 | On | CARE Dose4D | 12.83 | 30s | 25% |  |  | 1 |  | 2 | 0.8 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | Hip 5.0 B60s | [22 empty] | 5 | 5 | 21 | B60s sharp | Pelvis | [13 empty] | manual
|  | Hip 0.6 B60s | [22 empty] | 0.6 | 0.4 | 251 | B60s sharp | Pelvis | [13 empty] | manual
PelvisRoutine (Adult)
PelvisRoutine (Adult) | Topogram | Topogram 0.6 T80f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Body
| Pelvis |  | 120 | 120 | 210 | 210 |  |  | On | 7 | On | CARE Dose4D | 14.5 | 30s | 25% |  |  | 0.5 |  | 2 | 0.8 |  | 5 | 32x1.2mm | [15 empty] | Off
|  | Pelvis 5.0 B31f | [22 empty] | 5 | 5 | 41 | B31f medium smooth + | Pelvis | [13 empty] | manual
Pelvis_IR (Adult)
Pelvis_IR (Adult) | Topogram | Topogram 0.6 T80f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Body
| Pelvis |  | 120 | 120 | 147 | 147 |  |  | On | 7 | On | CARE Dose4D | 10.19 | 30s | 25% |  |  | 0.5 |  | 2 | 0.8 |  | 5 | 32x1.2mm | [15 empty] | Off
|  | Pelvis 5.0 I31f 3 | [22 empty] | 5 | 5 | 41 | I31f medium smooth | Pelvis | [13 empty] | manual
Specials
BodyMultiPhase_VAngio (Adult)
BodyMultiPhase_VAngio (Adult) | DynMulti4D_L |  |  | 100 |  | 150 |  |  |  | 3 | Off | CARE Dose | 76.8 |  | 25% |  |  | 0.285 |  | 2 | 1 |  | 5 | 32x1.2mm
|  | DynMulti4D_L 5.0 B30f | [22 empty] | 5 | 3 | 710 | B30f medium smooth | Abdomen
|  | DynMulti4D_L 3.0 B20f | [22 empty] | 3 | 2 | 1070 | B20f smooth | Abdomen
|  | DynMulti4D_L 1.5 B20f | [22 empty] | 1.5 | 1 | 2150 | B20f smooth | Abdomen
BodyPCT (Adult)
BodyPCT (Adult) | DynMulti |  |  | 100 |  | 150 |  |  |  | 3 | Off | CARE Dose | 252.56 |  | 25% |  |  | 1 |  | 2 | 0 |  | 10 | 32x1.2mm | [15 empty] | Off
|  | DynMulti 10.0 B30s | [22 empty] | 10 | 7 | 195 | B30s medium smooth | Abdomen | [13 empty] | manual
BodyPCTSeq (Adult)
BodyPCTSeq (Adult) | DynSerio |  |  | 100 | 150 | 150 |  |  |  | 3 | Off | CARE Dose | 265.19 |  | 25% |  |  | 0.5 |  | 2 | 0 | 0 | 10 | 32x1.2mm | [15 empty] | Off
|  | DynSerio 10.0 B30f | [22 empty] | 10 |  | 210 | B30f medium smooth | Abdomen | [13 empty] | manual
BodyPCTSingleSeq (Adult)
BodyPCTSingleSeq (Adult) | DynSerio |  |  | 100 | 150 | 150 |  |  |  | 3 | Off | CARE Dose | 236.85 |  | 25% |  |  | 0.5 |  | 2 | 0 | 0 | 10 | 1x10.0mm | [15 empty] | Off
|  | DynSerio 10.0 B30s | [22 empty] | 10 |  | 42 | B30s medium smooth | Abdomen | [13 empty] | manual
BodyVPCT (Adult)
BodyVPCT (Adult) | DynMulti4D |  |  | 100 |  | 150 |  |  |  | 3 | Off | CARE Dose | 146.58 |  | 25% |  |  | 0.285 |  | 2 | 0.5 |  | 5 | 32x1.2mm
|  | VPCT DynMulti4D 5.0 B20f | [22 empty] | 5 | 3 | 864 | B20f smooth | Abdomen
|  | VPCT DynMulti4D 3.0 B20f | [22 empty] | 3 | 2 | 1323 | B20f smooth | Abdomen
|  | DynMulti4D 1.5 B20f | [22 empty] | 1.5 | 1 | 2673 | B20f smooth | Abdomen
BodyVPCT_long (Adult)
BodyVPCT_long (Adult) | DynMulti4D_L |  |  | 100 |  | 150 |  |  |  | 3 | Off | CARE Dose | 222.29 |  | 25% |  |  | 0.285 |  | 2 | 1.02 |  | 5 | 32x1.2mm
|  | VPCT DynMulti4D_L 5.0 B20f | [22 empty] | 5 | 3 | 1323 | B20f smooth | Abdomen
|  | VPCT DynMulti4D_L 3.0 B20f | [22 empty] | 3 | 2 | 1998 | B20f smooth | Abdomen
|  | DynMulti4D_L 1.5 B20f | [22 empty] | 1.5 | 1 | 4023 | B20f smooth | Abdomen
DynMultiECG (Adult)
DynMultiECG (Adult) | DynMulti |  | 80 | 80 |  | 100 |  |  | Off | 3 | Off | CARE Dose | 169.9 |  | 25% |  |  | 0.285 |  | 2 | 0 |  | 3 | 32x1.2mm | [15 empty] | Normal | 70% | 70%
|  | DynMulti 3.0 B30f 70% | [22 empty] | 3 | 1.5 | 720 | B30f medium smooth | Mediastinum | [13 empty] | manual | 70%
DynSerioECGSeq (Adult)
DynSerioECGSeq (Adult) | DynSerio |  | 80 | 80 | 100 | 100 |  |  | Off | 3 | Off | CARE Dose | 35 |  | 25% |  |  | 0.285 |  | 2 | 0 | 0 | 3 | 32x1.2mm | [13 empty] | 70% | 70% | Off
|  | DynSerio 3.0 B30f 70% | [22 empty] | 3 |  | 720 | B30f medium smooth | Mediastinum | [13 empty] | manual | 70%
ExtrMultiPhase_VAngioPlus (Adult)
ExtrMultiPhase_VAngioPlus (Adult) | DynMulti4D_XL |  |  | 100 |  | 100 |  |  |  | 3 | Off | CARE Dose | 41.49 |  | 25% |  |  | 0.285 |  | 2 | 1.48 |  | 5 | 32x1.2mm
|  | DynMulti4D_XL 5.0 B30f | [22 empty] | 5 | 3 | 1008 | B30f medium smooth | CT Angio
|  | DynMulti4D_XL 3.0 B20f | [22 empty] | 3 | 2 | 1520 | B20f smooth | CT Angio
|  | DynMulti4D_XL 1.5 B20f | [22 empty] | 1.5 | 1 | 3048 | B20f smooth | CT Angio
Flash_HeadTrauma (Adult)
Flash_HeadTrauma (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.29 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head
| Fl_Head |  | 120 | 120 | 210 |  | 210 |  | On | 3 | On | CARE Dose4D | 32.14 | 5s | 25% |  |  | 0.285 |  | 2 | 1.55 |  | 5 | 128x0.6mm
|  | Fl_Head 5.0 H31f | [22 empty] | 5 | 5 | 25 | H31f medium smooth + | Cerebrum
|  | Fl_Head 5.0 H60f | [22 empty] | 5 | 5 | 25 | H60f sharp FR | Bone
Flash_Trauma (Adult)
Flash_Trauma (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 365 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Fl_Trauma |  | 140 | 140 | 140 |  | 140 |  | On | 3 | On | CARE Dose4D | 15.43 | 30s | 25% |  |  | 0.285 |  | 2 | 2.1 |  | 7 | 128x0.6mm
|  | Fl_Trauma 7.0 B30f | [22 empty] | 7 | 7 | 108 | B30f medium smooth | Abdomen
|  | Fl_Trauma 1.5 B30f | [22 empty] | 1.5 | 1 | 751 | B30f medium smooth | Abdomen
|  | Fl_Trauma 1.5 B70f | [22 empty] | 1.5 | 1 | 751 | B70f very sharp | Bone
HeadTrauma (Adult)
HeadTrauma (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.29 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head
| Head |  | 120 | 120 | 390 | 390 |  |  | On | 3 | On | CARE Dose4D | 55.12 | 9s | 25% |  |  | 1 |  | 2 | 0.55 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | Head 5.0 H31s | [22 empty] | 5 | 5 | 25 | H31s medium smooth + | Cerebrum | [13 empty] | manual
|  | Head 5.0 H60s | [22 empty] | 5 | 5 | 25 | H60s sharp FR | Bone | [13 empty] | manual
HeadTraumaSeq (Adult)
HeadTraumaSeq (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.29 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head
| HeadSeq |  | 120 | 120 | 390 | 390 |  |  | On | 3 | On | CARE Dose4D | 62.46 |  | 25% |  |  | 1 |  | 2 | 0 | 34 | 5 | 32x1.2mm | [15 empty] | Off
|  | HeadSeq 5.0 H31s | [22 empty] | 5 |  | 27 | H31s medium smooth + | Base Orbita | [13 empty] | manual
|  | HeadSeq 5.0 H60s | [22 empty] | 5 |  | 27 | H60s sharp FR | Bone | [13 empty] | manual
NeuroMultiPhase_VAngio (Adult)
NeuroMultiPhase_VAngio (Adult) | DynMulti4D_L |  | 80 | 80 |  | 150 |  |  | Off | 3 | Off | CARE Dose | 297.06 |  | 25% |  |  | 0.285 |  | 2 | 1.02 |  | 5 | 32x1.2mm
|  | DynMulti4D_L 5.0 H30f | [22 empty] | 5 | 3 | 1470 | H30f medium smooth | Cerebrum
|  | DynMulti4D_L 3.0 H30f | [22 empty] | 3 | 2 | 2220 | H30f medium smooth | Cerebrum
|  | DynMulti4D_L 1.5 H30f | [22 empty] | 1.5 | 1 | 4470 | H30f medium smooth | Cerebrum
NeuroPCT (Adult)
NeuroPCT (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.29 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head
| Head |  | 120 | 120 | 390 | 390 |  |  | On | 3 | On | CARE Dose4D | 71.99 | 76s | 25% |  |  | 1 |  | 2 | 0.55 |  | 5 | 20x0.6mm | [15 empty] | Off
|  | Head 5.0 H31s | [22 empty] | 5 | 5 | 25 | H31s medium smooth + | Base Orbita | [13 empty] | manual
| DynMulti |  | 80 | 80 |  | 200 |  |  | Off | 3 | Off | CARE Dose | 405 |  | 25% |  |  | 1 |  | 2 | 0 |  | 10 | 32x1.2mm | [15 empty] | Off
|  | PCT DynMulti 10.0 H30s | [22 empty] | 10 | 7 | 195 | H30s medium smooth | Cerebrum | [13 empty] | manual
NeuroPCTSingle (Adult)
NeuroPCTSingle (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.29 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head
| Head |  | 120 | 120 | 390 | 390 |  |  | On | 3 | On | CARE Dose4D | 71.99 | 76s | 25% |  |  | 1 |  | 2 | 0.55 |  | 5 | 20x0.6mm | [15 empty] | Off
|  | Head 5.0 H31s | [22 empty] | 5 | 5 | 25 | H31s medium smooth + | Base Orbita | [13 empty] | manual
| DynMulti |  | 80 | 80 |  | 200 |  |  | Off | 3 | Off | CARE Dose | 361.73 |  | 25% |  |  | 1 |  | 2 | 0 |  | 10 | 1x10.0mm | [15 empty] | Off
|  | PCT DynMulti 10.0 H30s | [22 empty] | 10 | 10 | 39 | H30s medium smooth | Cerebrum | [13 empty] | manual
NeuroVPCT (Adult)
NeuroVPCT (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.29 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head
| Head |  | 120 | 120 | 320 | 320 |  |  | On | 3 | On | CARE Dose4D | 58.93 | 76s | 25% |  |  | 1 |  | 2 | 0.55 |  | 5 | 20x0.6mm | [15 empty] | Off
|  | Head 5.0 H31s | [22 empty] | 5 | 5 | 25 | H31s medium smooth + | Base Orbita | [13 empty] | manual
|  | Head 1.0 H23s | [22 empty] | 1 | 0.7 | 172 | H23s smooth PBV FR | CT Angio | [13 empty] | manual
| DynMulti4D |  | 80 | 80 |  | 200 |  |  | Off | 3 | Off | CARE Dose | 259.69 |  | 25% |  |  | 0.285 |  | 2 | 0.5 |  | 10 | 32x1.2mm
|  | VPCT DynMulti4D 10.0 H20f | [22 empty] | 10 | 4 | 690 | H20f smooth | Cerebrum
|  | VPCT DynMulti4D 5.0 H20f | [22 empty] | 5 | 3 | 960 | H20f smooth | Cerebrum
|  | VPCT DynMulti4D 1.5 H20f | [22 empty] | 1.5 | 1 | 2970 | H20f smooth | Cerebrum
| HeadAngio |  | 100 | 100 | 175 | 175 |  |  | On | 11 | On | CARE Dose4D | 16.89 | 5s | 25% |  |  | 0.5 |  | 2 | 0.8 |  | 1 | 128x0.6mm | [15 empty] | Off
|  | HeadAngio 1.0 H22f | [22 empty] | 1 | 0.7 | 172 | H22f smooth FR | CT Angio | [13 empty] | manual
PolyTrauma (Adult)
PolyTrauma (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.29 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head
| Head |  | 120 | 120 | 390 | 390 |  |  | On | 3 | On | CARE Dose4D | 60.77 | 7s | 25% |  |  | 0.5 |  | 2 | 0.55 |  | 6 | 32x1.2mm | [15 empty] | Off
|  | Head 6.0 H31f | [22 empty] | 6 | 6 | 21 | H31f medium smooth + | Base Orbita | [13 empty] | manual
| Neck |  | 120 | 120 | 165 | 165 |  |  | On | 3 | On | CARE Dose4D | 11.39 | 10s | 25% |  |  | 0.5 |  | 5 | 0.8 |  | 5 | 32x1.2mm | [15 empty] | Off
|  | Neck 5.0 B31f | [22 empty] | 5 | 5 | 29 | B31f medium smooth + | Larynx | [13 empty] | manual
| Topogram | Topogram 0.6 T20f |  | 120 |  | 275 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| ThorAbd |  | 120 | 120 | 210 | 210 |  |  | On | 3 | On | CARE Dose4D | 14.51 | 30s | 25% |  |  | 0.5 |  | 2 | 1.4 |  | 7 | 32x1.2mm | [15 empty] | Off
|  | Tho ThorAbd 7.0 B31f | [22 empty] | 7 | 7 | 29 | B31f medium smooth + | Mediastinum | [13 empty] | manual
|  | Lun ThorAbd 7.0 B80f | [22 empty] | 7 | 7 | 29 | B80f ultra sharp | Lung | [13 empty] | manual
|  | Abd ThorAbd 7.0 B30f | [22 empty] | 7 | 7 | 58 | B30f medium smooth | Abdomen | [13 empty] | manual
|  | MPR ThorAbd 1.5 B30f | [22 empty] | 1.5 | 1 | 601 | B30f medium smooth | Mediastinum | [13 empty] | manual
Trauma (Adult)
Trauma (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 365 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Trauma |  | 120 | 120 | 180 | 180 |  |  | On | 3 | On | CARE Dose4D | 12.42 | 30s | 25% |  |  | 0.5 |  | 2 | 1.5 |  | 7 | 32x1.2mm | [15 empty] | Off
|  | Trauma 7.0 B30f | [22 empty] | 7 | 7 | 108 | B30f medium smooth | Abdomen | [13 empty] | manual
|  | Trauma 1.5 B30f | [22 empty] | 1.5 | 1 | 751 | B30f medium smooth | Abdomen | [13 empty] | manual
|  | Trauma 1.5 B70f | [22 empty] | 1.5 | 1 | 751 | B70f very sharp | Bone | [13 empty] | manual
TraumaThin (Adult)
TraumaThin (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 365 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| TraumaThin |  | 120 | 120 | 180 | 180 |  |  | On | 3 | On | CARE Dose4D | 12.14 | 30s | 25% |  |  | 0.5 |  | 2 | 1.5 |  | 7 | 128x0.6mm | [15 empty] | Off
|  | TraumaThin 7.0 B30f | [22 empty] | 7 | 7 | 108 | B30f medium smooth | Abdomen | [13 empty] | manual
|  | TraumaThin 0.6 B30f | [22 empty] | 0.6 | 0.3 | 2501 | B30f medium smooth | Abdomen | [13 empty] | manual
|  | TraumaThin 0.6 B70f | [22 empty] | 0.6 | 0.3 | 2501 | B70f very sharp | Bone | [13 empty] | manual
UpperExtremities
Extremity (Adult)
Extremity (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Extremity |  |  | 120 | 100 | 100 |  |  |  | 3 | On | CARE Dose | 6.75 | 8s | 25% |  |  | 1 |  | 2 | 0.8 |  | 4 | 128x0.6mm | [15 empty] | Off
|  | Extremity 4.0 B60s | [22 empty] | 4 | 4 | 38 | B60s sharp | Extremity | [13 empty] | manual
|  | Extremity 1.0 B60s | [22 empty] | 1 | 0.7 | 215 | B60s sharp | Bone | [13 empty] | manual
ExtrRoutineZUHR (Adult)
ExtrRoutineZUHR (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| ExtrUHR |  |  | 120 | 180 | 180 |  |  |  | 3 | On | CARE Dose | 17.65 | 20s | 25% |  |  | 1 |  | 2 | 0.85 |  | 2 | 16x0.3mm | [15 empty] | Off
|  | ExtrUHR 2.0 U90u | [22 empty] | 2 | 2 | 31 | U90u ultra sharp | Extremity | [13 empty] | manual
|  | ExtrUHR 0.4 U90u | [22 empty] | 0.4 | 0.2 | 301 | U90u ultra sharp | Bone | [13 empty] | manual
ExtrUHR (Adult)
ExtrUHR (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| ExtrUHR |  |  | 120 | 120 | 120 |  |  |  | 3 | On | CARE Dose | 11.77 | 20s | 25% |  |  | 1 |  | 2 | 0.85 |  | 2 | 16x0.6mm | [15 empty] | Off
|  | ExtrUHR 2.0 U90u | [22 empty] | 2 | 2 | 31 | U90u ultra sharp | Extremity | [13 empty] | manual
|  | ExtrUHR 0.6 U90u | [22 empty] | 0.6 | 0.4 | 151 | U90u ultra sharp | Bone | [13 empty] | manual
ExtrZUHR_IR (Adult)
ExtrZUHR_IR (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| ExtrUHR |  |  | 120 | 126 | 126 |  |  |  | 3 | On | CARE Dose | 12.46 | 20s | 25% |  |  | 1 |  | 2 | 0.85 |  | 2 | 16x0.3mm | [15 empty] | Off
|  | ExtrUHR 2.0 V90u 3 | [22 empty] | 2 | 2 | 31 | V90u ultra sharp | Extremity | [13 empty] | manual
|  | ExtrUHR 0.4 V90u 3 | [22 empty] | 0.4 | 0.2 | 301 | V90u ultra sharp | Bone | [13 empty] | manual
WristZUHR (Adult)
WristZUHR (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| WristUHR |  |  | 120 | 180 | 180 |  |  |  | 3 | On | CARE Dose | 17.65 | 20s | 25% |  |  | 1 |  | 2 | 0.85 |  | 2 | 16x0.3mm | [15 empty] | Off
|  | WristUHR 2.0 U90u | [22 empty] | 2 | 2 | 31 | U90u ultra sharp | Extremity | [13 empty] | manual
|  | WristUHR 0.4 U90u | [22 empty] | 0.4 | 0.2 | 301 | U90u ultra sharp | Bone | [13 empty] | manual
LowerExtremities
Extremity (Adult)
Extremity (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Extremity |  |  | 120 | 100 | 100 |  |  |  | 3 | On | CARE Dose | 6.75 | 8s | 25% |  |  | 1 |  | 2 | 0.8 |  | 4 | 128x0.6mm | [15 empty] | Off
|  | Extremity 4.0 B60s | [22 empty] | 4 | 4 | 38 | B60s sharp | Extremity | [13 empty] | manual
|  | Extremity 1.0 B60s | [22 empty] | 1 | 0.7 | 215 | B60s sharp | Bone | [13 empty] | manual
ExtrRoutineZUHR (Adult)
ExtrRoutineZUHR (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| ExtrUHR |  |  | 120 | 180 | 180 |  |  |  | 3 | On | CARE Dose | 17.65 | 20s | 25% |  |  | 1 |  | 2 | 0.85 |  | 2 | 16x0.3mm | [15 empty] | Off
|  | ExtrUHR 2.0 U90u | [22 empty] | 2 | 2 | 31 | U90u ultra sharp | Extremity | [13 empty] | manual
|  | ExtrUHR 0.4 U90u | [22 empty] | 0.4 | 0.2 | 301 | U90u ultra sharp | Bone | [13 empty] | manual
ExtrUHR (Adult)
ExtrUHR (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| ExtrUHR |  |  | 120 | 120 | 120 |  |  |  | 3 | On | CARE Dose | 11.77 | 20s | 25% |  |  | 1 |  | 2 | 0.85 |  | 2 | 16x0.6mm | [15 empty] | Off
|  | ExtrUHR 2.0 U90u | [22 empty] | 2 | 2 | 31 | U90u ultra sharp | Extremity | [13 empty] | manual
|  | ExtrUHR 0.6 U90u | [22 empty] | 0.6 | 0.4 | 151 | U90u ultra sharp | Bone | [13 empty] | manual
ExtrZUHR_IR (Adult)
ExtrZUHR_IR (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| ExtrUHR |  |  | 120 | 126 | 126 |  |  |  | 3 | On | CARE Dose | 12.46 | 20s | 25% |  |  | 1 |  | 2 | 0.85 |  | 2 | 16x0.3mm | [15 empty] | Off
|  | ExtrUHR 2.0 V90u 3 | [22 empty] | 2 | 2 | 31 | V90u ultra sharp | Extremity | [13 empty] | manual
|  | ExtrUHR 0.4 V90u 3 | [22 empty] | 0.4 | 0.2 | 301 | V90u ultra sharp | Bone | [13 empty] | manual
FootUHR (Adult)
FootUHR (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| FootUHR |  |  | 120 | 120 | 120 |  |  |  | 3 | On | CARE Dose | 11.77 | 47s | 25% |  |  | 1 |  | 2 | 0.85 |  | 3 | 16x0.6mm | [15 empty] | Off
|  | FootUHR 3.0 U90u | [22 empty] | 3 | 3 | 51 | U90u ultra sharp | Extremity | [13 empty] | manual
|  | FootUHR 0.6 U90u | [22 empty] | 0.6 | 0.4 | 376 | U90u ultra sharp | Bone | [13 empty] | manual
KneeUHR (Adult)
KneeUHR (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| KneeUHR |  |  | 120 | 140 | 140 |  |  |  | 3 | On | CARE Dose | 13.73 | 47s | 25% |  |  | 1 |  | 2 | 0.85 |  | 3 | 16x0.6mm | [15 empty] | Off
|  | KneeUHR 3.0 U90u | [22 empty] | 3 | 3 | 51 | U90u ultra sharp | Extremity | [13 empty] | manual
|  | KneeUHR 0.6 U90u | [22 empty] | 0.6 | 0.4 | 376 | U90u ultra sharp | Bone | [13 empty] | manual
Cardiac
DSXXL_CoronaryCTA (Adult)
DSXXL_CoronaryCTA (Adult) | Topogram | Topogram 0.6 T20f |  | 140 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.2 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
DS_CaScore (Adult)
DS_CaScore (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| DS_CaSc |  | 120 | 120 | 80 |  | 80 |  | Semi | 3 | On | CARE Dose4D | 26.25 |  | 25% |  |  | 0.285 |  | 2 | 0.17 |  | 3 | 128x0.6mm | [15 empty] | Advance | 70% | 70%
|  | DS_CaSc 3.0 B35f BestDiast | [22 empty] | 3 | 1.5 | 101 | B35f HeartView medium CaSco | Mediastinum | [13 empty] | BestDiast | 70%
DS_CaScoreSeq (Adult)
DS_CaScoreSeq (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| DS_CaScSeq |  | 120 | 120 | 80 |  | 80 |  | Semi | 3 | On | CARE Dose4D | 3 |  | 25% |  |  | 0.285 |  | 2 | 0 | 34 | 3 | 32x1.2mm | [13 empty] | 70% | 70% | Off
|  | DS_CaScSeq 3.0 B35f 70% | [22 empty] | 3 |  | 91 | B35f HeartView medium CaSco | Mediastinum | [13 empty] | manual | 70%
DS_CoroCTA_AdaptSeqArrhythmicHR (Adult)
DS_CoroCTA_AdaptSeqArrhythmicHR (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| DS_CorAdSeq |  | 100 | 100 | 366 |  | 366 |  | On | 11 | On | CARE Dose4D | 47.37 |  | 25% |  |  | 0.285 |  | 2 | 0 | 34 | 3 | 128x0.6mm | [13 empty] | 250 ms | 450 ms | Normal | 300 ms | 400 ms
|  | DS_CorAdSeq 3.0 B30f BestSyst | [22 empty] | 3 |  | 46 | B30f medium smooth | Cardiac | [13 empty] | BestSyst | 350 ms
|  | DS_CorAdSeq 0.6 B26f BestSyst | [22 empty] | 0.6 |  | 344 | B26f medium smooth ASA | Cardiac | [13 empty] | BestSyst | 350 ms
DS_CoroCTA_AdaptSeq_Edge (Adult)
DS_CoroCTA_AdaptSeq_Edge (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| DS_CorAdSeq |  | 100 | 100 | 256 |  | 256 |  | On | 11 | On | CARE Dose4D | 33.15 |  | 25% |  |  | 0.285 |  | 2 | 0 | 34 | 3 | 128x0.6mm | [13 empty] | 70% | 70% | Normal | 70% | 70%
|  | DS_CorAdSeq 3.0 I30f 3 BestDiast | [22 empty] | 3 |  | 46 | I30f medium smooth | Cardiac | [13 empty] | BestDiast | 70%
|  | DS_CorAdSeq 0.75 I26f 3 BestDiast | [22 empty] | 0.75 |  | 344 | I26f medium smooth ASA | Cardiac | [13 empty] | BestDiast | 70%
|  | Edge DS_CorAdSeq 0.5 I46f 3 BestDiast | [22 empty] | 0.5 |  | 917 | I46f HeartView sharp ASA | Cardiac | [13 empty] | BestDiast | 70%
DS_CoroCTA_ArrhythmicHR (Adult)
DS_CoroCTA_ArrhythmicHR (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| DS_CorCTA |  | 100 | 100 | 366 |  | 366 |  | On | 11 | On | CARE Dose4D | 71.65 |  | 25% |  |  | 0.285 |  | 2 | 0.17 |  | 3 | 128x0.6mm | [15 empty] | Normal | 300 ms | 400 ms
|  | DS_CorCTA 3.0 B30f BestSyst | [22 empty] | 3 | 3 | 41 | B30f medium smooth | Cardiac | [13 empty] | BestSyst | 350 ms
|  | Coro DS_CorCTA 0.6 B26f BestSyst | [22 empty] | 0.6 | 0.3 | 401 | B26f medium smooth ASA | Cardiac | [13 empty] | BestSyst | 350 ms
DS_CoronaryCTA (Adult)
DS_CoronaryCTA (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| DS_CorCTA |  | 100 | 100 | 366 |  | 366 |  | On | 11 | On | CARE Dose4D | 71.65 |  | 25% |  |  | 0.285 |  | 2 | 0.17 |  | 3 | 128x0.6mm | [15 empty] | Normal | 70% | 70%
|  | DS_CorCTA 3.0 B30f BestDiast | [22 empty] | 3 | 3 | 41 | B30f medium smooth | Cardiac | [13 empty] | BestDiast | 70%
|  | Coro DS_CorCTA 0.6 B26f BestDiast | [22 empty] | 0.6 | 0.3 | 401 | B26f medium smooth ASA | Cardiac | [13 empty] | BestDiast | 70%
|  | Func DS_CorCTA 1.5 B26f 10 - 100 % Matrix 256 | [22 empty] | 1.5 | 1 | 1210 | B26f medium smooth ASA | Cardiac |  | Functional imaging 256 matrix | [11 empty] | manual |  | 10 % 20 % 30 % 40 % 50 % 60 % 70 % 80 % 90 % 100 %
DS_CoronaryCTABiSeg (Adult)
DS_CoronaryCTABiSeg (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| DS_CorCTABi |  | 100 | 100 | 366 |  | 366 |  | On | 11 | On | CARE Dose4D | 71.65 |  | 25% |  |  | 0.285 |  | 2 | 0.17 |  | 3 | 128x0.6mm | [15 empty] | Normal | 70% | 70%
|  | DS_CorCTABi 3.0 B30f BestDiast | [22 empty] | 3 | 3 | 41 | B30f medium smooth | Cardiac | [13 empty] | BestDiast | 70%
|  | Coro DS_CorCTABi 0.6 B26f BestDiast | [22 empty] | 0.6 | 0.3 | 401 | B26f medium smooth ASA | Cardiac | [13 empty] | BestDiast | 70%
|  | Func DS_CorCTABi 1.5 B26f 10 - 100 % Matrix 256 | [22 empty] | 1.5 | 1 | 1210 | B26f medium smooth ASA | Cardiac |  | Functional imaging 256 matrix | [11 empty] | manual |  | 10 % 20 % 30 % 40 % 50 % 60 % 70 % 80 % 90 % 100 %
DS_CoronaryCTA_AdaptSeq (Adult)
DS_CoronaryCTA_AdaptSeq (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| DS_CorAdSeq |  | 100 | 100 | 366 |  | 366 |  | On | 11 | On | CARE Dose4D | 47.37 |  | 25% |  |  | 0.285 |  | 2 | 0 | 34 | 3 | 128x0.6mm | [13 empty] | 70% | 70% | Normal | 70% | 70%
|  | DS_CorAdSeq 3.0 B30f BestDiast | [22 empty] | 3 |  | 46 | B30f medium smooth | Cardiac | [13 empty] | BestDiast | 70%
|  | DS_CorAdSeq 0.6 B26f BestDiast | [22 empty] | 0.6 |  | 344 | B26f medium smooth ASA | Cardiac | [13 empty] | BestDiast | 70%
DS_CoronaryCTA_IR (Adult)
DS_CoronaryCTA_IR (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| DS_CorCTA |  | 100 | 100 | 256 |  | 256 |  | On | 11 | On | CARE Dose4D | 50.14 |  | 25% |  |  | 0.285 |  | 2 | 0.17 |  | 3 | 128x0.6mm | [15 empty] | Normal | 70% | 70%
|  | DS_CorCTA 3.0 I30f 3 BestDiast | [22 empty] | 3 | 3 | 41 | I30f medium smooth | Cardiac | [13 empty] | BestDiast | 70%
|  | Coro DS_CorCTA 0.6 I26f 3 BestDiast | [22 empty] | 0.6 | 0.3 | 401 | I26f medium smooth ASA | Cardiac | [13 empty] | BestDiast | 70%
|  | Func DS_CorCTA 1.5 I26f 3 10 - 100 % Matrix 256 | [22 empty] | 1.5 | 1 | 1210 | I26f medium smooth ASA | Cardiac |  | Functional imaging 256 matrix | [11 empty] | manual |  | 10 % 20 % 30 % 40 % 50 % 60 % 70 % 80 % 90 % 100 %
FAST_Flash_Cardio (Adult)
FAST_Flash_Cardio (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Fl_CaSc |  | 120 | 120 | 80 |  | 80 |  | Semi | 3 | On | CARE Dose4D | 1.31 |  | 25% |  |  | 0.285 |  | 2 | 3.4 |  | 3 | 128x0.6mm
|  | Fl_CaSc 3.0 B35f 70% | [22 empty] | 3 | 1.5 | 91 | B35f HeartView medium CaSco | Mediastinum | [14 empty] | 70%
| Fl_CorCTA |  | 100 | 100 | 366 |  | 366 |  | On | 11 | On | CARE Dose4D | 3.57 |  | 25% |  |  | 0.285 |  | 2 | 3.4 |  | 3 | 128x0.6mm
|  | Fl_CorCTA 3.0 B30f 60% | [22 empty] | 3 | 3 | 46 | B30f medium smooth | Cardiac | [14 empty] | 60%
|  | Fl_CorCTA 0.6 B26f 60% | [22 empty] | 0.6 | 0.3 | 451 | B26f medium smooth ASA | Cardiac | [14 empty] | 60%
| DS_CorAdSeq |  | 100 | 100 | 366 |  | 366 |  | On | 11 | On | CARE Dose4D | 47.37 |  | 25% |  |  | 0.285 |  | 2 | 0 | 34 | 3 | 128x0.6mm | [13 empty] | 70% | 70% | Normal | 70% | 70%
|  | DS_CorAdSeq 3.0 B30f BestDiast | [22 empty] | 3 |  | 46 | B30f medium smooth | Cardiac | [13 empty] | BestDiast | 70%
|  | DS_CorAdSeq 0.6 B26f BestDiast | [22 empty] | 0.6 |  | 344 | B26f medium smooth ASA | Cardiac | [13 empty] | BestDiast | 70%
Flash_Cardio_IR (Adult)
Flash_Cardio_IR (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Fl_CaSc |  | 120 | 120 | 80 |  | 80 |  | Semi | 3 | On | CARE Dose4D | 1.31 |  | 25% |  |  | 0.285 |  | 2 | 3.4 |  | 3 | 128x0.6mm
|  | Fl_CaSc 3.0 B35f 70% | [22 empty] | 3 | 1.5 | 91 | B35f HeartView medium CaSco | Mediastinum | [14 empty] | 70%
| Fl_CorCTA |  | 100 | 100 | 256 |  | 256 |  | On | 11 | On | CARE Dose4D | 2.5 |  | 25% |  |  | 0.285 |  | 2 | 3.4 |  | 3 | 128x0.6mm
|  | Fl_CorCTA 3.0 I30f 3 60% | [22 empty] | 3 | 3 | 46 | I30f medium smooth | Cardiac | [14 empty] | 60%
|  | Fl_CorCTA 0.6 I26f 3 60% | [22 empty] | 0.6 | 0.3 | 451 | I26f medium smooth ASA | Cardiac | [14 empty] | 60%
| DS_CorAdSeq |  | 100 | 100 | 256 |  | 256 |  | On | 11 | On | CARE Dose4D | 33.15 |  | 25% |  |  | 0.285 |  | 2 | 0 | 34 | 3 | 128x0.6mm | [13 empty] | 70% | 70% | Normal | 70% | 70%
|  | DS_CorAdSeq 3.0 I30f 3 BestDiast | [22 empty] | 3 |  | 46 | I30f medium smooth | Cardiac | [13 empty] | BestDiast | 70%
|  | DS_CorAdSeq 0.6 I26f 3 BestDiast | [22 empty] | 0.6 |  | 344 | I26f medium smooth ASA | Cardiac | [13 empty] | BestDiast | 70%
Flash_CaScore (Adult)
Flash_CaScore (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Fl_CaSc |  | 120 | 120 | 80 |  | 80 |  | Semi | 3 | On | CARE Dose4D | 1.31 |  | 25% |  |  | 0.285 |  | 2 | 3.4 |  | 3 | 128x0.6mm
|  | Fl_CaSc 3.0 B35f 70% | [22 empty] | 3 | 1.5 | 91 | B35f HeartView medium CaSco | Mediastinum | [14 empty] | 70%
Vascular
AngioRunOff (Adult)
AngioRunOff (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 365 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| AngioRunOff |  | 120 | 120 | 120 | 120 |  |  | On | 11 | On | CARE Dose4D | 8.11 | 27s | 25% |  |  | 0.5 |  | 2 | 0.6 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | AngioRunOff 5.0 B30f | [22 empty] | 5 | 5 | 201 | B30f medium smooth | Abdomen | [13 empty] | manual
|  | AngioRunOff 1.0 B26f | [22 empty] | 1 | 0.7 | 1429 | B26f medium smooth ASA | CT Angio | [13 empty] | manual
AngioRunOff_IR (Adult)
AngioRunOff_IR (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 365 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| AngioRunOff |  | 120 | 120 | 84 | 84 |  |  | On | 11 | On | CARE Dose4D | 5.69 | 27s | 25% |  |  | 0.5 |  | 2 | 0.6 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | AngioRunOff 5.0 I30f 3 | [22 empty] | 5 | 5 | 201 | I30f medium smooth | Abdomen | [13 empty] | manual
|  | AngioRunOff 1.0 I26f 3 | [22 empty] | 1 | 0.7 | 1429 | I26f medium smooth ASA | CT Angio | [13 empty] | manual
BodyAngioRoutine (Adult)
BodyAngioRoutine (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| BodyAngio |  | 120 | 120 | 120 | 120 |  |  | On | 11 | On | CARE Dose4D | 8.11 | 6s | 25% |  |  | 0.5 |  | 2 | 1.2 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | BodyAngio 5.0 B30f | [22 empty] | 5 | 5 | 81 | B30f medium smooth | Abdomen | [13 empty] | manual
|  | BodyAngio 0.75 B26f | [22 empty] | 0.75 | 0.5 | 801 | B26f medium smooth ASA | CT Angio | [13 empty] | manual
BodyAngio_IR (Adult)
BodyAngio_IR (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| BodyAngio |  | 120 | 120 | 84 | 84 |  |  | On | 11 | On | CARE Dose4D | 5.69 | 6s | 25% |  |  | 0.5 |  | 2 | 1.2 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | BodyAngio 5.0 I30f 3 | [22 empty] | 5 | 5 | 81 | I30f medium smooth | Abdomen | [13 empty] | manual
|  | BodyAngio 0.75 I26f 3 | [22 empty] | 0.75 | 0.5 | 801 | I26f medium smooth ASA | CT Angio | [13 empty] | manual
BodyAngio_Vol (Adult)
BodyAngio_Vol (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| BodyAngio |  | 120 | 120 | 120 | 120 |  |  | On | 11 | On | CARE Dose4D | 8.11 | 6s | 25% |  |  | 0.5 |  | 2 | 1.2 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | BodyAngio 5.0 B30f | [22 empty] | 5 | 5 | 81 | B30f medium smooth | Abdomen | [13 empty] | manual
|  | BodyAngio 3.0 MIP cor | [22 empty] | 3 | 3 | 34 | B30f medium smooth | CT Angio | [13 empty] | manual
CarotidAngioRoutine (Adult)
CarotidAngioRoutine (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| CarotidAngio |  | 120 | 120 | 120 | 120 |  |  | On | 11 | On | CARE Dose4D | 8.12 | 5s | 25% |  |  | 0.285 |  | 2 | 1.2 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | CarotidAngio 5.0 B30f | [22 empty] | 5 | 5 | 61 | B30f medium smooth | Larynx | [13 empty] | manual
|  | CarotidAngio 0.6 B26f | [22 empty] | 0.6 | 0.4 | 751 | B26f medium smooth ASA | CT Angio | [13 empty] | manual
CarotidAngio_IR (Adult)
CarotidAngio_IR (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| CarotidAngio |  | 120 | 120 | 84 | 84 |  |  | On | 11 | On | CARE Dose4D | 5.68 | 5s | 25% |  |  | 0.285 |  | 2 | 1.2 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | CarotidAngio 5.0 I30f 3 | [22 empty] | 5 | 5 | 61 | I30f medium smooth | Larynx | [13 empty] | manual
|  | CarotidAngio 0.6 I26f 3 | [22 empty] | 0.6 | 0.4 | 751 | I26f medium smooth ASA | CT Angio | [13 empty] | manual
CarotidAngio_Vol (Adult)
CarotidAngio_Vol (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| CarotidAngio |  | 120 | 120 | 120 | 120 |  |  | On | 11 | On | CARE Dose4D | 8.12 | 5s | 25% |  |  | 0.285 |  | 2 | 1.2 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | CarotidAngio 5.0 B30f | [22 empty] | 5 | 5 | 61 | B30f medium smooth | Larynx | [13 empty] | manual
|  | CarotidAngio 3.0 MIP cor | [22 empty] | 3 | 3 | 27 | B30f medium smooth | CT Angio | [13 empty] | manual
CarotidDSACT (Adult)
CarotidDSACT (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Non Contrast |  | 120 | 120 | 60 | 60 |  |  | On | 3 | On | CARE Dose4D | 4.06 | 5s | 25% |  |  | 0.285 |  | 2 | 1.2 |  | 0.6 | 128x0.6mm | [15 empty] | Off
|  | DSAnc Non Contrast 0.6 B26f | [22 empty] | 0.6 | 0.4 | 751 | B26f medium smooth ASA | CT Angio | [13 empty] | manual
| CarotidAngio |  | 120 | 120 | 120 | 120 |  |  | On | 11 | On | CARE Dose4D | 8.12 | 5s | 25% |  |  | 0.285 |  | 2 | 1.2 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | CarotidAngio 5.0 B30f | [22 empty] | 5 | 5 | 61 | B30f medium smooth | Larynx | [13 empty] | manual
|  | DSA CarotidAngio 0.6 B26f | [22 empty] | 0.6 | 0.4 | 751 | B26f medium smooth ASA | CT Angio | [13 empty] | manual
DSXXL_ChestPainECG (Adult)
DSXXL_ChestPainECG (Adult) | Topogram | Topogram 0.6 T20f |  | 140 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.2 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
DS_ChestECG_Abd (Adult)
DS_ChestECG_Abd (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 365 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| DS_ChestPain |  | 120 | 120 | 320 |  | 320 |  | On | 11 | On | CARE Dose4D | 104.63 |  | 25% |  |  | 0.285 |  | 2 | 0.17 |  | 3 | 128x0.6mm | [15 empty] | Advance | 70% | 70%
|  | Cor DS_ChestPain 3.0 B30f BestDiast | [22 empty] | 3 | 3 | 51 | B30f medium smooth | Cardiac | [13 empty] | BestDiast | 70%
|  | QCA_D DS_ChestPain 0.6 B26f BestDiast | [22 empty] | 0.6 | 0.3 | 501 | B26f medium smooth ASA | Cardiac | [13 empty] | BestDiast | 70%
|  | QCA_S DS_ChestPain 0.6 B26f BestSyst | [22 empty] | 0.6 | 0.3 | 501 | B26f medium smooth ASA | Cardiac | [13 empty] | BestSyst | 20%
|  | Tho DS_ChestPain 3.0 B30f BestDiast | [22 empty] | 3 | 3 | 84 | B30f medium smooth | Mediastinum | [13 empty] | BestDiast | 70%
|  | Lun DS_ChestPain 3.0 B70f BestDiast | [22 empty] | 3 | 3 | 84 | B70f very sharp | Lung | [13 empty] | BestDiast | 70%
|  | PE DS_ChestPain 1.0 B30f BestDiast | [22 empty] | 1 | 0.5 | 501 | B30f medium smooth | Mediastinum | [13 empty] | BestDiast | 70%
| BodyAngio |  | 120 | 120 | 120 | 120 |  |  | On | 11 | On | CARE Dose4D | 6.66 | 5s | 25% |  |  | 0.285 |  | 3 | 1.2 |  | 3 | 128x0.6mm | [15 empty] | Off
|  | BodyAngio 3.0 B30f | [22 empty] | 3 | 3 | 101 | B30f medium smooth | Abdomen | [13 empty] | manual
|  | BodyAngio 0.6 B26f | [22 empty] | 0.6 | 0.4 | 751 | B26f medium smooth ASA | CT Angio | [13 empty] | manual
DS_ChestPainECG (Adult)
DS_ChestPainECG (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| DS_ChestPain |  | 120 | 120 | 320 |  | 320 |  | On | 11 | On | CARE Dose4D | 104.63 |  | 25% |  |  | 0.285 |  | 2 | 0.17 |  | 3 | 128x0.6mm | [15 empty] | Advance | 70% | 70%
|  | Cor DS_ChestPain 3.0 B30f BestDiast | [22 empty] | 3 | 3 | 51 | B30f medium smooth | Cardiac | [13 empty] | BestDiast | 70%
|  | QCA_D DS_ChestPain 0.6 B26f BestDiast | [22 empty] | 0.6 | 0.3 | 501 | B26f medium smooth ASA | Cardiac | [13 empty] | BestDiast | 70%
|  | QCA_S DS_ChestPain 0.6 B26f BestSyst | [22 empty] | 0.6 | 0.3 | 501 | B26f medium smooth ASA | Cardiac | [13 empty] | BestSyst | 20%
|  | Tho DS_ChestPain 3.0 B30f BestDiast | [22 empty] | 3 | 3 | 84 | B30f medium smooth | Mediastinum | [13 empty] | BestDiast | 70%
|  | Lun DS_ChestPain 3.0 B70f BestDiast | [22 empty] | 3 | 3 | 84 | B70f very sharp | Lung | [13 empty] | BestDiast | 70%
|  | PE DS_ChestPain 1.0 B30f BestDiast | [22 empty] | 1 | 0.5 | 501 | B30f medium smooth | Mediastinum | [13 empty] | BestDiast | 70%
Embolism (Adult)
Embolism (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Embolism |  | 100 | 100 | 150 | 150 |  |  | On | 11 | On | CARE Dose4D | 6.18 | 5s | 25% |  |  | 0.285 |  | 2 | 1.2 |  | 3 | 128x0.6mm | [15 empty] | Off
|  | Embolism 3.0 B30f | [22 empty] | 3 | 3 | 101 | B30f medium smooth | Mediastinum | [13 empty] | manual
|  | Embolism 3.0 B70f | [22 empty] | 3 | 3 | 101 | B70f very sharp | Lung | [13 empty] | manual
|  | Embolism 1.0 B30f | [22 empty] | 1 | 0.5 | 601 | B30f medium smooth | CT Angio | [13 empty] | manual
Embolism_IR (Adult)
Embolism_IR (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Embolism |  | 100 | 100 | 105 | 105 |  |  | On | 11 | On | CARE Dose4D | 4.33 | 5s | 25% |  |  | 0.285 |  | 2 | 1.2 |  | 3 | 128x0.6mm | [15 empty] | Off
|  | Embolism 3.0 I30f 3 | [22 empty] | 3 | 3 | 101 | I30f medium smooth | Mediastinum | [13 empty] | manual
|  | Embolism 3.0 I70f 3 | [22 empty] | 3 | 3 | 101 | I70f very sharp ASA | Lung | [13 empty] | manual
|  | Embolism 1.0 I30f 3 | [22 empty] | 1 | 0.5 | 601 | I30f medium smooth | CT Angio | [13 empty] | manual
Flash_ChestPainECG (Adult)
Flash_ChestPainECG (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Fl_ChestPain |  | 120 | 120 | 320 |  | 320 |  | On | 11 | On | CARE Dose4D | 5.54 |  | 25% |  |  | 0.285 |  | 2 | 3.2 |  | 3 | 128x0.6mm
|  | Cor Fl_ChestPain 3.0 B30f | [22 empty] | 3 | 3 | 51 | B30f medium smooth | Cardiac
|  | QCA Fl_ChestPain 0.6 B26f 60% | [22 empty] | 0.6 | 0.3 | 501 | B26f medium smooth ASA | Cardiac | [14 empty] | 60%
|  | Tho Fl_ChestPain 3.0 B30f 60% | [22 empty] | 3 | 3 | 51 | B30f medium smooth | Mediastinum | [14 empty] | 60%
|  | Lun Fl_ChestPain 3.0 B70f 60% | [22 empty] | 3 | 3 | 84 | B70f very sharp | Lung | [14 empty] | 60%
|  | PE Fl_ChestPain 1.0 B30f 60% | [22 empty] | 1 | 0.5 | 501 | B30f medium smooth | Mediastinum | [14 empty] | 60%
| DS_CorAdSeq |  | 120 | 120 | 320 |  | 320 |  | On | 11 | On | CARE Dose4D | 69.17 |  | 25% |  |  | 0.285 |  | 2 | 0 | 34 | 3 | 128x0.6mm | [13 empty] | 70% | 70% | Normal | 70% | 70%
|  | Cor DS_CorAdSeq 3.0 B30f BestDiast | [22 empty] | 3 |  | 46 | B30f medium smooth | Cardiac | [13 empty] | BestDiast | 70%
|  | Cor DS_CorAdSeq 0.6 B26f BestDiast | [22 empty] | 0.6 |  | 344 | B26f medium smooth ASA | Cardiac | [13 empty] | BestDiast | 70%
Flash_Embolism (Adult)
Flash_Embolism (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Fl_Embolism |  | 100 | 100 | 172 |  | 172 |  | On | 11 | On | CARE Dose4D | 5.58 | 5s | 25% |  |  | 0.285 |  | 2 | 2.2 |  | 3 | 128x0.6mm
|  | Fl_Embolism 3.0 B30f | [22 empty] | 3 | 3 | 101 | B30f medium smooth | Mediastinum
|  | Fl_Embolism 3.0 B70f | [22 empty] | 3 | 3 | 101 | B70f very sharp | Lung
|  | Fl_Embolism 1.0 B30f | [22 empty] | 1 | 0.5 | 601 | B30f medium smooth | CT Angio
Flash_Embolism_IR (Adult)
Flash_Embolism_IR (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Fl_Embolism |  | 100 | 100 | 120 |  | 120 |  | On | 11 | On | CARE Dose4D | 3.98 | 5s | 25% |  |  | 0.285 |  | 2 | 2.2 |  | 3 | 128x0.6mm
|  | Fl_Embolism 3.0 I30f 3 | [22 empty] | 3 | 3 | 101 | I30f medium smooth | Mediastinum
|  | Fl_Embolism 3.0 I70f 3 | [22 empty] | 3 | 3 | 101 | I70f very sharp ASA | Lung
|  | Fl_Embolism 1.0 I30f 3 | [22 empty] | 1 | 0.5 | 601 | I30f medium smooth | CT Angio
Flash_ThorAngio (Adult)
Flash_ThorAngio (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Fl_ThorAngio |  | 100 | 100 | 160 |  | 160 |  | On | 11 | On | CARE Dose4D | 5.3 | 5s | 25% |  |  | 0.285 |  | 2 | 2.2 |  | 5 | 128x0.6mm
|  | Fl_ThorAngio 5.0 B30f | [22 empty] | 5 | 5 | 81 | B30f medium smooth | Mediastinum
|  | Fl_ThorAngio 5.0 B70f | [22 empty] | 5 | 5 | 81 | B70f very sharp | Lung
|  | Fl_ThorAngio 0.75 B26f | [22 empty] | 0.75 | 0.5 | 801 | B26f medium smooth ASA | CT Angio
Flash_ThorAngio_IR (Adult)
Flash_ThorAngio_IR (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Fl_ThorAngio |  | 100 | 100 | 112 |  | 112 |  | On | 11 | On | CARE Dose4D | 3.72 | 5s | 25% |  |  | 0.285 |  | 2 | 2.2 |  | 5 | 128x0.6mm
|  | Fl_ThorAngio 5.0 I30f 3 | [22 empty] | 5 | 5 | 81 | I30f medium smooth | Mediastinum
|  | Fl_ThorAngio 5.0 I70f 3 | [22 empty] | 5 | 5 | 81 | I70f very sharp ASA | Lung
|  | Fl_ThorAngio 0.75 I26f 3 | [22 empty] | 0.75 | 0.5 | 801 | I26f medium smooth ASA | CT Angio
HeadAngioRoutine (Adult)
HeadAngioRoutine (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.29 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head
| HeadAngio |  | 100 | 100 | 175 | 175 |  |  | On | 11 | On | CARE Dose4D | 16.91 | 5s | 25% |  |  | 0.5 |  | 2 | 1.2 |  | 4 | 128x0.6mm | [15 empty] | Off
|  | HeadAngio 4.0 H20f | [22 empty] | 4 | 4 | 31 | H20f smooth | Base Orbita | [13 empty] | manual
|  | HeadAngio 0.6 H10f | [22 empty] | 0.6 | 0.4 | 301 | H10f very smooth | CT Angio | [13 empty] | manual
HeadAngio_IR (Adult)
HeadAngio_IR (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.29 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head
| HeadAngio |  | 100 | 100 | 123 | 123 |  |  | On | 11 | On | CARE Dose4D | 11.92 | 5s | 25% |  |  | 0.5 |  | 2 | 1.2 |  | 4 | 128x0.6mm | [15 empty] | Off
|  | HeadAngio 4.0 J30f 3 | [22 empty] | 4 | 4 | 31 | J30f medium smooth | Base Orbita | [13 empty] | manual
|  | HeadAngio 0.6 J30f 3 | [22 empty] | 0.6 | 0.4 | 301 | J30f medium smooth | CT Angio | [13 empty] | manual
HeadAngio_Vol (Adult)
HeadAngio_Vol (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.29 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head
| HeadAngio |  | 100 | 100 | 175 | 175 |  |  | On | 11 | On | CARE Dose4D | 16.91 | 5s | 25% |  |  | 0.5 |  | 2 | 1.2 |  | 4 | 128x0.6mm | [15 empty] | Off
|  | HeadAngio 4.0 H20f | [22 empty] | 4 | 4 | 31 | H20f smooth | Base Orbita | [13 empty] | manual
|  | HeadAngio 4.0 MIP cor | [22 empty] | 4 | 4 | 21 | H10f very smooth | CT Angio | [13 empty] | manual
HeadAngio_XCARE (Adult)
HeadAngio_XCARE (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.29 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head
| HeadAngio |  | 100 | 100 | 175 | 175 |  |  | On | 11 | On | CARE Dose4D | 16.94 | 5s | 25% |  |  | 0.285 |  | 2 | 0.6 |  | 4 | 128x0.6mm | [15 empty] | Off
|  | HeadAngio 4.0 H20f | [22 empty] | 4 | 4 | 31 | H20f smooth | Base Orbita |  |  | X-CARE | [10 empty] | manual
|  | HeadAngio 0.6 H10f | [22 empty] | 0.6 | 0.4 | 301 | H10f very smooth | CT Angio |  |  | X-CARE | [10 empty] | manual
HeadAngio_XCARE_IR (Adult)
HeadAngio_XCARE_IR (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.29 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head
| HeadAngio |  | 100 | 100 | 123 | 123 |  |  | On | 11 | On | CARE Dose4D | 11.89 | 5s | 25% |  |  | 0.285 |  | 2 | 0.6 |  | 4 | 128x0.6mm | [15 empty] | Off
|  | HeadAngio 4.0 J30f 3 | [22 empty] | 4 | 4 | 31 | J30f medium smooth | Base Orbita |  |  | X-CARE | [10 empty] | manual
|  | HeadAngio 0.6 J30f 3 | [22 empty] | 0.6 | 0.4 | 301 | J30f medium smooth | CT Angio |  |  | X-CARE | [10 empty] | manual
NeuroDSACT (Adult)
NeuroDSACT (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.29 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head
| Non Contrast |  | 100 | 100 | 90 | 90 |  |  | On | 3 | On | CARE Dose4D | 8.7 | 5s | 25% |  |  | 0.5 |  | 2 | 1.2 |  | 0.6 | 128x0.6mm | [15 empty] | Off
|  | DSAnc Non Contrast 0.6 H10f | [22 empty] | 0.6 | 0.4 | 301 | H10f very smooth | CT Angio | [13 empty] | manual
| HeadAngio |  | 100 | 100 | 175 | 175 |  |  | On | 11 | On | CARE Dose4D | 16.91 | 5s | 25% |  |  | 0.5 |  | 2 | 1.2 |  | 4 | 128x0.6mm | [15 empty] | Off
|  | HeadAngio 4.0 H20f | [22 empty] | 4 | 4 | 31 | H20f smooth | Base Orbita | [13 empty] | manual
|  | DSA HeadAngio 0.6 H10f | [22 empty] | 0.6 | 0.4 | 301 | H10f very smooth | CT Angio | [13 empty] | manual
ThorAngioRoutine (Adult)
ThorAngioRoutine (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| ThorAngio |  | 80 | 80 | 224 | 224 |  |  | On | 11 | On | CARE Dose4D | 4.39 | 6s | 25% |  |  | 0.5 |  | 2 | 1.2 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | ThorAngio 5.0 B30f | [22 empty] | 5 | 5 | 81 | B30f medium smooth | Mediastinum | [13 empty] | manual
|  | ThorAngio 5.0 B70f | [22 empty] | 5 | 5 | 81 | B70f very sharp | Lung | [13 empty] | manual
|  | ThorAngio 0.75 B26f | [22 empty] | 0.75 | 0.5 | 801 | B26f medium smooth ASA | CT Angio | [13 empty] | manual
ThorAngio_IR (Adult)
ThorAngio_IR (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| ThorAngio |  | 80 | 80 | 157 | 157 |  |  | On | 11 | On | CARE Dose4D | 3.08 | 6s | 25% |  |  | 0.5 |  | 2 | 1.2 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | ThorAngio 5.0 I30f 3 | [22 empty] | 5 | 5 | 81 | I30f medium smooth | Mediastinum | [13 empty] | manual
|  | ThorAngio 5.0 I70f 3 | [22 empty] | 5 | 5 | 81 | I70f very sharp ASA | Lung | [13 empty] | manual
|  | ThorAngio 0.75 I26f 3 | [22 empty] | 0.75 | 0.5 | 801 | I26f medium smooth ASA | CT Angio | [13 empty] | manual
ThorAngio_Vol (Adult)
ThorAngio_Vol (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| ThorAngio |  | 80 | 80 | 224 | 224 |  |  | On | 11 | On | CARE Dose4D | 4.39 | 6s | 25% |  |  | 0.5 |  | 2 | 1.2 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | ThorAngio 5.0 B30f | [22 empty] | 5 | 5 | 81 | B30f medium smooth | Mediastinum | [13 empty] | manual
|  | ThorAngio 5.0 B70f | [22 empty] | 5 | 5 | 81 | B70f very sharp | Lung | [13 empty] | manual
|  | ThorAngio 3.0 MIP cor | [22 empty] | 3 | 3 | 29 | B30f medium smooth | CT Angio | [13 empty] | manual
ThorAngio_XCARE (Adult)
ThorAngio_XCARE (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| ThorAngio |  | 80 | 80 | 224 | 224 |  |  | On | 11 | On | CARE Dose4D | 4.39 | 6s | 25% |  |  | 0.285 |  | 2 | 0.6 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | ThorAngio 5.0 B30f | [22 empty] | 5 | 5 | 81 | B30f medium smooth | Mediastinum |  |  | X-CARE | [10 empty] | manual
|  | ThorAngio 5.0 B70f | [22 empty] | 5 | 5 | 81 | B70f very sharp | Lung |  |  | X-CARE | [10 empty] | manual
|  | ThorAngio 0.75 B26f | [22 empty] | 0.75 | 0.5 | 801 | B26f medium smooth ASA | CT Angio |  |  | X-CARE | [10 empty] | manual
ThorAngio_XCARE_IR (Adult)
ThorAngio_XCARE_IR (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| ThorAngio |  | 80 | 80 | 157 | 157 |  |  | On | 11 | On | CARE Dose4D | 3.08 | 6s | 25% |  |  | 0.285 |  | 2 | 0.6 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | ThorAngio 5.0 I30f 3 | [22 empty] | 5 | 5 | 81 | I30f medium smooth | Mediastinum |  |  | X-CARE | [10 empty] | manual
|  | ThorAngio 5.0 I70f 3 | [22 empty] | 5 | 5 | 81 | I70f very sharp ASA | Lung |  |  | X-CARE | [10 empty] | manual
|  | ThorAngio 0.75 I26f 3 | [22 empty] | 0.75 | 0.5 | 801 | I26f medium smooth ASA | CT Angio |  |  | X-CARE | [10 empty] | manual
WholeBodyAngio (Adult)
WholeBodyAngio (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 365 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| WholeBody |  | 120 | 120 | 120 | 120 |  |  | On | 11 | On | CARE Dose4D | 8.11 | 11s | 25% |  |  | 0.5 |  | 2 | 1.2 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | WholeBody 5.0 B30f | [22 empty] | 5 | 5 | 161 | B30f medium smooth | Abdomen | [13 empty] | manual
|  | WholeBody 0.75 B26f | [22 empty] | 0.75 | 0.5 | 1601 | B26f medium smooth ASA | CT Angio | [13 empty] | manual
WholeBodyAngio_IR (Adult)
WholeBodyAngio_IR (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 365 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| WholeBody |  | 120 | 120 | 84 | 84 |  |  | On | 11 | On | CARE Dose4D | 5.69 | 11s | 25% |  |  | 0.5 |  | 2 | 1.2 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | WholeBody 5.0 I30f 3 | [22 empty] | 5 | 5 | 161 | I30f medium smooth | Abdomen | [13 empty] | manual
|  | WholeBody 0.75 I26f 3 | [22 empty] | 0.75 | 0.5 | 1601 | I26f medium smooth ASA | CT Angio | [13 empty] | manual
RT
RT_Abdomen (Adult)
RT_Abdomen (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Abdomen |  | 120 | 120 | 250 | 250 |  |  | Semi | 3 | On | CARE Dose4D | 17.28 | 99s | 25% |  |  | 1 |  | 2 | 0.6 |  | 4 | 32x1.2mm | [15 empty] | Off
|  | Abdomen 4.0 B30s | [22 empty] | 4 | 4 | 51 | B30s medium smooth | Abdomen | [13 empty] | manual
RT_Head (Adult)
RT_Head (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.29 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head
| Head |  | 120 | 120 | 410 | 410 |  |  | Off | 3 | On | CARE Dose4D | 64.03 | 99s | 25% |  |  | 1 |  | 2 | 0.55 |  | 3 | 32x1.2mm | [15 empty] | Off
|  | Head 3.0 H31s | [22 empty] | 3 | 3 | 41 | H31s medium smooth + | Cerebrum | [13 empty] | manual
RT_HeadNeckShoulder (Adult)
RT_HeadNeckShoulder (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| HeadNeckShoul |  | 120 | 120 | 200 | 200 |  |  | Semi | 3 | On | CARE Dose4D | 16.69 | 99s | 25% |  |  | 1 |  | 2 | 0.8 |  | 4 | 32x1.2mm | [15 empty] | Off
|  | HeadNeckShoul 4.0 B30s | [22 empty] | 4 | 4 | 38 | B30s medium smooth | Larynx | [13 empty] | manual
RT_Pelvis (Adult)
RT_Pelvis (Adult) | Topogram | Topogram 0.6 T80f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T80f sharp | Topogram Body
| Pelvis |  | 120 | 120 | 270 | 270 |  |  | Semi | 3 | On | CARE Dose4D | 18.65 | 99s | 25% |  |  | 1 |  | 2 | 0.8 |  | 4 | 32x1.2mm | [15 empty] | Off
|  | Pelvis 4.0 B30s | [22 empty] | 4 | 4 | 51 | B30s medium smooth | Pelvis | [13 empty] | manual
RT_Thorax (Adult)
RT_Thorax (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Thorax |  | 120 | 120 | 140 | 140 |  |  | Semi | 3 | On | CARE Dose4D | 9.68 | 99s | 25% |  |  | 1 |  | 2 | 1.2 |  | 4 | 32x1.2mm | [15 empty] | Off
|  | Thorax 4.0 B30s | [22 empty] | 4 | 4 | 76 | B30s medium smooth | Mediastinum | [13 empty] | manual
DualEnergy
DE_AbdomenSeq (Adult)
DE_AbdomenSeq (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| DE_AbdSeq |  |  | 100 | 171 |  | 171 | 132 |  | 3 | On | CARE Dose4D | 14.96 |  | 25% |  |  | 0.5 |  | 2 | 0 | 17 | 5 | 32x0.6mm | [15 empty] | Off
|  | DE_AbdSeq 5.0 B31f | [22 empty] | 5 |  | 37 | B31f medium smooth + | Abdomen | [13 empty] | manual
|  | DE #PP DE_AbdSeq 1.5 D30f | [22 empty] | 1.5 |  | 186 | D30f medium smooth | Abdomen | [13 empty] | manual
DE_Abdomen_KidneyStones (Adult)
DE_Abdomen_KidneyStones (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| DE_Abdomen |  |  | 100 | 210 |  | 210 | 162 |  | 3 | On | CARE Dose4D | 16.67 | 30s | 25% |  |  | 0.33 |  | 2 | 0.7 |  | 5 | 64x0.6mm | [15 empty] | Off
|  | DE_Abdomen 5.0 B30f | [22 empty] | 5 | 5 | 41 | B30f medium smooth | Abdomen | [13 empty] | manual
|  | DE #PP DE_Abdomen 1.5 D30f | [22 empty] | 1.5 | 1 | 201 | D30f medium smooth | Abdomen | [13 empty] | manual
DE_Abdomen_VNC (Adult)
DE_Abdomen_VNC (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| DE_Abdomen |  |  | 100 | 230 |  | 230 | 178 |  | 3 | On | CARE Dose4D | 17.88 | 30s | 25% |  |  | 0.5 |  | 2 | 0.6 |  | 5 | 32x0.6mm | [15 empty] | Off
|  | DE_Abdomen 5.0 B30f | [22 empty] | 5 | 5 | 41 | B30f medium smooth | Abdomen | [13 empty] | manual
|  | DE #PP DE_Abdomen 1.5 D30f | [22 empty] | 1.5 | 1 | 201 | D30f medium smooth | Abdomen | [13 empty] | manual
DE_BodyAngioFast_BoneRem (Adult)
DE_BodyAngioFast_BoneRem (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| DE_BodyAngio |  |  | 100 | 119 |  | 119 | 92 |  | 3 | On | CARE Dose4D | 8.28 | 6s | 25% |  |  | 0.33 |  | 2 | 0.7 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | DE_BodyAngio 5.0 B30f | [22 empty] | 5 | 5 | 81 | B30f medium smooth | Abdomen | [13 empty] | manual
|  | DE #PP DE_BodyAngio 1.5 D26f | [22 empty] | 1.5 | 1 | 401 | D26f med. smooth ASA | Abdomen | [13 empty] | manual
DE_BodyAngio_BoneRem (Adult)
DE_BodyAngio_BoneRem (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| DE_BodyAngio |  |  | 100 | 119 |  | 119 | 92 |  | 3 | On | CARE Dose4D | 9.29 | 15s | 25% |  |  | 0.5 |  | 2 | 0.9 |  | 5 | 32x0.6mm | [15 empty] | Off
|  | DE_BodyAngio 5.0 B30f | [22 empty] | 5 | 5 | 81 | B30f medium smooth | Abdomen | [13 empty] | manual
|  | DE #PP DE_BodyAngio 1.5 D26f | [22 empty] | 1.5 | 1 | 401 | D26f med. smooth ASA | Abdomen | [13 empty] | manual
DE_BodyAngio_BoneRem_IR (Adult)
DE_BodyAngio_BoneRem_IR (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| DE_BodyAngio |  |  | 100 | 95 |  | 95 | 73 |  | 3 | On | CARE Dose4D | 7.39 | 15s | 25% |  |  | 0.5 |  | 2 | 0.9 |  | 5 | 32x0.6mm | [15 empty] | Off
|  | DE_BodyAngio 5.0 I30f 3 | [22 empty] | 5 | 5 | 81 | I30f medium smooth | Abdomen | [13 empty] | manual
|  | DE #PP DE_BodyAngio 1.5 Q30f 3 | [22 empty] | 1.5 | 1 | 401 | Q30f medium smooth | Abdomen | [13 empty] | manual
DE_Body_BoneMarrow (Adult)
DE_Body_BoneMarrow (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 365 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Topogram | Topogram 0.6 T20f |  | 120 |  | 365 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| DE_Body |  |  | 100 | 230 |  | 230 | 178 |  | 3 | On | CARE Dose4D | 17.88 | 30s | 25% |  |  | 0.5 |  | 2 | 0.6 |  | 5 | 32x0.6mm | [15 empty] | Off
|  | DE_Body 5.0 B30f | [22 empty] | 5 | 5 | 141 | B30f medium smooth | Abdomen | [13 empty] | manual
|  | DE #PP DE_Body 1.5 D34f | [22 empty] | 1.5 | 1 | 701 | D34f med. smooth BHCBone | Abdomen | [13 empty] | manual
DE_Body_OptiCon (Adult)
DE_Body_OptiCon (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| DE_Body |  |  | 140 | 32 |  | 32 | 176 |  | 3 | On | CARE Dose4D | 8.67 | 19s | 25% |  |  | 0.5 |  | 2 | 0.7 |  | 5 | 32x0.6mm | [15 empty] | Off
|  | DE_Body 5.0 B30f | [22 empty] | 5 | 5 | 81 | B30f medium smooth | Abdomen | [13 empty] | manual
|  | DE #PP DE_Body 1.5 D30f | [22 empty] | 1.5 | 1 | 401 | D30f medium smooth | Abdomen | [13 empty] | manual
DE_CarotidAngioFast_BoneRem (Adult)
DE_CarotidAngioFast_BoneRem (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| DE_CarotidAngio |  |  | 100 | 104 |  | 104 | 104 |  | 3 | On | CARE Dose4D | 8.09 | 10s | 25% |  |  | 0.285 |  | 2 | 0.7 |  | 5 | 128x0.6mm | [15 empty] | Off
|  | DE_CarotidAngio 5.0 B30f | [22 empty] | 5 | 5 | 51 | B30f medium smooth | Larynx | [13 empty] | manual
|  | DE #PP DE_CarotidAngio 1.0 D26f | [22 empty] | 1 | 0.7 | 358 | D26f med. smooth ASA | Larynx | [13 empty] | manual
DE_CarotidAngio_BoneRem (Adult)
DE_CarotidAngio_BoneRem (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| DE_CarotidAngio |  |  | 100 | 114 |  | 114 | 114 |  | 3 | On | CARE Dose4D | 8.03 | 10s | 25% |  |  | 0.285 |  | 2 | 0.9 |  | 5 | 64x0.6mm | [15 empty] | Off
|  | DE_CarotidAngio 5.0 B30f | [22 empty] | 5 | 5 | 51 | B30f medium smooth | Larynx | [13 empty] | manual
|  | DE #PP DE_CarotidAngio 1.0 D26f | [22 empty] | 1 | 0.7 | 358 | D26f med. smooth ASA | Larynx | [13 empty] | manual
DE_CoronaryCTA_HeartPBV (Adult)
DE_CoronaryCTA_HeartPBV (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| DE_CorCTA |  |  | 100 | 165 |  | 165 | 140 |  | 3 | On | CARE Dose4D | 64.62 |  | 25% |  |  | 0.285 |  | 2 | 0.17 |  | 3 | 64x0.6mm | [15 empty] | Normal | 60% | 80%
|  | DE_CorCTA 3.0 B30f 70% | [22 empty] | 3 | 3 | 41 | B30f medium smooth | Cardiac | [13 empty] | manual | 70%
|  | QCA DE_CorCTA 0.6 B26f 70% | [22 empty] | 0.6 | 0.3 | 401 | B26f medium smooth ASA | Cardiac | [13 empty] | manual | 70%
|  | DE #PP DE_CorCTA 1.0 D33f 60% | [22 empty] | 1 | 0.7 | 172 | D33f med. smooth BHCIodine | Cardiac | [13 empty] | manual | 60%
DE_CoronaryCTA_Seq (Adult)
DE_CoronaryCTA_Seq (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| DE_CorSeq |  |  | 100 | 165 |  | 165 | 140 |  | 3 | On | CARE Dose4D | 12.16 |  | 25% |  |  | 0.285 |  | 2 | 0 | 17 | 3 | 64x0.6mm | [13 empty] | 60% | 60% | Off
|  | DE_CorSeq 3.0 B30f 60% | [22 empty] | 3 |  | 39 | B30f medium smooth | Cardiac | [13 empty] | manual | 60%
|  | DE #PP DE_CorSeq 1.5 D33f 60% | [22 empty] | 1.5 |  | 118 | D33f med. smooth BHCIodine | Cardiac | [13 empty] | manual | 60%
DE_ExtrAngio_BoneRem (Adult)
DE_ExtrAngio_BoneRem (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 365 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Topogram | Topogram 0.6 T20f |  | 120 |  | 365 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| DE_ExtrAngio |  |  | 80 | 230 |  | 230 | 98 |  | 3 | On | CARE Dose4D | 9.24 | 30s | 25% |  |  | 0.33 |  | 2 | 0.7 |  | 5 | 64x0.6mm | [15 empty] | Off
|  | DE_ExtrAngio 5.0 B30f | [22 empty] | 5 | 5 | 201 | B30f medium smooth | Abdomen | [13 empty] | manual
|  | DE #PP DE_ExtrAngio 1.5 D30f | [22 empty] | 1.5 | 1 | 1001 | D30f medium smooth | Abdomen | [13 empty] | manual
DE_Foot_Gout (Adult)
DE_Foot_Gout (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| DE_Foot |  |  | 80 | 250 |  | 250 | 125 |  | 3 | On | CARE Dose | 10.8 | 10s | 25% |  |  | 0.33 |  | 2 | 0.7 |  | 2 | 64x0.6mm | [15 empty] | Off
|  | DE_Foot 2.0 B20f | [22 empty] | 2 | 2 | 76 | B20f smooth | Extremity | [13 empty] | manual
|  | DE #PP DE_Foot 0.75 D34f | [22 empty] | 0.75 | 0.5 | 301 | D34f med. smooth BHCBone | Bone | [13 empty] | manual
DE_Hand_Gout (Adult)
DE_Hand_Gout (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| DE_Hand |  |  | 80 | 205 |  | 205 | 103 |  | 3 | On | CARE Dose | 8.89 | 10s | 25% |  |  | 0.33 |  | 2 | 0.7 |  | 2 | 64x0.6mm | [15 empty] | Off
|  | DE_Hand 2.0 B20f | [22 empty] | 2 | 2 | 76 | B20f smooth | Extremity | [13 empty] | manual
|  | DE #PP DE_Hand 0.75 D34f | [22 empty] | 0.75 | 0.5 | 301 | D34f med. smooth BHCBone | Extremity | [13 empty] | manual
DE_HeadAngio_BoneRem (Adult)
DE_HeadAngio_BoneRem (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.29 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head
| Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.29 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head
| DE_HeadAngio |  |  | 80 | 222 |  | 222 | 111 |  | 3 | On | CARE Dose4D | 19.4 | 10s | 25% |  |  | 0.33 |  | 2 | 0.7 |  | 4 | 64x0.6mm | [15 empty] | Off
|  | DE_HeadAngio 4.0 H20f | [22 empty] | 4 | 4 | 31 | H20f smooth | Base Orbita | [13 empty] | manual
|  | DE #PP DE_HeadAngio 1.0 D26f | [22 empty] | 1 | 0.7 | 172 | D26f med. smooth ASA | Larynx | [13 empty] | manual
DE_Head_BrainHem (Adult)
DE_Head_BrainHem (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.29 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head
| Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.29 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head
| DE_Head |  |  | 80 | 310 |  | 310 | 155 |  | 3 | On | CARE Dose4D | 28.96 | 16s | 25% |  |  | 0.5 |  | 2 | 0.7 |  | 4 | 40x0.6mm | [15 empty] | Off
|  | DE_Head 4.0 H20f | [22 empty] | 4 | 4 | 31 | H20f smooth | Base Orbita | [13 empty] | manual
|  | DE #PP DE_Head 1.0 D34f | [22 empty] | 1 | 0.7 | 172 | D34f med. smooth BHCBone | Base Orbita | [13 empty] | manual
DE_Knee_BoneMarrow (Adult)
DE_Knee_BoneMarrow (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| DE_Knee |  |  | 80 | 240 |  | 240 | 120 |  | 3 | On | CARE Dose | 10.18 | 8s | 25% |  |  | 0.5 |  | 2 | 0.7 |  | 3 | 32x0.6mm | [15 empty] | Off
|  | DE_Knee 3.0 B30f | [22 empty] | 3 | 3 | 51 | B30f medium smooth | Extremity | [13 empty] | manual
|  | DE #PP DE_Knee 0.75 D30f | [22 empty] | 0.75 | 0.5 | 301 | D30f medium smooth | Bone | [13 empty] | manual
DE_Knee_Tendon (Adult)
DE_Knee_Tendon (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| DE_Knee |  |  | 140 | 45 |  | 45 | 191 |  | 3 | On | CARE Dose | 11.57 | 23s | 25% |  |  | 1 |  | 2 | 0.7 |  | 2 | 40x0.6mm | [15 empty] | Off
|  | DE_Knee 2.0 B20s | [22 empty] | 2 | 2 | 76 | B20s smooth | Extremity | [13 empty] | manual
|  | DE #PP DE_Knee 0.75 D34s | [22 empty] | 0.75 | 0.5 | 301 | D34s med. smooth BHCBone | Bone | [13 empty] | manual
DE_Spine_Metal_Vol (Adult)
DE_Spine_Metal_Vol (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| DE_Spine |  |  | 100 | 250 |  | 250 | 483 |  | 3 | Off | CARE Dose4D | 31.12 | 40s | 25% |  |  | 0.5 |  | 2 | 0.5 |  | 3 | 32x0.6mm | [15 empty] | Off
|  | DE_Spine 3.0 B30f | [22 empty] | 3 | 3 | 54 | B30f medium smooth | Spine | [13 empty] | manual
|  | DE_Spine 1.5 D30f | [22 empty] | 1.5 | 1 | 161 | D30f medium smooth | Spine | [13 empty] | manual
|  | DE_Spine 1.5 B20f | [22 empty] | 1.5 | 1 | 161 | B20f smooth | Spine | [13 empty] | manual
|  | DE_Spine 2.0 MPR spine multi | [22 empty] | 2 | 2 | 0 | B20f smooth | Spine | [13 empty] | manual
|  | DE_Spine 2.0 MPR spine multi | [22 empty] | 2 | 2 | 0 | B60f sharp | Spine | [13 empty] | manual
DE_Thorax_Analysis (Adult)
DE_Thorax_Analysis (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| DE_Thorax |  |  | 100 | 89 |  | 89 | 76 |  | 3 | On | CARE Dose4D | 7.4 | 10s | 25% |  |  | 0.285 |  | 2 | 0.55 |  | 5 | 64x0.6mm | [15 empty] | Off
|  | DE_Thorax 5.0 B31f | [22 empty] | 5 | 5 | 61 | B31f medium smooth + | Mediastinum | [13 empty] | manual
|  | DE_Thorax 5.0 B80f | [22 empty] | 5 | 5 | 61 | B80f ultra sharp | Lung | [13 empty] | manual
|  | DE #PP DE_Thorax 2.0 D33f | [22 empty] | 2 | 1.4 | 215 | D33f med. smooth BHCIodine | Mediastinum | [13 empty] | manual
DE_Thorax_Analysis_IR (Adult)
DE_Thorax_Analysis_IR (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| DE_Thorax |  |  | 100 | 71 |  | 71 | 60 |  | 3 | On | CARE Dose4D | 5.89 | 10s | 25% |  |  | 0.285 |  | 2 | 0.55 |  | 5 | 64x0.6mm | [15 empty] | Off
|  | DE_Thorax 5.0 I31f 3 | [22 empty] | 5 | 5 | 61 | I31f medium smooth | Mediastinum | [13 empty] | manual
|  | DE_Thorax 5.0 I70f 3 | [22 empty] | 5 | 5 | 61 | I70f very sharp ASA | Lung | [13 empty] | manual
|  | DE #PP DE_Thorax 2.0 Q30f 3 | [22 empty] | 2 | 1.4 | 215 | Q30f medium smooth | Mediastinum | [13 empty] | manual
DE_Thorax_Analysis_XCARE (Adult)
DE_Thorax_Analysis_XCARE (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| DE_Thorax |  |  | 100 | 89 |  | 89 | 76 |  | 3 | On | CARE Dose4D | 7.4 | 10s | 25% |  |  | 0.285 |  | 2 | 0.55 |  | 5 | 64x0.6mm | [15 empty] | Off
|  | DE_Thorax 5.0 B31f | [22 empty] | 5 | 5 | 61 | B31f medium smooth + | Mediastinum |  |  | X-CARE | [10 empty] | manual
|  | DE_Thorax 5.0 B80f | [22 empty] | 5 | 5 | 61 | B80f ultra sharp | Lung |  |  | X-CARE | [10 empty] | manual
|  | DE #PP DE_Thorax 2.0 D30f | [22 empty] | 2 | 1.4 | 215 | D30f medium smooth | Mediastinum |  |  | X-CARE | [10 empty] | manual
DE_Thorax_LungNodules (Adult)
DE_Thorax_LungNodules (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| DE_Thorax |  |  | 80 | 205 |  | 205 | 87 |  | 3 | On | CARE Dose4D | 8.25 | 10s | 25% |  |  | 0.285 |  | 2 | 0.55 |  | 5 | 64x0.6mm | [15 empty] | Off
|  | DE_Thorax 5.0 B31f | [22 empty] | 5 | 5 | 61 | B31f medium smooth + | Mediastinum | [13 empty] | manual
|  | DE_Thorax 5.0 B80f | [22 empty] | 5 | 5 | 61 | B80f ultra sharp | Lung | [13 empty] | manual
|  | DE #PP DE_Thorax 1.5 D30f | [22 empty] | 1.5 | 1 | 301 | D30f medium smooth | Mediastinum | [13 empty] | manual
DE_Thorax_Xenon (Adult)
DE_Thorax_Xenon (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Topogram | Topogram 0.6 T20f |  | 120 |  | 185 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| DE_Thorax |  |  | 100 | 89 |  | 89 | 76 |  | 3 | On | CARE Dose4D | 7.4 | 10s | 25% |  |  | 0.285 |  | 2 | 0.55 |  | 5 | 64x0.6mm | [15 empty] | Off
|  | DE_Thorax 5.0 B31f | [22 empty] | 5 | 5 | 61 | B31f medium smooth + | Mediastinum | [13 empty] | manual
|  | DE_Thorax 5.0 B80f | [22 empty] | 5 | 5 | 61 | B80f ultra sharp | Lung | [13 empty] | manual
|  | DE #PP DE_Thorax 2.0 D34f | [22 empty] | 2 | 1.4 | 215 | D34f med. smooth BHCBone | Lung | [13 empty] | manual
DE_Wrist_Tendon (Adult)
DE_Wrist_Tendon (Adult) | Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| Topogram | Topogram 0.6 T20f |  | 120 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.14 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body
| DE_Wrist |  |  | 140 | 37 |  | 37 | 157 |  | 3 | On | CARE Dose | 8.85 | 5s | 25% |  |  | 0.33 |  | 2 | 0.7 |  | 2 | 64x0.6mm | [15 empty] | Off
|  | DE_Wrist 2.0 B20f | [22 empty] | 2 | 2 | 31 | B20f smooth | Extremity | [13 empty] | manual
|  | DE #PP DE_Wrist 0.75 D34f | [22 empty] | 0.75 | 0.5 | 121 | D34f med. smooth BHCBone | Bone | [13 empty] | manual

## Sheet1
Protocol name | Range name | Series description | Ref. kV | kV | Quality ref. mAs | (Eff.) mAs | (Eff.) mAs(Tube A) | (Eff.) mAs(Tube B) | CARE kV | Tissue of Interest (Slider position 1 - 12) | Dose modulation | CARE Dose type | CTDIvol (mGy) | FAST Adjust: Upper limit scan time | FAST Adjust: Lower limit max. mAs | Dose Notification value CTDIvol (mGy) | Dose Notification value DLP (mGy*cm) | Rotation time (s) | Scan time (s) | Delay (s) | Pitch | Feed per scan (mm) | Slice (mm) | Acq. | Effective Slice (mm) | Position increment (mm) | No. of images | Kernel/Algorithm | Window | API | Comment 1 | Comment 2 | Auto transfer 1 | Auto transfer 2 | Auto transfer 3 | syngo.via: Workflow | syngo.via: Data role | Scan start | Scan end | Pulsing | Pulsing start | Pulsing end | Best phase | Phase start | Multi phase
1150_Hea_wo_cst_N09_02A (Adult)
1150_Hea_wo_cst_N09_02A (Adult) | Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.26 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS |  | TRACKER | [9 empty] | None
| Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.26 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS |  | TRACKER | [9 empty] | None
| Hea |  | 120 | 120 | 390 | 290 |  |  | Off | 3 | Off | CARE Dose4D | 44.37 | 15s | 25% | 80 |  | 0.5 |  | 2 | 0.6 |  | 4 | 128x0.6mm | [15 empty] | Off |  |  |  |  | None
|  | Hea h150 | [22 empty] | 4 | 4 | 31 | J40f medium | Head Soft |  |  | X-CARE | GEPACS | [9 empty] | manual |  | None
|  | Hea bon | [22 empty] | 4 | 4 | 31 | H60f sharp FR | Bone |  |  | X-CARE | GEPACS | [9 empty] | manual |  | None
|  | Hea cor | [22 empty] | 3 | 3 | 84 | J40f medium | Head Soft |  |  | X-CARE | GEPACS | [9 empty] | manual |  | None
|  | Hea sag | [22 empty] | 3 | 3 | 84 | J40f medium | Head Soft |  |  | X-CARE | GEPACS | [9 empty] | manual |  | None
|  | Hea h150 :-D | [22 empty] | 1 | 1 | 121 | J40f medium | Head Soft |  |  | X-CARE | ROUTER | GEPACS | [8 empty] | manual |  | None
5140_Che_2 (Adult)
5140_Che_2 (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER | [9 empty] | None
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER | [9 empty] | None
| art Che |  | 120 | 120 | 150 | 110 |  |  | On | 7 | On | CARE Dose4D | 10.12 | 30s | 25% | 50 |  | 0.5 |  | 25 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off |  |  |  |  | None
|  | art Che | [22 empty] | 3 | 2 | 145 | I41f medium | Chest |  |  |  | GEPACS | [9 empty] | manual |  | None
|  | art Che sag | [22 empty] | 3 | 2 | 191 | I41f medium | Bone |  |  |  | GEPACS | [9 empty] | manual |  | None
|  | art Che mip | [22 empty] | 10 | 3 | 121 | I41f medium | Lung |  |  |  | GEPACS | [9 empty] | manual |  | None
|  | art Che c140 3dt | [22 empty] | 1 | 1 | 289 | I50f medium sharp ASA | Lung |  |  |  | ROUTER | GEPACS | [8 empty] | manual |  | None
5170_Che_2_AbdPel (Adult)
5170_Che_2_AbdPel (Adult) | Topogram | Topogram |  | 100 |  | 393 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 393 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| art Che |  | 120 | 120 | 150 | 110 |  |  | On | 7 | On | CARE Dose4D | 10.12 | 30s | 25% | 50 |  | 0.5 |  | 25 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | art Che | [22 empty] | 3 | 2 | 145 | I41f medium | Chest |  |  |  | GEPACS | [9 empty] | manual
|  | art Che sag | [22 empty] | 3 | 2 | 191 | I41f medium | Bone |  |  |  | GEPACS | [9 empty] | manual
|  | art Che mip | [22 empty] | 10 | 3 | 121 | I41f medium | Lung |  |  |  | GEPACS | [9 empty] | manual
|  | art Che c140 3dt | [22 empty] | 1 | 1 | 289 | I50f medium sharp ASA | Lung |  |  |  | ROUTER | GEPACS | [8 empty] | manual
| ven AbdPel |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 60 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven AbdPel a030 | [22 empty] | 3 | 3 | 131 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel cor | [22 empty] | 3 | 3 | 67 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel sag | [22 empty] | 3 | 3 | 121 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel a030 :-D | [22 empty] | 0.6 | 0.6 | 653 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
6030_AbdPel_A02_00A (Adult)
6030_AbdPel_A02_00A (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
| ven AbdPel |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 90 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
|  | ven AbdPel a030 | [22 empty] | 3 | 3 | 131 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel cor | [22 empty] | 3 | 3 | 67 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel sag | [22 empty] | 3 | 3 | 121 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
|  | ven AbdPel a030 :-D | [22 empty] | 0.6 | 0.6 | 653 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
9060_Foo_sim_wo_cst (Adult)
9060_Foo_sim_wo_cst (Adult) | Topogram PA | Topogram PA 0.6 T20s |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.09 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER | [9 empty] | None
| Topogram LAT | Topogram LAT 0.6 T20s |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.09 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER | [9 empty] | None
| Foot simple |  |  | 120 | 120 | 250 |  |  |  | 3 | Off | CARE Dose | 24.04 | 30s | 25% | 60 |  | 1 |  | 2 | 1 |  | 0.6 | 16x0.3mm | [15 empty] | Off |  |  |  |  | None
|  | Foo bon l060 :-D | [22 empty] | 0.6 | 0.4 | 587 | U70u sharp | Bone |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual |  | None
|  | Foo bon l060 | [22 empty] | 2 | 2 | 117 | U70u sharp | Bone |  |  |  | GEPACS | [9 empty] | manual |  | None
|  | Foo :-D | [22 empty] | 0.6 | 0.4 | 584 | U30u medium smooth | Soft Tissue |  |  |  | ROUTER | GEPACS | [8 empty] | manual |  | None
|  | Foo | [22 empty] | 2 | 2 | 117 | U30u medium smooth | Soft Tissue |  |  |  | GEPACS | [9 empty] | manual |  | None

## Sheet3
Protocol name | Range name | Series description | Ref. kV | kV | Quality ref. mAs | (Eff.) mAs | (Eff.) mAs(Tube A) | (Eff.) mAs(Tube B) | CARE kV | Tissue of Interest (Slider position 1 - 12) | Dose modulation | CARE Dose type | CTDIvol (mGy) | FAST Adjust: Upper limit scan time | FAST Adjust: Lower limit max. mAs | Dose Notification value CTDIvol (mGy) | Dose Notification value DLP (mGy*cm) | Rotation time (s) | Scan time (s) | Delay (s) | Pitch | Feed per scan (mm) | Slice (mm) | Acq. | Effective Slice (mm) | Position increment (mm) | No. of images | Kernel/Algorithm | Window | API | Comment 1 | Comment 2 | Auto transfer 1 | Auto transfer 2 | Auto transfer 3 | syngo.via: Workflow | syngo.via: Data role | Scan start | Scan end | Pulsing | Pulsing start | Pulsing end | Best phase | Phase start | Multi phase
9060_Foo_sim_wo_cst (Adult)
9060_Foo_sim_wo_cst (Adult) | Topogram PA | Topogram PA 0.6 T20s |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.09 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER | [9 empty] | None
| Topogram LAT | Topogram LAT 0.6 T20s |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.09 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER | [9 empty] | None
| Foot simple |  |  | 120 | 120 | 250 |  |  |  | 3 | Off | CARE Dose | 24.04 | 30s | 25% | 60 |  | 1 |  | 2 | 1 |  | 0.6 | 16x0.3mm | [15 empty] | Off |  |  |  |  | None
|  | Foo bon l060 :-D | [22 empty] | 0.6 | 0.4 | 587 | U70u sharp | Bone |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual |  | None
|  | Foo bon l060 | [22 empty] | 2 | 2 | 117 | U70u sharp | Bone |  |  |  | GEPACS | [9 empty] | manual |  | None
|  | Foo :-D | [22 empty] | 0.6 | 0.4 | 584 | U30u medium smooth | Soft Tissue |  |  |  | ROUTER | GEPACS | [8 empty] | manual |  | None
|  | Foo | [22 empty] | 2 | 2 | 117 | U30u medium smooth | Soft Tissue |  |  |  | GEPACS | [9 empty] | manual |  | None

## Indexing
| Protocol name | Range name | Series description | Ref. kV | kV | Quality ref. mAs | (Eff.) mAs | (Eff.) mAs(Tube A) | (Eff.) mAs(Tube B) | CARE kV | Tissue of Interest (Slider position 1 - 12) | Dose modulation | CARE Dose type | CTDIvol (mGy) | FAST Adjust: Upper limit scan time | FAST Adjust: Lower limit max. mAs | Dose Notification value CTDIvol (mGy) | Dose Notification value DLP (mGy*cm) | Rotation time (s) | Scan time (s) | Delay (s) | Pitch | Feed per scan (mm) | Slice (mm) | Acq. | Effective Slice (mm) | Position increment (mm) | No. of images | Kernel/Algorithm | Window | API | Comment 1 | Comment 2 | Auto transfer 1 | Auto transfer 2 | Auto transfer 3 | syngo.via: Workflow | syngo.via: Data role | Scan start | Scan end | Pulsing | Pulsing start | Pulsing end | Best phase | Phase start | Multi phase
0 | 1150_Hea_wo_cst_N09_02A (Adult)
1 | 1150_Hea_wo_cst_N09_02A (Adult) | Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.26 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS |  | TRACKER | [9 empty] | None
2 |  | Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.26 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS |  | TRACKER | [9 empty] | None
3 |  | Hea |  | 120 | 120 | 390 | 290 |  |  | Off | 3 | Off | CARE Dose4D | 44.37 | 15s | 25% | 80 |  | 0.5 |  | 2 | 0.6 |  | 4 | 128x0.6mm | [15 empty] | Off |  |  |  |  | None
4 |  |  | Hea h150 | [22 empty] | 4 | 4 | 31 | J40f medium | Head Soft |  |  | X-CARE | GEPACS | [9 empty] | manual |  | None
5 |  |  | Hea bon | [22 empty] | 4 | 4 | 31 | H60f sharp FR | Bone |  |  | X-CARE | GEPACS | [9 empty] | manual |  | None
6 |  |  | Hea cor | [22 empty] | 3 | 3 | 84 | J40f medium | Head Soft |  |  | X-CARE | GEPACS | [9 empty] | manual |  | None
7 |  |  | Hea sag | [22 empty] | 3 | 3 | 84 | J40f medium | Head Soft |  |  | X-CARE | GEPACS | [9 empty] | manual |  | None
8 |  |  | Hea h150 :-D | [22 empty] | 1 | 1 | 121 | J40f medium | Head Soft |  |  | X-CARE | ROUTER | GEPACS | [8 empty] | manual |  | None
9 | 5140_Che_2 (Adult)
10 | 5140_Che_2 (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER | [9 empty] | None
11 |  | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER | [9 empty] | None
12 |  | art Che |  | 120 | 120 | 150 | 110 |  |  | On | 7 | On | CARE Dose4D | 10.12 | 30s | 25% | 50 |  | 0.5 |  | 25 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off |  |  |  |  | None
13 |  |  | art Che | [22 empty] | 3 | 2 | 145 | I41f medium | Chest |  |  |  | GEPACS | [9 empty] | manual |  | None
14 |  |  | art Che sag | [22 empty] | 3 | 2 | 191 | I41f medium | Bone |  |  |  | GEPACS | [9 empty] | manual |  | None
15 |  |  | art Che mip | [22 empty] | 10 | 3 | 121 | I41f medium | Lung |  |  |  | GEPACS | [9 empty] | manual |  | None
16 |  |  | art Che c140 3dt | [22 empty] | 1 | 1 | 289 | I50f medium sharp ASA | Lung |  |  |  | ROUTER | GEPACS | [8 empty] | manual |  | None
17 | 5170_Che_2_AbdPel (Adult)
18 | 5170_Che_2_AbdPel (Adult) | Topogram | Topogram |  | 100 |  | 393 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
19 |  | Topogram | Topogram |  | 100 |  | 393 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
20 |  | art Che |  | 120 | 120 | 150 | 110 |  |  | On | 7 | On | CARE Dose4D | 10.12 | 30s | 25% | 50 |  | 0.5 |  | 25 | 1 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
21 |  |  | art Che | [22 empty] | 3 | 2 | 145 | I41f medium | Chest |  |  |  | GEPACS | [9 empty] | manual
22 |  |  | art Che sag | [22 empty] | 3 | 2 | 191 | I41f medium | Bone |  |  |  | GEPACS | [9 empty] | manual
23 |  |  | art Che mip | [22 empty] | 10 | 3 | 121 | I41f medium | Lung |  |  |  | GEPACS | [9 empty] | manual
24 |  |  | art Che c140 3dt | [22 empty] | 1 | 1 | 289 | I50f medium sharp ASA | Lung |  |  |  | ROUTER | GEPACS | [8 empty] | manual
25 |  | ven AbdPel |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 60 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
26 |  |  | ven AbdPel a030 | [22 empty] | 3 | 3 | 131 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
27 |  |  | ven AbdPel cor | [22 empty] | 3 | 3 | 67 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
28 |  |  | ven AbdPel sag | [22 empty] | 3 | 3 | 121 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
29 |  |  | ven AbdPel a030 :-D | [22 empty] | 0.6 | 0.6 | 653 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
30 | 6030_AbdPel_A02_00A (Adult)
31 | 6030_AbdPel_A02_00A (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
32 |  | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
33 |  | ven AbdPel |  | 120 | 120 | 180 | 210 |  |  | On | 6 | On | CARE Dose4D | 12.15 | 30s | 25% | 50 |  | 0.5 |  | 90 | 0.8 |  | 3 | 128x0.6mm | [5 empty] | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 | [9 empty] | Off
34 |  |  | ven AbdPel a030 | [22 empty] | 3 | 3 | 131 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
35 |  |  | ven AbdPel cor | [22 empty] | 3 | 3 | 67 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
36 |  |  | ven AbdPel sag | [22 empty] | 3 | 3 | 121 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
37 |  |  | ven AbdPel a030 :-D | [22 empty] | 0.6 | 0.6 | 653 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
38 | 9060_Foo_sim_wo_cst (Adult)
39 | 9060_Foo_sim_wo_cst (Adult) | Topogram PA | Topogram PA 0.6 T20s |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.09 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER | [9 empty] | None
40 |  | Topogram LAT | Topogram LAT 0.6 T20s |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.09 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER | [9 empty] | None
41 |  | Foot simple |  |  | 120 | 120 | 250 |  |  |  | 3 | Off | CARE Dose | 24.04 | 30s | 25% | 60 |  | 1 |  | 2 | 1 |  | 0.6 | 16x0.3mm | [15 empty] | Off |  |  |  |  | None
42 |  |  | Foo bon l060 :-D | [22 empty] | 0.6 | 0.4 | 587 | U70u sharp | Bone |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual |  | None
43 |  |  | Foo bon l060 | [22 empty] | 2 | 2 | 117 | U70u sharp | Bone |  |  |  | GEPACS | [9 empty] | manual |  | None
44 |  |  | Foo :-D | [22 empty] | 0.6 | 0.4 | 584 | U30u medium smooth | Soft Tissue |  |  |  | ROUTER | GEPACS | [8 empty] | manual |  | None
45 |  |  | Foo | [22 empty] | 2 | 2 | 117 | U30u medium smooth | Soft Tissue |  |  |  | GEPACS | [9 empty] | manual |  | None

## Record Pattern 1
| Protocol name | Range name | Series description | Ref. kV | kV | Quality ref. mAs | (Eff.) mAs | (Eff.) mAs(Tube A) | (Eff.) mAs(Tube B) | CARE kV | Tissue of Interest (Slider position 1 - 12) | Dose modulation | CARE Dose type | CTDIvol (mGy) | FAST Adjust: Upper limit scan time | FAST Adjust: Lower limit max. mAs | Dose Notification value CTDIvol (mGy) | Dose Notification value DLP (mGy*cm) | Rotation time (s) | Scan time (s) | Delay (s) | Pitch | Feed per scan (mm) | Slice (mm) | Acq. | Effective Slice (mm) | Position increment (mm) | No. of images | Kernel/Algorithm | Window | API | Comment 1 | Comment 2 | Auto transfer 1 | Auto transfer 2 | Auto transfer 3 | syngo.via: Workflow | syngo.via: Data role | Scan start | Scan end | Pulsing | Pulsing start | Pulsing end | Best phase | Phase start | Multi phase
1 | 1150_Hea_wo_cst_N09_02A (Adult) | Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.26 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS |  | TRACKER | [9 empty] | None
2 |  | Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.26 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS |  | TRACKER | [9 empty] | None
10 | 5140_Che_2 (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER | [9 empty] | None
11 |  | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS | ROEMIR001a | TRACKER | [9 empty] | None
18 | 5170_Che_2_AbdPel (Adult) | Topogram | Topogram |  | 100 |  | 393 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
19 |  | Topogram | Topogram |  | 100 |  | 393 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
31 | 6030_AbdPel_A02_00A (Adult) | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
32 |  | Topogram | Topogram |  | 100 |  | 265 |  |  |  | 3 | Off | CARE Dose | 0.12 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | 400/40 Abd | 732b8b21-043d-46f3-80cf-17c5c2d1ddb4 |  |  | GEPACS |  | TRACKER
39 | 9060_Foo_sim_wo_cst (Adult) | Topogram PA | Topogram PA 0.6 T20s |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.09 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER | [9 empty] | None
40 |  | Topogram LAT | Topogram LAT 0.6 T20s |  | 100 |  | 96 |  |  |  | 3 | Off | CARE Dose | 0.09 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Body | [5 empty] | TRACKER | [9 empty] | None

## Sheet2
Protocol name | Range name | Series description | Ref. kV | kV | Quality ref. mAs | (Eff.) mAs | (Eff.) mAs(Tube A) | (Eff.) mAs(Tube B) | CARE kV | Tissue of Interest (Slider position 1 - 12) | Dose modulation | CARE Dose type | CTDIvol (mGy) | FAST Adjust: Upper limit scan time | FAST Adjust: Lower limit max. mAs | Dose Notification value CTDIvol (mGy) | Dose Notification value DLP (mGy*cm) | Rotation time (s) | Scan time (s) | Delay (s) | Pitch | Feed per scan (mm) | Slice (mm) | Acq. | Effective Slice (mm) | Position increment (mm) | No. of images | Kernel/Algorithm | Window | API | Comment 1 | Comment 2 | Auto transfer 1 | Auto transfer 2 | Auto transfer 3 | syngo.via: Workflow | syngo.via: Data role | Scan start | Scan end | Pulsing | Pulsing start | Pulsing end | Best phase | Phase start | Multi phase
1150_Hea_wo_cst_N09_02A (Adult)
1150_Hea_wo_cst_N09_02A (Adult) | Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.26 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS |  | TRACKER | [9 empty] | None
| Topogram | Topogram |  | 100 |  | 137 |  |  |  | 3 | Off | CARE Dose | 0.26 |  | 25% |  |  | 0.5 |  | 2 | 0 |  |  | 6x0.6mm | 0.6 |  | 1 | T20f medium | Topogram Head |  |  |  | GEPACS |  | TRACKER | [9 empty] | None
| Hea |  | 120 | 120 | 390 | 290 |  |  | Off | 3 | Off | CARE Dose4D | 44.37 | 15s | 25% | 80 |  | 0.5 |  | 2 | 0.6 |  | 4 | 128x0.6mm | [15 empty] | Off |  |  |  |  | None
|  | Hea h150 | [22 empty] | 4 | 4 | 31 | J40f medium | Head Soft |  |  | X-CARE | GEPACS | [9 empty] | manual |  | None
|  | Hea bon | [22 empty] | 4 | 4 | 31 | H60f sharp FR | Bone |  |  | X-CARE | GEPACS | [9 empty] | manual |  | None
|  | Hea cor | [22 empty] | 3 | 3 | 84 | J40f medium | Head Soft |  |  | X-CARE | GEPACS | [9 empty] | manual |  | None
|  | Hea sag | [22 empty] | 3 | 3 | 84 | J40f medium | Head Soft |  |  | X-CARE | GEPACS | [9 empty] | manual |  | None
|  | Hea h150 :-D | [22 empty] | 1 | 1 | 121 | J40f medium | Head Soft |  |  | X-CARE | ROUTER | GEPACS | [8 empty] | manual |  | None

## Record Pattern 3
| Protocol name | Range name | Series description | Ref. kV | kV | Quality ref. mAs | (Eff.) mAs | (Eff.) mAs(Tube A) | (Eff.) mAs(Tube B) | CARE kV | Tissue of Interest (Slider position 1 - 12) | Dose modulation | CARE Dose type | CTDIvol (mGy) | FAST Adjust: Upper limit scan time | FAST Adjust: Lower limit max. mAs | Dose Notification value CTDIvol (mGy) | Dose Notification value DLP (mGy*cm) | Rotation time (s) | Scan time (s) | Delay (s) | Pitch | Feed per scan (mm) | Slice (mm) | Acq. | Effective Slice (mm) | Position increment (mm) | No. of images | Kernel/Algorithm | Window | API | Comment 1 | Comment 2 | Auto transfer 1 | Auto transfer 2 | Auto transfer 3 | syngo.via: Workflow | syngo.via: Data role | Scan start | Scan end | Pulsing | Pulsing start | Pulsing end | Best phase | Phase start | Multi phase
4 |  |  | Hea h150 | [22 empty] | 4 | 4 | 31 | J40f medium | Head Soft |  |  | X-CARE | GEPACS | [9 empty] | manual |  | None
5 |  |  | Hea bon | [22 empty] | 4 | 4 | 31 | H60f sharp FR | Bone |  |  | X-CARE | GEPACS | [9 empty] | manual |  | None
6 |  |  | Hea cor | [22 empty] | 3 | 3 | 84 | J40f medium | Head Soft |  |  | X-CARE | GEPACS | [9 empty] | manual |  | None
7 |  |  | Hea sag | [22 empty] | 3 | 3 | 84 | J40f medium | Head Soft |  |  | X-CARE | GEPACS | [9 empty] | manual |  | None
8 |  |  | Hea h150 :-D | [22 empty] | 1 | 1 | 121 | J40f medium | Head Soft |  |  | X-CARE | ROUTER | GEPACS | [8 empty] | manual |  | None
13 |  |  | art Che | [22 empty] | 3 | 2 | 145 | I41f medium | Chest |  |  |  | GEPACS | [9 empty] | manual |  | None
14 |  |  | art Che sag | [22 empty] | 3 | 2 | 191 | I41f medium | Bone |  |  |  | GEPACS | [9 empty] | manual |  | None
15 |  |  | art Che mip | [22 empty] | 10 | 3 | 121 | I41f medium | Lung |  |  |  | GEPACS | [9 empty] | manual |  | None
16 |  |  | art Che c140 3dt | [22 empty] | 1 | 1 | 289 | I50f medium sharp ASA | Lung |  |  |  | ROUTER | GEPACS | [8 empty] | manual |  | None
21 |  |  | art Che | [22 empty] | 3 | 2 | 145 | I41f medium | Chest |  |  |  | GEPACS | [9 empty] | manual
22 |  |  | art Che sag | [22 empty] | 3 | 2 | 191 | I41f medium | Bone |  |  |  | GEPACS | [9 empty] | manual
23 |  |  | art Che mip | [22 empty] | 10 | 3 | 121 | I41f medium | Lung |  |  |  | GEPACS | [9 empty] | manual
24 |  |  | art Che c140 3dt | [22 empty] | 1 | 1 | 289 | I50f medium sharp ASA | Lung |  |  |  | ROUTER | GEPACS | [8 empty] | manual
26 |  |  | ven AbdPel a030 | [22 empty] | 3 | 3 | 131 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
27 |  |  | ven AbdPel cor | [22 empty] | 3 | 3 | 67 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
28 |  |  | ven AbdPel sag | [22 empty] | 3 | 3 | 121 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
29 |  |  | ven AbdPel a030 :-D | [22 empty] | 0.6 | 0.6 | 653 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
34 |  |  | ven AbdPel a030 | [22 empty] | 3 | 3 | 131 | I30f medium smooth | Abdomen |  |  |  | GEPACS | [9 empty] | manual
35 |  |  | ven AbdPel cor | [22 empty] | 3 | 3 | 67 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
36 |  |  | ven AbdPel sag | [22 empty] | 3 | 3 | 121 | I30f medium smooth | 400/40 Abd |  |  |  | GEPACS | [9 empty] | manual
37 |  |  | ven AbdPel a030 :-D | [22 empty] | 0.6 | 0.6 | 653 | I30f medium smooth | 400/40 Abd |  |  |  | ROUTER | GEPACS | MCHNeuroBodyAW | [7 empty] | manual
42 |  |  | Foo bon l060 :-D | [22 empty] | 0.6 | 0.4 | 587 | U70u sharp | Bone |  |  |  | ROUTER | GEPACS | MCH3DLabAW | [7 empty] | manual |  | None
43 |  |  | Foo bon l060 | [22 empty] | 2 | 2 | 117 | U70u sharp | Bone |  |  |  | GEPACS | [9 empty] | manual |  | None
44 |  |  | Foo :-D | [22 empty] | 0.6 | 0.4 | 584 | U30u medium smooth | Soft Tissue |  |  |  | ROUTER | GEPACS | [8 empty] | manual |  | None
45 |  |  | Foo | [22 empty] | 2 | 2 | 117 | U30u medium smooth | Soft Tissue |  |  |  | GEPACS | [9 empty] | manual |  | None